## Title
Varian TrueBeam
6MV Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables |  | 2. Plots
| Open field |  | Open Field
| Open Field Depth Doses |  | Wedge 15
| Open Field Profiles at 1.5cm |  | Wedge 30 Plots
| Open Field Profiles at 5cm |  | Wedge 45 Plots
| Open Field Profiles at 10cm |  | Wedge 60 Plots
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
Wedge 15
Wedge 15 Depth Doses
SSD80 Wedge 15 Depth Doses
SSD120 Wedge 15 Depth Doses
Wedge 15 Lateral Profiles
Wedge 15 Longitudinal Profiles
Wedge 15 Output Factors
Wedge 30
Wedge 30 Depth Doses
SSD80 Wedge 30 Depth Doses
SSD120 Wedge 30 Depth Doses
Wedge 30 Lateral Profiles
Wedge 30 Longitudinal Profiles
Wedge 30 Output Factors
Wedge 45
Wedge 45 Depth Doses
SSD80 Wedge 45 Depth Doses
Wedge 45 Longitudinal Profiles
Wedge 45 Lateral Profiles
Wedge 45 Longitudinal Profiles
Wedge 45 Output Factors
Wedge 60
Wedge 60 Depth Doses
SSD80 Wedge 60 Depth Doses
SSD120 Wedge 60 Depth Doses
Wedge 60 Lateral Profiles
Wedge 60 Longitudinal Profiles
Wedge 60 Output Factors

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 47.8 | 48.3 | 49.8 | 52.1 | 53.6 | 62.2 | 68.5 | 72.2
0.1 | 53.075 | 53.8 | 55.325 | 57.025 | 59.1 | 66.6 | 73 | 76.35
0.2 | 62.05 | 61.8 | 63.475 | 65.6 | 66.225 | 74.275 | 78.475 | 81.56
0.3 | 72.325 | 72.15 | 73.725 | 74.6 | 75.7 | 81.4 | 85.25 | 87.6
0.4 | 79.3 | 79.95 | 80.4 | 81.5 | 82.775 | 87.2 | 90 | 91.6
0.5 | 85.4 | 85.675 | 86.1 | 86.8 | 87.95 | 91.4 | 93.55 | 94.5
0.6 | 89.6 | 89.95 | 90.3 | 90.9 | 91.525 | 94.2 | 95.825 | 96.55
0.7 | 93 | 93.02 | 93.5 | 93.9 | 94.275 | 96.3 | 97.5333 | 97.9
0.8 | 95.6 | 95.3 | 95.6 | 96 | 96.2 | 97.7333 | 98.5 | 98.82
0.9 | 97.05 | 96.95 | 97.3 | 97.4333 | 97.65 | 98.72 | 99.3 | 99.35
1 | 98.325 | 98.2333 | 98.4333 | 98.5 | 98.8 | 99.525 | 99.7 | 99.82
1.1 | 99.225 | 99.25 | 99.2 | 99.4 | 99.45 | 99.9 | 100 | 99.94
1.2 | 99.75 | 99.75 | 99.75 | 99.8 | 99.8 | 100 | 100 | 100
1.3 | 100 | 100 | 99.975 | 100 | 100 | 100 | 100 | 100
1.4 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.8667
1.5 | 100 | 100 | 100 | 100 | 100 | 99.9 | 99.825 | 99.7
1.6 | 99.9 | 99.825 | 99.8333 | 99.8 | 99.9 | 99.6333 | 99.575 | 99.45
1.7 | 99.6 | 99.6 | 99.6667 | 99.6 | 99.6333 | 99.4 | 99.325 | 99.225
1.8 | 99.25 | 99.35 | 99.3 | 99.3 | 99.4 | 99.1 | 99.1 | 98.84
1.9 | 98.8 | 98.9 | 99 | 98.98 | 98.95 | 98.8 | 98.64 | 98.5
2 | 98.375 | 98.5 | 98.7 | 98.575 | 98.7333 | 98.375 | 98.4 | 98.2333
2.1 | 97.85 | 98 | 98.22 | 98.25 | 98.3 | 98.125 | 98.0333 | 97.9
2.2 | 97.35 | 97.6 | 97.85 | 97.675 | 97.875 | 97.7 | 97.8 | 97.6
2.3 | 96.95 | 97.1 | 97.375 | 97.425 | 97.5 | 97.325 | 97.35 | 97.26
2.4 | 96.425 | 96.65 | 96.875 | 97.1 | 97.1 | 97.05 | 97 | 96.95
2.5 | 95.925 | 96.175 | 96.35 | 96.525 | 96.75 | 96.575 | 96.6667 | 96.55
2.6 | 95.4 | 95.65 | 95.9 | 96.2 | 96.3 | 96.3 | 96.3 | 96.3
2.7 | 94.85 | 95.075 | 95.5 | 95.7 | 95.84 | 95.98 | 95.95 | 95.975
2.8 | 94.3 | 94.58 | 95 | 95.2 | 95.5 | 95.55 | 95.55 | 95.625
2.9 | 93.7 | 94.15 | 94.5 | 94.8 | 95.05 | 95.1 | 95.275 | 95.28
3 | 93.2 | 93.5 | 94.2333 | 94.3 | 94.55 | 94.8 | 94.925 | 94.92
3.1 | 92.7333 | 93.05 | 93.7 | 93.95 | 94.15 | 94.4 | 94.6 | 94.5333
3.2 | 92.05 | 92.6 | 93.16 | 93.56 | 93.65 | 94 | 94.25 | 94.2333
3.3 | 91.625 | 92.08 | 92.66 | 92.96 | 93.36 | 93.6 | 93.9 | 93.9167
3.4 | 90.95 | 91.55 | 92.25 | 92.5 | 92.85 | 93.2 | 93.55 | 93.575
3.5 | 90.375 | 90.975 | 91.75 | 92.15 | 92.5 | 92.95 | 93.12 | 93.34
3.6 | 90 | 90.55 | 91.3 | 91.75 | 92.05 | 92.5 | 92.8 | 93
3.7 | 89.4667 | 90.15 | 90.75 | 91.24 | 91.6 | 92.15 | 92.45 | 92.625
3.8 | 88.9 | 89.55 | 90.325 | 90.8 | 91.1 | 91.75 | 92 | 92.3
3.9 | 88.4 | 88.925 | 89.86 | 90.4 | 90.7 | 91.375 | 91.76 | 91.96
4 | 87.8 | 88.375 | 89.28 | 89.95 | 90.2667 | 90.95 | 91.38 | 91.575
4.1 | 87.25 | 88.1 | 88.85 | 89.55 | 89.86 | 90.65 | 91 | 91.28
4.2 | 86.8 | 87.6 | 88.4333 | 89.15 | 89.46 | 90.25 | 90.7 | 90.94
4.3 | 86.2333 | 87.1 | 87.9 | 88.625 | 89.05 | 89.9 | 90.3 | 90.58
4.4 | 85.6 | 86.45 | 87.4333 | 88.15 | 88.55 | 89.5 | 89.95 | 90.2
4.5 | 85.14 | 85.9 | 87.025 | 87.8 | 88.125 | 89.0667 | 89.6 | 89.86
4.6 | 84.55 | 85.48 | 86.55 | 87.4 | 87.7 | 88.7333 | 89.25 | 89.54
4.7 | 84.1667 | 84.875 | 86.15 | 86.9 | 87.2 | 88.3 | 89 | 89.2
4.8 | 83.65 | 84.45 | 85.5 | 86.35 | 86.8 | 88 | 88.55 | 88.9
4.9 | 83.05 | 84.025 | 85.15 | 85.9 | 86.375 | 87.6667 | 88.25 | 88.45
5 | 82.58 | 83.45 | 84.7 | 85.52 | 86.125 | 87.275 | 87.75 | 88.2
5.1 | 82.025 | 83 | 84.2 | 85.05 | 85.625 | 86.88 | 87.5 | 87.8333
5.2 | 81.5667 | 82.4 | 83.75 | 84.525 | 85.25 | 86.55 | 87.1 | 87.46
5.3 | 81 | 81.9 | 83.375 | 84.2 | 84.8 | 86.1667 | 86.76 | 87.16
5.4 | 80.5667 | 81.5 | 82.95 | 83.7 | 84.4 | 85.7 | 86.425 | 86.75
5.5 | 79.9 | 81 | 82.46 | 83.35 | 83.9 | 85.3667 | 86.05 | 86.42
5.6 | 79.4667 | 80.65 | 81.9 | 82.85 | 83.46 | 85.05 | 85.68 | 86.1
5.7 | 79.1 | 80.15 | 81.4 | 82.35 | 83.15 | 84.675 | 85.4 | 85.76
5.8 | 78.7 | 79.64 | 81.0333 | 81.98 | 82.775 | 84.275 | 85.05 | 85.42
5.9 | 78.1 | 79.08 | 80.625 | 81.55 | 82.325 | 83.925 | 84.7 | 84.96
6 | 77.62 | 78.7 | 80.125 | 81.075 | 81.84 | 83.58 | 84.25 | 84.675
6.1 | 77.15 | 78.2667 | 79.6 | 80.8333 | 81.5 | 83.2 | 84.025 | 84.4333
6.2 | 76.65 | 77.775 | 79.3 | 80.25 | 81.1 | 82.75 | 83.675 | 84
6.3 | 76.15 | 77.275 | 78.8 | 79.9 | 80.6 | 82.4 | 83.25 | 83.7
6.4 | 75.55 | 76.95 | 78.3 | 79.35 | 80.15 | 82.1 | 82.88 | 83.3833
6.5 | 75.2 | 76.4 | 77.94 | 79.0333 | 79.825 | 81.7333 | 82.6 | 83.0667
6.6 | 74.75 | 75.8 | 77.45 | 78.6667 | 79.45 | 81.275 | 82.3 | 82.7
6.7 | 74.2 | 75.4 | 77.08 | 78.15 | 79.08 | 80.975 | 81.9667 | 82.36
6.8 | 73.82 | 75.05 | 76.6333 | 77.725 | 78.6667 | 80.6333 | 81.45 | 82.025
6.9 | 73.375 | 74.525 | 76.1667 | 77.4 | 78.2 | 80.225 | 81.2 | 81.75
7 | 72.875 | 74.05 | 75.7 | 77.025 | 77.84 | 79.975 | 80.8667 | 81.35
7.1 | 72.45 | 73.64 | 75.25 | 76.475 | 77.45 | 79.55 | 80.6 | 80.9667
7.2 | 72 | 73.2 | 74.85 | 76.05 | 77.05 | 79.175 | 80.2 | 80.7
7.3 | 71.45 | 72.7 | 74.55 | 75.8 | 76.575 | 78.9 | 79.85 | 80.4
7.4 | 71 | 72.4 | 74.025 | 75.4 | 76.18 | 78.4 | 79.4333 | 80.1286
7.5 | 70.6333 | 71.95 | 73.6 | 74.9 | 75.8 | 78.15 | 79.2667 | 79.75
7.6 | 70.175 | 71.45 | 73.15 | 74.52 | 75.4 | 77.825 | 78.875 | 79.4
7.7 | 69.85 | 71.125 | 72.8 | 74.1333 | 75 | 77.425 | 78.525 | 79.0667
7.8 | 69.28 | 70.5667 | 72.45 | 73.72 | 74.675 | 77.075 | 78.16 | 78.775
7.9 | 69 | 70.1 | 72.05 | 73.275 | 74.35 | 76.825 | 77.85 | 78.5667
8 | 68.6 | 69.7 | 71.65 | 72.95 | 73.85 | 76.3 | 77.54 | 78.2
8.1 | 68.15 | 69.2 | 71.15 | 72.6 | 73.425 | 76.05 | 77.225 | 77.78
8.2 | 67.6667 | 68.9 | 70.7 | 72.2 | 73.2 | 75.8 | 76.825 | 77.52
8.3 | 67.1 | 68.35 | 70.25 | 71.75 | 72.8 | 75.35 | 76.6 | 77.22
8.4 | 66.775 | 67.94 | 69.9 | 71.3 | 72.35 | 74.94 | 76.18 | 76.9
8.5 | 66.44 | 67.56 | 69.475 | 70.95 | 72.05 | 74.65 | 75.9 | 76.5667
8.6 | 66 | 67.2 | 69.05 | 70.525 | 71.6 | 74.225 | 75.5 | 76.3
8.7 | 65.5 | 66.8 | 68.65 | 70.16 | 71.3 | 73.875 | 75.2 | 75.96
8.8 | 65.04 | 66.32 | 68.24 | 69.8 | 70.8333 | 73.55 | 74.8667 | 75.65
8.9 | 64.675 | 65.975 | 67.8667 | 69.3 | 70.5333 | 73.2 | 74.65 | 75.25
9 | 64.25 | 65.55 | 67.6 | 68.95 | 70.2 | 72.95 | 74.25 | 75
9.1 | 63.85 | 65.1333 | 67.1 | 68.7 | 69.8 | 72.6 | 73.825 | 74.6833
9.2 | 63.525 | 64.7 | 66.725 | 68.2 | 69.25 | 72.2667 | 73.575 | 74.3
9.3 | 63.06 | 64.2 | 66.25 | 67.85 | 68.92 | 71.85 | 73.15 | 73.975
9.4 | 62.7 | 63.85 | 65.85 | 67.525 | 68.575 | 71.6 | 72.9 | 73.7333
9.5 | 62.2 | 63.5 | 65.4 | 67.05 | 68.3 | 71.225 | 72.6 | 73.3
9.6 | 61.8 | 63.0667 | 65.15 | 66.65 | 67.775 | 70.75 | 72.2333 | 73.0333
9.7 | 61.4 | 62.75 | 64.6 | 66.28 | 67.45 | 70.5667 | 71.85 | 72.7667
9.8 | 61.1 | 62.4 | 64.3 | 65.9 | 67.2 | 70.2 | 71.6 | 72.5
9.9 | 60.76 | 61.925 | 63.9 | 65.5 | 66.75 | 69.85 | 71.3 | 72.0667
10 | 60.45 | 61.54 | 63.575 | 65.22 | 66.4 | 69.54 | 70.9667 | 71.78
10.1 | 59.975 | 61.15 | 63.325 | 64.7 | 66.02 | 69.275 | 70.6667 | 71.5
10.2 | 59.5333 | 60.8 | 62.775 | 64.375 | 65.625 | 68.85 | 70.3 | 71.1667
10.3 | 59.075 | 60.4 | 62.525 | 64.125 | 65.2333 | 68.4667 | 70 | 70.825
10.4 | 58.775 | 60.025 | 62.1 | 63.775 | 64.925 | 68.175 | 69.7 | 70.5333
10.5 | 58.54 | 59.675 | 61.7 | 63.275 | 64.7 | 67.84 | 69.3667 | 70.28
10.6 | 58.1 | 59.35 | 61.4 | 63 | 64.25 | 67.6 | 69.1 | 69.86
10.7 | 57.7 | 58.975 | 61 | 62.6 | 64 | 67.2 | 68.7333 | 69.5667
10.8 | 57.2 | 58.58 | 60.66 | 62.2333 | 63.52 | 66.9333 | 68.5 | 69.3
10.9 | 56.9333 | 58.2333 | 60.275 | 61.9 | 63.3 | 66.575 | 68.15 | 69.04
11 | 56.575 | 57.85 | 59.85 | 61.65 | 62.96 | 66.25 | 67.8 | 68.75
11.1 | 56.1 | 57.55 | 59.575 | 61.3 | 62.525 | 65.875 | 67.6 | 68.4167
11.2 | 55.9 | 57.14 | 59.25 | 60.9 | 62.175 | 65.64 | 67.3 | 68.14
11.3 | 55.55 | 56.775 | 58.9 | 60.55 | 61.825 | 65.3 | 67.025 | 67.8333
11.4 | 55.18 | 56.525 | 58.6 | 60.1 | 61.5 | 64.95 | 66.675 | 67.6
11.5 | 54.8 | 56.15 | 58.2 | 59.825 | 61.15 | 64.6 | 66.425 | 67.3
11.6 | 54.6 | 55.75 | 57.9 | 59.55 | 60.9 | 64.375 | 66.1 | 66.95
11.7 | 54.16 | 55.45 | 57.5333 | 59.15 | 60.45 | 64.025 | 65.75 | 66.725
11.8 | 53.925 | 55.1 | 57.2 | 58.9 | 60.25 | 63.7 | 65.5 | 66.4333
11.9 | 53.45 | 54.75 | 56.8 | 58.5 | 59.875 | 63.4333 | 65.2 | 66.06
12 | 53.15 | 54.4 | 56.56 | 58.1 | 59.525 | 63.2 | 64.9 | 65.75
12.1 | 52.775 | 54.05 | 56.1667 | 57.875 | 59.18 | 62.8667 | 64.54 | 65.55
12.2 | 52.525 | 53.7 | 55.775 | 57.525 | 58.9 | 62.6 | 64.34 | 65.2667
12.3 | 52.2 | 53.45 | 55.525 | 57.26 | 58.6 | 62.25 | 63.925 | 64.9333
12.4 | 51.9 | 53.025 | 55.2 | 56.85 | 58.25 | 62.02 | 63.7 | 64.62
12.5 | 51.5 | 52.75 | 54.85 | 56.55 | 57.975 | 61.625 | 63.4 | 64.4143
12.6 | 51.3 | 52.5 | 54.5 | 56.2 | 57.725 | 61.35 | 63.15 | 64.18
12.7 | 50.9 | 52.1 | 54.2 | 55.9333 | 57.3 | 61.1 | 62.9 | 63.85
12.8 | 50.65 | 51.8 | 53.925 | 55.625 | 56.9667 | 60.75 | 62.7 | 63.5
12.9 | 50.25 | 51.425 | 53.6333 | 55.275 | 56.7 | 60.52 | 62.225 | 63.3
13 | 50 | 51.075 | 53.25 | 55.05 | 56.4 | 60.175 | 62.05 | 63.05
13.1 | 49.6 | 50.825 | 52.95 | 54.75 | 56.075 | 59.875 | 61.725 | 62.72
13.2 | 49.3 | 50.5667 | 52.7 | 54.3 | 55.8 | 59.55 | 61.375 | 62.525
13.3 | 48.95 | 50.1 | 52.35 | 54 | 55.45 | 59.175 | 61.15 | 62.2333
13.4 | 48.6 | 49.8667 | 52 | 53.7 | 55.125 | 58.9333 | 60.9 | 61.8
13.5 | 48.36 | 49.7 | 51.65 | 53.4 | 54.8 | 58.6667 | 60.7 | 61.5667
13.6 | 48.05 | 49.275 | 51.3667 | 53.1 | 54.55 | 58.4 | 60.25 | 61.3
13.7 | 47.725 | 48.925 | 51.1333 | 52.75 | 54.2 | 58.15 | 60.0667 | 61.075
13.8 | 47.475 | 48.675 | 50.85 | 52.475 | 53.8333 | 57.86 | 59.8 | 60.74
13.9 | 47.25 | 48.35 | 50.48 | 52.14 | 53.55 | 57.55 | 59.4333 | 60.54
14 | 47 | 48.1 | 50.2 | 51.9 | 53.32 | 57.275 | 59.2 | 60.2667
14.1 | 46.65 | 47.8667 | 49.9 | 51.55 | 53.0333 | 57.025 | 58.925 | 60.0333
14.2 | 46.3 | 47.5 | 49.55 | 51.28 | 52.7667 | 56.7 | 58.675 | 59.7
14.3 | 46.1 | 47.26 | 49.25 | 51 | 52.4 | 56.5 | 58.4 | 59.4333
14.4 | 45.8 | 47.025 | 48.975 | 50.6667 | 52.15 | 56.2 | 58.1 | 59.14
14.5 | 45.52 | 46.625 | 48.65 | 50.45 | 51.9 | 55.9 | 57.8 | 58.9
14.6 | 45.225 | 46.35 | 48.375 | 50.1 | 51.5 | 55.55 | 57.5 | 58.675
14.7 | 44.925 | 46.125 | 48.125 | 49.875 | 51.26 | 55.3 | 57.275 | 58.3667
14.8 | 44.675 | 45.75 | 47.85 | 49.55 | 51.025 | 55.075 | 56.95 | 58.1167
14.9 | 44.3 | 45.44 | 47.6333 | 49.35 | 50.775 | 54.825 | 56.75 | 57.8
15 | 44.05 | 45.2 | 47.3 | 49.025 | 50.425 | 54.575 | 56.425 | 57.6286
15.1 | 43.8 | 45 | 47.05 | 48.7 | 50.075 | 54.225 | 56.2 | 57.3143
15.2 | 43.5 | 44.66 | 46.825 | 48.5 | 49.825 | 53.98 | 55.9 | 57.0286
15.3 | 43.225 | 44.5 | 46.475 | 48.175 | 49.5 | 53.65 | 55.7 | 56.82
15.4 | 42.975 | 44.2 | 46.125 | 47.925 | 49.3 | 53.4 | 55.4 | 56.575
15.5 | 42.725 | 43.85 | 45.875 | 47.675 | 49.02 | 53.15 | 55.2 | 56.36
15.6 | 42.575 | 43.6 | 45.625 | 47.325 | 48.775 | 52.9 | 54.8667 | 56.1
15.7 | 42.25 | 43.35 | 45.4 | 47.1 | 48.5 | 52.675 | 54.7 | 55.8286
15.8 | 41.9 | 43 | 45.2 | 46.725 | 48.25 | 52.35 | 54.45 | 55.5
15.9 | 41.7 | 42.8 | 44.8333 | 46.48 | 47.9 | 52.15 | 54.2 | 55.3
16 | 41.4 | 42.55 | 44.625 | 46.25 | 47.6 | 51.875 | 53.88 | 55.08
16.1 | 41.1667 | 42.275 | 44.3 | 45.9 | 47.45 | 51.65 | 53.65 | 54.8667
16.2 | 40.975 | 41.925 | 44.025 | 45.75 | 47.1 | 51.4 | 53.4 | 54.6167
16.3 | 40.725 | 41.775 | 43.775 | 45.5 | 46.9 | 51.1 | 53.1 | 54.3333
16.4 | 40.475 | 41.5667 | 43.55 | 45.25 | 46.5667 | 50.8 | 52.9 | 54.05
16.5 | 40.14 | 41.3 | 43.2667 | 45.02 | 46.3 | 50.6 | 52.7 | 53.875
16.6 | 39.95 | 41.075 | 43 | 44.675 | 46.1 | 50.325 | 52.3667 | 53.62
16.7 | 39.68 | 40.8 | 42.75 | 44.425 | 45.9 | 50.1 | 52.2 | 53.4
16.8 | 39.45 | 40.55 | 42.52 | 44.1 | 45.7 | 49.85 | 51.95 | 53.06
16.9 | 39.275 | 40.24 | 42.275 | 43.95 | 45.35 | 49.5 | 51.7333 | 52.8
17 | 39.02 | 40 | 42.05 | 43.75 | 45.175 | 49.375 | 51.45 | 52.5667
17.1 | 38.7 | 39.8 | 41.76 | 43.4 | 44.9 | 49.05 | 51.2 | 52.3667
17.2 | 38.54 | 39.55 | 41.45 | 43.15 | 44.6 | 48.9 | 50.95 | 52.1667
17.3 | 38.3 | 39.3 | 41.2 | 42.9 | 44.34 | 48.65 | 50.7 | 51.88
17.4 | 38.0667 | 39.1 | 41.05 | 42.775 | 44.1 | 48.4 | 50.475 | 51.68
17.5 | 37.8 | 38.9 | 40.8 | 42.4667 | 43.85 | 48.175 | 50.225 | 51.38
17.6 | 37.6 | 38.6333 | 40.475 | 42.275 | 43.6 | 47.925 | 49.98 | 51.2333
17.7 | 37.3333 | 38.4 | 40.3 | 41.95 | 43.4 | 47.675 | 49.8 | 51
17.8 | 37.1 | 38.225 | 40.025 | 41.75 | 43.125 | 47.44 | 49.6 | 50.7667
17.9 | 36.975 | 37.9 | 39.875 | 41.525 | 42.9 | 47.1 | 49.3667 | 50.56
18 | 36.725 | 37.75 | 39.625 | 41.275 | 42.725 | 46.95 | 49.1 | 50.325
18.1 | 36.5 | 37.6 | 39.4 | 41.05 | 42.46 | 46.7 | 48.875 | 49.96
18.2 | 36.25 | 37.25 | 39.2 | 40.84 | 42.25 | 46.5 | 48.625 | 49.825
18.3 | 36 | 37.0667 | 38.9 | 40.6 | 42 | 46.25 | 48.375 | 49.6
18.4 | 35.825 | 36.8 | 38.75 | 40.35 | 41.75 | 46 | 48.2 | 49.34
18.5 | 35.575 | 36.6 | 38.42 | 40.2 | 41.475 | 45.825 | 47.9 | 49.1333
18.6 | 35.325 | 36.3333 | 38.275 | 39.875 | 41.325 | 45.5 | 47.65 | 48.95
18.7 | 35.075 | 36.125 | 38.025 | 39.62 | 41.075 | 45.3 | 47.4 | 48.6714
18.8 | 34.925 | 35.9 | 37.775 | 39.4 | 40.825 | 45.1 | 47.25 | 48.4333
18.9 | 34.7 | 35.7 | 37.625 | 39.175 | 40.575 | 44.9 | 47.02 | 48.2333
19 | 34.55 | 35.5 | 37.28 | 38.925 | 40.325 | 44.6 | 46.775 | 47.95
19.1 | 34.3333 | 35.3 | 37.15 | 38.675 | 40.2 | 44.425 | 46.6 | 47.86
19.2 | 34.1667 | 35.1 | 37 | 38.525 | 39.9 | 44.18 | 46.375 | 47.625
19.3 | 33.85 | 34.875 | 36.66 | 38.3 | 39.7 | 43.9 | 46.14 | 47.35
19.4 | 33.675 | 34.625 | 36.475 | 38.05 | 39.4 | 43.7 | 45.9 | 47.15
19.5 | 33.5 | 34.4 | 36.2 | 37.8 | 39.25 | 43.5 | 45.7 | 46.9
19.6 | 33.26 | 34.24 | 36 | 37.6 | 38.925 | 43.3 | 45.45 | 46.68
19.7 | 33.1 | 34.1 | 35.8 | 37.425 | 38.78 | 43 | 45.325 | 46.48
19.8 | 32.85 | 33.75 | 35.7 | 37.275 | 38.5 | 42.8333 | 45.075 | 46.3
19.9 | 32.675 | 33.6 | 35.35 | 37.025 | 38.35 | 42.5667 | 44.825 | 46.08
20 | 32.45 | 33.375 | 35.2 | 36.8 | 38.1 | 42.4 | 44.525 | 45.7667
20.1 | 32.3 | 33.125 | 34.95 | 36.6333 | 37.925 | 42.1667 | 44.3 | 45.6
20.2 | 32.05 | 32.975 | 34.82 | 36.35 | 37.7 | 42 | 44.2 | 45.4333
20.3 | 31.8 | 32.825 | 34.575 | 36.1 | 37.5 | 41.75 | 44 | 45.1833
20.4 | 31.7 | 32.6 | 34.425 | 35.975 | 37.24 | 41.5 | 43.75 | 45.0167
20.5 | 31.5 | 32.45 | 34.175 | 35.725 | 37.1 | 41.25 | 43.5 | 44.7333
20.6 | 31.3 | 32.2 | 34 | 35.5 | 36.9 | 41.1 | 43.275 | 44.52
20.7 | 31.125 | 32.05 | 33.8 | 35.3 | 36.6 | 40.86 | 43.025 | 44.3667
20.8 | 31 | 31.85 | 33.55 | 35.15 | 36.45 | 40.625 | 42.875 | 44.16
20.9 | 30.775 | 31.7 | 33.4 | 34.9 | 36.3 | 40.5 | 42.7 | 43.9
21 | 30.6 | 31.4 | 33.26 | 34.75 | 36.1 | 40.3 | 42.5333 | 43.7
21.1 | 30.4 | 31.3 | 33.1 | 34.6 | 35.925 | 40.1333 | 42.2667 | 43.5833
21.2 | 30.25 | 31.05 | 32.8 | 34.375 | 35.7 | 39.9 | 42.12 | 43.4
21.3 | 30.1 | 30.9 | 32.625 | 34.125 | 35.55 | 39.75 | 41.875 | 43.1714
21.4 | 29.9 | 30.75 | 32.4 | 34 | 35.3 | 39.5 | 41.7 | 42.9333
21.5 | 29.6 | 30.55 | 32.3 | 33.8 | 35.1 | 39.26 | 41.575 | 42.76
21.6 | 29.525 | 30.4 | 32.1667 | 33.6 | 34.9 | 39.025 | 41.25 | 42.5
21.7 | 29.375 | 30.1667 | 31.925 | 33.4 | 34.725 | 38.875 | 41.08 | 42.3667
21.8 | 29.15 | 30 | 31.7 | 33.2 | 34.575 | 38.725 | 40.95 | 42.2
21.9 | 29 | 29.8 | 31.525 | 33.025 | 34.325 | 38.475 | 40.7 | 41.9333
22 | 28.85 | 29.6 | 31.4 | 32.875 | 34.2 | 38.35 | 40.55 | 41.75
22.1 | 28.6 | 29.45 | 31.2 | 32.625 | 34 | 38.1 | 40.4 | 41.58
22.2 | 28.45 | 29.3 | 31.1 | 32.475 | 33.8 | 38 | 40.15 | 41.38
22.3 | 28.3 | 29.15 | 30.85 | 32.3 | 33.6 | 37.76 | 39.9 | 41.16
22.4 | 28.175 | 28.9 | 30.72 | 32.2 | 33.45 | 37.62 | 39.8 | 40.94
22.5 | 28.025 | 28.775 | 30.4 | 31.95 | 33.3 | 37.3667 | 39.6 | 40.85
22.6 | 27.775 | 28.625 | 30.225 | 31.8 | 33.0333 | 37.2 | 39.375 | 40.6
22.7 | 27.625 | 28.4 | 30.1 | 31.6 | 32.85 | 36.9667 | 39.2 | 40.3833
22.8 | 27.5 | 28.3 | 29.925 | 31.4 | 32.7 | 36.8 | 39.05 | 40.2286
22.9 | 27.3 | 28.1 | 29.8 | 31.2 | 32.5 | 36.55 | 38.8 | 40.0667
23 | 27.2 | 27.95 | 29.6 | 31.025 | 32.35 | 36.4 | 38.6 | 39.9
23.1 | 27.06 | 27.8 | 29.42 | 30.8 | 32.1 | 36.25 | 38.425 | 39.7
23.2 | 26.8 | 27.56 | 29.25 | 30.7 | 31.925 | 36 | 38.3 | 39.5
23.3 | 26.75 | 27.425 | 29.075 | 30.5 | 31.8 | 35.9 | 38.05 | 39.26
23.4 | 26.525 | 27.3 | 28.875 | 30.35 | 31.625 | 35.7 | 37.85 | 39.05
23.5 | 26.325 | 27.125 | 28.7667 | 30.2 | 31.475 | 35.525 | 37.7 | 38.98
23.6 | 26.275 | 27 | 28.575 | 30 | 31.275 | 35.3 | 37.55 | 38.8
23.7 | 26.1 | 26.8 | 28.4 | 29.825 | 31.025 | 35.125 | 37.32 | 38.6
23.8 | 25.9 | 26.7 | 28.3 | 29.575 | 30.9 | 34.975 | 37.175 | 38.4167
23.9 | 25.7 | 26.55 | 28.1 | 29.45 | 30.7 | 34.84 | 37.025 | 38.2
24 | 25.5333 | 26.4 | 27.9 | 29.3 | 30.6 | 34.6 | 36.8 | 38.0833
24.1 | 25.4 | 26.2 | 27.8 | 29.2 | 30.4 | 34.4333 | 36.6 | 37.9167
24.2 | 25.3 | 26.1 | 27.6 | 29 | 30.2333 | 34.26 | 36.45 | 37.74
24.3 | 25.1 | 25.875 | 27.5 | 28.8 | 30.075 | 34.1 | 36.3 | 37.55
24.4 | 24.95 | 25.725 | 27.3 | 28.7 | 29.9 | 33.9 | 36.1 | 37.3667
24.5 | 24.8 | 25.575 | 27.175 | 28.5 | 29.7667 | 33.725 | 35.9 | 37.18
24.6 | 24.7 | 25.5 | 27 | 28.32 | 29.6 | 33.575 | 35.8 | 37
24.7 | 24.6 | 25.25 | 26.8 | 28.175 | 29.4 | 33.325 | 35.6 | 36.88
24.8 | 24.4 | 25.2 | 26.7 | 28 | 29.3 | 33.18 | 35.44 | 36.65
24.9 | 24.2667 | 24.96 | 26.5 | 27.9 | 29.1 | 33 | 35.2 | 36.5
25 | 24.075 | 24.8667 | 26.325 | 27.7 | 28.9 | 32.85 | 35.1 | 36.3
25.1 | 23.925 | 24.7 | 26.175 | 27.54 | 28.8 | 32.8 | 34.925 | 36.1
25.2 | 23.8 | 24.575 | 26.05 | 27.4 | 28.6 | 32.55 | 34.775 | 35.98
25.3 | 23.6 | 24.4 | 25.9 | 27.25 | 28.44 | 32.375 | 34.6 | 35.78
25.4 | 23.5 | 24.25 | 25.75 | 27.1 | 28.3 | 32.225 | 34.4 | 35.7
25.5 | 23.4 | 24.1 | 25.6 | 26.9 | 28.14 | 32.08 | 34.2 | 35.4
25.6 | 23.3 | 23.95 | 25.4 | 26.775 | 28 | 31.9 | 34 | 35.3
25.7 | 23.1667 | 23.8667 | 25.325 | 26.6 | 27.8333 | 31.7 | 33.9 | 35.14
25.8 | 23.05 | 23.675 | 25.175 | 26.475 | 27.7 | 31.6 | 33.75 | 34.96
25.9 | 22.8 | 23.6 | 25 | 26.325 | 27.55 | 31.5 | 33.62 | 34.775
26 | 22.7 | 23.475 | 24.875 | 26.24 | 27.425 | 31.3 | 33.375 | 34.56
26.1 | 22.6 | 23.325 | 24.7 | 26.1 | 27.2 | 31.125 | 33.225 | 34.4
26.2 | 22.5 | 23.1667 | 24.6 | 25.95 | 27.1 | 30.9 | 33.2 | 34.3
26.3 | 22.3 | 23 | 24.5 | 25.8 | 26.9 | 30.8 | 32.9333 | 34.1
26.4 | 22.2 | 22.9 | 24.3 | 25.6 | 26.775 | 30.65 | 32.8 | 34
26.5 | 22.025 | 22.7333 | 24.2 | 25.5 | 26.625 | 30.5 | 32.6 | 33.8429
26.6 | 21.9 | 22.625 | 24.025 | 25.32 | 26.4667 | 30.36 | 32.425 | 33.5714
26.7 | 21.825 | 22.475 | 23.9 | 25.175 | 26.3667 | 30.225 | 32.275 | 33.4167
26.8 | 21.7 | 22.325 | 23.7 | 25.025 | 26.2 | 30 | 32.125 | 33.35
26.9 | 21.5 | 22.2 | 23.64 | 24.875 | 26.0667 | 29.925 | 32 | 33.14
27 | 21.425 | 22.1 | 23.5 | 24.74 | 25.9 | 29.7 | 31.85 | 33
27.1 | 21.3 | 22 | 23.36 | 24.6 | 25.8 | 29.55 | 31.7 | 32.86
27.2 | 21.2 | 21.8 | 23.3 | 24.5 | 25.6 | 29.4 | 31.5 | 32.76
27.3 | 21.0667 | 21.7 | 23.125 | 24.375 | 25.525 | 29.2 | 31.4 | 32.525
27.4 | 20.9 | 21.6 | 22.975 | 24.225 | 25.4 | 29.1 | 31.25 | 32.3667
27.5 | 20.8667 | 21.5 | 22.84 | 24.1 | 25.25 | 28.975 | 31.1 | 32.25
27.6 | 20.7 | 21.375 | 22.8 | 23.95 | 25.1 | 28.825 | 30.9667 | 32.0833
27.7 | 20.575 | 21.2 | 22.6 | 23.8 | 24.95 | 28.7 | 30.725 | 32
27.8 | 20.4 | 21.1 | 22.4 | 23.7 | 24.8 | 28.5 | 30.675 | 31.825
27.9 | 20.3 | 21 | 22.3 | 23.5 | 24.7 | 28.4 | 30.5 | 31.6833
28 | 20.25 | 20.9 | 22.2 | 23.475 | 24.575 | 28.25 | 30.3 | 31.5
28.1 | 20.1 | 20.8 | 22.075 | 23.325 | 24.4 | 28.1 | 30.2 | 31.3167
28.2 | 20 | 20.675 | 21.925 | 23.1 | 24.275 | 27.95 | 30 | 31.2
28.3 | 19.875 | 20.525 | 21.88 | 23 | 24.125 | 27.825 | 29.9 | 31.08
28.4 | 19.75 | 20.475 | 21.75 | 22.9 | 24 | 27.6667 | 29.7 | 31
28.5 | 19.7 | 20.3 | 21.6 | 22.76 | 23.85 | 27.52 | 29.6 | 30.74
28.6 | 19.55 | 20.15 | 21.4 | 22.625 | 23.7 | 27.375 | 29.475 | 30.625
28.7 | 19.5 | 20 | 21.4 | 22.575 | 23.6 | 27.3 | 29.3 | 30.5
28.8 | 19.35 | 19.9 | 21.275 | 22.4 | 23.5 | 27.1 | 29.1 | 30.3
28.9 | 19.3 | 19.8 | 21.1 | 22.3 | 23.5 | 27 | 29 | 30.175
29 | 19.1 | 19.7 | 21 | 22.2 | 23.2667 | 26.8 | 28.9 | 30.1
29.1 | 19 | 19.6 | 20.9 | 22.025 | 23.08 | 26.7 | 28.7667 | 29.8667
29.2 | 18.9 | 19.4 | 20.8 | 21.9 | 23 | 26.5667 | 28.6 | 29.7833
29.3 | 18.8 | 19.35 | 20.6 | 21.8 | 22.8 | 26.475 | 28.5 | 29.6167
29.4 | 18.65 | 19.3 | 20.5 | 21.7 | 22.7 | 26.325 | 28.3333 | 29.55
29.5 | 18.5 | 19.2 | 20.4 | 21.6 | 22.6 | 26.1 | 28.1333 | 29.35
29.6 | 18.4667 | 19.0667 | 20.3 | 21.45 | 22.5 | 26.05 | 28 | 29.2333
29.7 | 18.3 | 19 | 20.1333 | 21.3 | 22.325 | 25.9 | 27.9333 | 29.1
29.8 | 18.2 | 18.875 | 20.1 | 21.2 | 22.2 | 25.8 | 27.8 | 28.9667
29.9 | 18.1 | 18.74 | 19.95 | 21.1 | 22.1 | 25.65 | 27.6667 | 28.8143
30 | 18 | 18.58 | 19.9 | 21 | 21.98 | 25.52 | 27.5 | 28.7

## Open Field Profiles at 1.5cm
Lateral Profiles
Open field
SSD=100cm
Depth=1.5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.3 | [7 empty] | 4.5
-25.2 | [7 empty] | 4.68
-25.1 | [7 empty] | 4.7
-25 | [7 empty] | 4.8
-24.9 | [7 empty] | 4.9
-24.8 | [7 empty] | 4.9833
-24.7 | [7 empty] | 5.1
-24.6 | [7 empty] | 5.1833
-24.5 | [7 empty] | 5.26
-24.4 | [7 empty] | 5.4
-24.3 | [7 empty] | 5.5
-24.2 | [7 empty] | 5.6
-24.1 | [7 empty] | 5.7667
-24 | [7 empty] | 5.9
-23.9 | [7 empty] | 6
-23.8 | [7 empty] | 6.16
-23.7 | [7 empty] | 6.25
-23.6 | [7 empty] | 6.34
-23.5 | [7 empty] | 6.525
-23.4 | [7 empty] | 6.65
-23.3 | [7 empty] | 6.8429
-23.2 | [7 empty] | 7
-23.1 | [7 empty] | 7.1
-23 | [7 empty] | 7.2667
-22.9 | [7 empty] | 7.4333
-22.8 | [7 empty] | 7.62
-22.7 | [7 empty] | 7.8167
-22.6 | [7 empty] | 8.0167
-22.5 | [7 empty] | 8.1714
-22.4 | [7 empty] | 8.38
-22.3 | [7 empty] | 8.56
-22.2 | [7 empty] | 8.76
-22.1 | [7 empty] | 8.95
-22 | [7 empty] | 9.1333
-21.9 | [7 empty] | 9.3
-21.8 | [7 empty] | 9.5333
-21.7 | [7 empty] | 9.7333
-21.6 | [7 empty] | 9.9667
-21.5 | [7 empty] | 10.2333
-21.4 | [7 empty] | 10.55
-21.3 | [7 empty] | 10.9667
-21.2 | [7 empty] | 11.45
-21.1 | [7 empty] | 12.1
-21 | [7 empty] | 12.925
-20.9 | [7 empty] | 14.3
-20.8 | [7 empty] | 16.4
-20.7 | [7 empty] | 20.58
-20.6 | [7 empty] | 28.4
-20.5 | [7 empty] | 38.2857
-20.4 | [7 empty] | 49.82
-20.3 | [7 empty] | 62.4429
-20.2 | [7 empty] | 74.3
-20.1 | [6 empty] | 4.3 | 84.12
-20 | [6 empty] | 4.4 | 92.45
-19.9 | [6 empty] | 4.5 | 97.8
-19.8 | [6 empty] | 4.6 | 100.3
-19.7 | [6 empty] | 4.6667 | 101.7
-19.6 | [6 empty] | 4.76 | 102.5
-19.5 | [6 empty] | 4.8 | 103
-19.4 | [6 empty] | 4.9 | 103.3333
-19.3 | [6 empty] | 5 | 103.6667
-19.2 | [6 empty] | 5.1167 | 103.8
-19.1 | [6 empty] | 5.2167 | 103.96
-19 | [6 empty] | 5.3 | 104.1
-18.9 | [6 empty] | 5.4 | 104.2
-18.8 | [6 empty] | 5.5667 | 104.3
-18.7 | [6 empty] | 5.66 | 104.32
-18.6 | [6 empty] | 5.78 | 104.5
-18.5 | [6 empty] | 5.9 | 104.5143
-18.4 | [6 empty] | 6 | 104.6
-18.3 | [6 empty] | 6.1 | 104.6
-18.2 | [6 empty] | 6.3 | 104.6
-18.1 | [6 empty] | 6.4 | 104.6
-18 | [6 empty] | 6.5 | 104.7
-17.9 | [6 empty] | 6.625 | 104.7
-17.8 | [6 empty] | 6.8 | 104.8
-17.7 | [6 empty] | 6.96 | 104.8
-17.6 | [6 empty] | 7.075 | 104.8
-17.5 | [6 empty] | 7.225 | 104.8
-17.4 | [6 empty] | 7.4333 | 104.85
-17.3 | [6 empty] | 7.54 | 104.9
-17.2 | [6 empty] | 7.76 | 104.9
-17.1 | [6 empty] | 7.9 | 104.9
-17 | [6 empty] | 8.1 | 104.9167
-16.9 | [6 empty] | 8.3 | 105
-16.8 | [6 empty] | 8.45 | 105
-16.7 | [6 empty] | 8.6667 | 105
-16.6 | [6 empty] | 8.88 | 105
-16.5 | [6 empty] | 9.1 | 105
-16.4 | [6 empty] | 9.3333 | 105
-16.3 | [6 empty] | 9.66 | 105
-16.2 | [6 empty] | 10.06 | 105
-16.1 | [6 empty] | 10.55 | 105
-16 | [6 empty] | 11.1 | 105
-15.9 | [6 empty] | 12.125 | 105
-15.8 | [6 empty] | 13.5667 | 105
-15.7 | [6 empty] | 15.92 | 105
-15.6 | [6 empty] | 20.9 | 105
-15.5 | [6 empty] | 28.65 | 105
-15.4 | [6 empty] | 38.6 | 105
-15.3 | [6 empty] | 51 | 105
-15.2 | [6 empty] | 63.4 | 104.9
-15.1 | [5 empty] | 3.15 | 74.7 | 104.9
-15 | [5 empty] | 3.2 | 84.45 | 104.9
-14.9 | [5 empty] | 3.3 | 92.1 | 104.9
-14.8 | [5 empty] | 3.3 | 96.7 | 104.9
-14.7 | [5 empty] | 3.4 | 98.95 | 104.9
-14.6 | [5 empty] | 3.45 | 100.3 | 104.9
-14.5 | [5 empty] | 3.5 | 101.08 | 104.9
-14.4 | [5 empty] | 3.6 | 101.55 | 104.9
-14.3 | [5 empty] | 3.7 | 101.9 | 104.8
-14.2 | [5 empty] | 3.7 | 102.1167 | 104.8
-14.1 | [5 empty] | 3.8 | 102.42 | 104.8
-14 | [5 empty] | 3.9 | 102.52 | 104.8
-13.9 | [5 empty] | 4 | 102.64 | 104.8
-13.8 | [5 empty] | 4.1 | 102.74 | 104.8
-13.7 | [5 empty] | 4.2 | 102.85 | 104.7167
-13.6 | [5 empty] | 4.3 | 102.9 | 104.7
-13.5 | [5 empty] | 4.375 | 103 | 104.7
-13.4 | [5 empty] | 4.5 | 103 | 104.6
-13.3 | [5 empty] | 4.6 | 103.1 | 104.6
-13.2 | [5 empty] | 4.7 | 103.1 | 104.6
-13.1 | [5 empty] | 4.8 | 103.2 | 104.5
-13 | [5 empty] | 4.9 | 103.2 | 104.5
-12.9 | [5 empty] | 5.08 | 103.2 | 104.5
-12.8 | [5 empty] | 5.2 | 103.2 | 104.5
-12.7 | [5 empty] | 5.28 | 103.2 | 104.5
-12.6 | [5 empty] | 5.4 | 103.3 | 104.5
-12.5 | [5 empty] | 5.55 | 103.3 | 104.4
-12.4 | [5 empty] | 5.7 | 103.3 | 104.4
-12.3 | [5 empty] | 5.875 | 103.34 | 104.4
-12.2 | [5 empty] | 6 | 103.4 | 104.4
-12.1 | [5 empty] | 6.2 | 103.4 | 104.35
-12 | [5 empty] | 6.35 | 103.4 | 104.3
-11.9 | [5 empty] | 6.5 | 103.5 | 104.3
-11.8 | [5 empty] | 6.7 | 103.5 | 104.3
-11.7 | [5 empty] | 6.85 | 103.5 | 104.2
-11.6 | [5 empty] | 7.08 | 103.5 | 104.2
-11.5 | [5 empty] | 7.3 | 103.5 | 104.2
-11.4 | [5 empty] | 7.5333 | 103.5 | 104.2
-11.3 | [5 empty] | 7.7667 | 103.5 | 104.2
-11.2 | [5 empty] | 8.075 | 103.5 | 104.2
-11.1 | [5 empty] | 8.45 | 103.5 | 104.12
-11 | [5 empty] | 8.975 | 103.5 | 104.1
-10.9 | [5 empty] | 9.68 | 103.5 | 104.1
-10.8 | [5 empty] | 10.6 | 103.5 | 104.1
-10.7 | [5 empty] | 12.2 | 103.5 | 104
-10.6 | [5 empty] | 15.0333 | 103.5 | 104
-10.5 | [5 empty] | 20.1 | 103.5 | 104
-10.4 | [5 empty] | 28.2667 | 103.5 | 104
-10.3 | [5 empty] | 39.1667 | 103.5 | 103.9
-10.2 | [5 empty] | 51.2 | 103.5 | 103.8
-10.1 | [5 empty] | 62.475 | 103.5 | 103.8
-10 |  |  |  |  | 1.7 | 73.9667 | 103.5 | 103.8
-9.9 |  |  |  |  | 1.7 | 84.475 | 103.5 | 103.8
-9.8 |  |  |  |  | 1.8 | 91.5 | 103.5 | 103.8
-9.7 |  |  |  |  | 1.8 | 95.65 | 103.5 | 103.8
-9.6 |  |  |  |  | 1.8 | 97.7 | 103.4 | 103.7
-9.5 |  |  |  |  | 1.9 | 98.9 | 103.4 | 103.7
-9.4 |  |  |  |  | 1.9 | 99.65 | 103.4 | 103.7
-9.3 |  |  |  |  | 1.95 | 100.25 | 103.4 | 103.7
-9.2 |  |  |  |  | 2 | 100.5333 | 103.4 | 103.6
-9.1 |  |  |  |  | 2.0333 | 100.9 | 103.4 | 103.6
-9 |  |  |  | 1.3 | 2.1 | 101.1333 | 103.4 | 103.6
-8.9 |  |  |  | 1.4 | 2.1 | 101.3 | 103.4 | 103.5
-8.8 |  |  |  | 1.4 | 2.2 | 101.5 | 103.4 | 103.5
-8.7 |  |  |  | 1.4 | 2.2 | 101.6 | 103.4 | 103.5
-8.6 |  |  |  | 1.45 | 2.3 | 101.7 | 103.4 | 103.5
-8.5 |  |  |  | 1.5 | 2.35 | 101.75 | 103.34 | 103.4
-8.4 |  |  |  | 1.5 | 2.4 | 101.825 | 103.3 | 103.4
-8.3 |  |  |  | 1.5 | 2.5 | 101.9 | 103.3 | 103.4
-8.2 |  |  |  | 1.6 | 2.525 | 102.1 | 103.3 | 103.3
-8.1 |  |  |  | 1.6 | 2.6 | 102.1 | 103.3 | 103.3
-8 |  |  | 1 | 1.6 | 2.7 | 102.2 | 103.3 | 103.3
-7.9 |  |  | 1 | 1.7 | 2.8 | 102.2 | 103.2 | 103.3
-7.8 |  |  | 1 | 1.7 | 2.9 | 102.3 | 103.2 | 103.3
-7.7 |  |  | 1 | 1.8 | 2.9 | 102.3 | 103.2 | 103.3
-7.6 |  |  | 1.1 | 1.8 | 3.0333 | 102.3 | 103.2 | 103.2
-7.5 |  |  | 1.1 | 1.9 | 3.1667 | 102.3 | 103.2 | 103.2
-7.4 |  |  | 1.1 | 1.9 | 3.2 | 102.3 | 103.2 | 103.1
-7.3 |  |  | 1.1 | 1.975 | 3.4 | 102.4 | 103.1 | 103.1
-7.2 |  |  | 1.1 | 2 | 3.48 | 102.4 | 103.1 | 103.1
-7.1 |  |  | 1.2 | 2.1 | 3.6 | 102.4 | 103.1 | 103
-7 |  | 0.7 | 1.2 | 2.1 | 3.7 | 102.4 | 103.1 | 103
-6.9 |  | 0.7 | 1.2 | 2.2 | 3.9 | 102.4 | 103.1 | 103
-6.8 |  | 0.7 | 1.3 | 2.3 | 4 | 102.4 | 103.04 | 103
-6.7 |  | 0.7 | 1.3 | 2.4 | 4.1333 | 102.5 | 103 | 102.9
-6.6 |  | 0.7 | 1.3 | 2.4 | 4.325 | 102.5 | 103 | 102.86
-6.5 | 0.5 | 0.7 | 1.4 | 2.5 | 4.45 | 102.5 | 103 | 102.8
-6.4 | 0.5 | 0.7 | 1.4 | 2.6 | 4.65 | 102.5 | 102.9 | 102.8
-6.3 | 0.5 | 0.8 | 1.5 | 2.7 | 4.9 | 102.5 | 102.9 | 102.8
-6.2 | 0.55 | 0.8 | 1.5 | 2.8 | 5.05 | 102.5 | 102.9 | 102.8
-6.1 | 0.6 | 0.8 | 1.55 | 2.9 | 5.4333 | 102.5 | 102.8857 | 102.7
-6 | 0.6 | 0.8 | 1.6 | 3.025 | 5.8 | 102.4 | 102.8 | 102.7
-5.9 | 0.6 | 0.8 | 1.7 | 3.1 | 6.3 | 102.4 | 102.8 | 102.7
-5.8 | 0.6 | 0.9 | 1.7 | 3.2667 | 6.9333 | 102.4 | 102.8 | 102.6
-5.7 | 0.6 | 0.9 | 1.8 | 3.4 | 8.0333 | 102.4 | 102.78 | 102.6
-5.6 | 0.6 | 0.9 | 1.9 | 3.575 | 9.6667 | 102.4 | 102.7 | 102.5333
-5.5 | 0.6 | 0.9 | 1.9 | 3.7 | 12.4333 | 102.4 | 102.7 | 102.5
-5.4 | 0.6 | 0.9 | 1.95 | 3.8333 | 18.1 | 102.4 | 102.68 | 102.5
-5.3 | 0.7 | 0.93333 | 2 | 4.04 | 26.6667 | 102.4 | 102.6 | 102.4
-5.2 | 0.7 | 1 | 2.1 | 4.3 | 37.675 | 102.4 | 102.5857 | 102.4
-5.1 | 0.7 | 1 | 2.2 | 4.6 | 49.9 | 102.3 | 102.5 | 102.34
-5 | 0.7 | 1.0333 | 2.3 | 4.8667 | 61.4 | 102.3 | 102.5 | 102.3
-4.9 | 0.7 | 1.1 | 2.4 | 5.34 | 73.5333 | 102.2 | 102.45 | 102.3
-4.8 | 0.7 | 1.1 | 2.5 | 5.925 | 82.9 | 102.2 | 102.36 | 102.2
-4.7 | 0.7 | 1.1 | 2.6 | 6.8 | 90.6 | 102.1 | 102.3 | 102.12
-4.6 | 0.8 | 1.2 | 2.75 | 8.3 | 94.7667 | 102.05 | 102.25 | 102.04
-4.5 | 0.8 | 1.2 | 2.8 | 10.7667 | 96.7333 | 102 | 102.2 | 101.9333
-4.4 | 0.8 | 1.3 | 3 | 15.4 | 98 | 101.96 | 102.06 | 101.8
-4.3 | 0.8 | 1.3 | 3.1333 | 22.7 | 98.65 | 101.9 | 102 | 101.8
-4.2 | 0.9 | 1.4 | 3.3 | 33 | 99.1 | 101.8 | 101.9 | 101.7
-4.1 | 0.9 | 1.4 | 3.5 | 44.8 | 99.45 | 101.7 | 101.82 | 101.6
-4 | 0.9 | 1.5 | 3.8 | 57.8667 | 99.7 | 101.6 | 101.75 | 101.5
-3.9 | 0.9 | 1.6 | 4.2 | 69.5 | 99.7 | 101.55 | 101.65 | 101.4
-3.8 | 1 | 1.6667 | 4.7 | 80.475 | 99.9 | 101.5 | 101.56 | 101.3167
-3.7 | 1 | 1.7 | 5.6 | 88.5 | 100 | 101.475 | 101.4667 | 101.3
-3.6 | 1.1 | 1.8 | 6.9 | 93.2 | 100.1 | 101.4 | 101.4 | 101.2
-3.5 | 1.1 | 1.9 | 8.9 | 95.7 | 100.2 | 101.3667 | 101.3 | 101.1
-3.4 | 1.2 | 2 | 12.9 | 97.3 | 100.2 | 101.3 | 101.3 | 101.1
-3.3 | 1.2 | 2.1 | 20 | 98.15 | 100.3 | 101.225 | 101.22 | 101.1
-3.2 | 1.3 | 2.25 | 29.6 | 98.7 | 100.3 | 101.2 | 101.2 | 101.1
-3.1 | 1.3 | 2.4 | 42.1 | 99.02 | 100.3667 | 101.2 | 101.2 | 101
-3 | 1.4 | 2.7 | 54.7 | 99.3 | 100.4 | 101.2 | 101.1 | 101
-2.9 | 1.5 | 3.05 | 67 | 99.475 | 100.4 | 101.1 | 101.1 | 101
-2.8 | 1.6 | 3.5 | 78 | 99.7 | 100.5 | 101.1 | 101.1 | 101
-2.7 | 1.75 | 4.2 | 86.75 | 99.9 | 100.6 | 101.1 | 101.1 | 100.9333
-2.6 | 1.85 | 5.35 | 92.6333 | 100 | 100.6 | 101.1 | 101.1 | 100.9
-2.5 | 2.05 | 7.3667 | 95.9 | 100.1 | 100.65 | 101.1 | 101.1 | 100.9
-2.4 | 2.35 | 11 | 97.4 | 100.2 | 100.7 | 101.1 | 101 | 100.9
-2.3 | 2.7667 | 17.5 | 98.3 | 100.3 | 100.7 | 101 | 101 | 100.8
-2.2 | 3.4 | 27.05 | 98.8 | 100.34 | 100.7 | 101 | 101 | 100.8
-2.1 | 4.6 | 38.9 | 99.3 | 100.4 | 100.7 | 100.9 | 101 | 100.8
-2 | 6.3 | 52 | 99.6 | 100.4 | 100.7 | 100.9 | 100.9 | 100.7
-1.9 | 9.9 | 64.7 | 99.7 | 100.5 | 100.7 | 100.9 | 100.85 | 100.7
-1.8 | 16 | 76.3 | 99.9 | 100.48 | 100.7 | 100.8 | 100.8 | 100.6333
-1.7 | 25.6 | 86.4 | 100 | 100.4 | 100.7 | 100.8 | 100.7 | 100.6
-1.6 | 36.5 | 92.3 | 100 | 100.4 | 100.7 | 100.7 | 100.7 | 100.6
-1.5 | 50.1 | 95.9 | 100.1 | 100.4 | 100.6 | 100.7 | 100.6 | 100.5
-1.4 | 63.1 | 97.65 | 100.1 | 100.4 | 100.575 | 100.6 | 100.6 | 100.46
-1.3 | 74.6 | 98.55 | 100.2 | 100.3 | 100.5 | 100.6 | 100.5 | 100.4
-1.2 | 84.7 | 99.1 | 100.2 | 100.3 | 100.4 | 100.5 | 100.5 | 100.35
-1.1 | 92 | 99.5 | 100.2 | 100.3 | 100.4 | 100.45 | 100.4 | 100.3
-1 | 95.7 | 99.7 | 100.2 | 100.3 | 100.3 | 100.4 | 100.3 | 100.3
-0.9 | 97.7 | 99.8 | 100.2 | 100.3 | 100.3 | 100.3 | 100.3 | 100.2
-0.8 | 98.7333 | 99.9 | 100.2 | 100.3 | 100.3 | 100.3 | 100.3 | 100.1833
-0.7 | 99.1 | 100 | 100.1 | 100.3 | 100.2 | 100.3 | 100.2 | 100.1
-0.6 | 99.5 | 100 | 100.1 | 100.2 | 100.2 | 100.2 | 100.15 | 100.05
-0.5 | 99.8 | 100 | 100.1 | 100.2 | 100.1 | 100.2 | 100.1 | 100
-0.4 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100.2 | 100.1 | 100
-0.3 | 99.9 | 100 | 100.1 | 100.1 | 100 | 100.1 | 100 | 100
-0.2 | 100 | 100 | 100 | 100.075 | 100 | 100.1 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100.0333 | 100 | 100 | 100 | 100
0.3 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100 | 100 | 100
0.4 | 99.8 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.1 | 100
0.5 | 99.7 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.1 | 100
0.6 | 99.5 | 99.9 | 100 | 100.1667 | 100.1 | 100.175 | 100.1 | 100.0167
0.7 | 99.1 | 99.9 | 100.1 | 100.2 | 100.2 | 100.2 | 100.2 | 100.1
0.8 | 98.8 | 99.9 | 100.1 | 100.2 | 100.2 | 100.2 | 100.3 | 100.2
0.9 | 97.5 | 99.8 | 100.1 | 100.2 | 100.2 | 100.3 | 100.3 | 100.2
1 | 95.5 | 99.6 | 100.1 | 100.2 | 100.2 | 100.3333 | 100.32 | 100.2833
1.1 | 91.9 | 99.3333 | 100.05 | 100.2 | 100.3 | 100.4 | 100.4 | 100.3
1.2 | 84.9 | 99.1 | 100 | 100.2 | 100.3 | 100.5 | 100.45 | 100.3167
1.3 | 74.6 | 98.4 | 100 | 100.2 | 100.3333 | 100.5 | 100.5 | 100.4
1.4 | 63.8 | 97.4 | 100 | 100.3 | 100.4 | 100.5 | 100.6 | 100.5
1.5 | 50.15 | 95.65 | 100 | 100.3 | 100.5 | 100.6 | 100.6 | 100.5143
1.6 | 36.3 | 92.35 | 99.9 | 100.3 | 100.5 | 100.6 | 100.7 | 100.6
1.7 | 25 | 86.4 | 99.8 | 100.3 | 100.5 | 100.7 | 100.75 | 100.6
1.8 | 15.55 | 76.2 | 99.7 | 100.3 | 100.5 | 100.7 | 100.8 | 100.6
1.9 | 9.5333 | 64.5 | 99.6 | 100.3 | 100.6 | 100.7 | 100.88 | 100.7
2 | 6.2 | 51.8 | 99.45 | 100.25 | 100.6 | 100.8 | 100.9 | 100.7
2.1 | 4.4667 | 38.8333 | 99.1 | 100.2 | 100.6 | 100.8 | 100.925 | 100.8
2.2 | 3.4 | 26.8 | 98.7 | 100.2 | 100.6 | 100.8 | 101 | 100.8
2.3 | 2.8 | 17.1 | 98.15 | 100.1 | 100.6 | 100.9 | 101 | 100.8
2.4 | 2.3 | 10.85 | 97.2 | 100 | 100.6 | 100.9 | 101 | 100.9
2.5 | 2.05 | 7.2333 | 95.5 | 99.9 | 100.5 | 100.9 | 101.1 | 100.9
2.6 | 1.85 | 5.3 | 92.5 | 99.8 | 100.4 | 100.9 | 101.1 | 100.9
2.7 | 1.7 | 4.1 | 86.7 | 99.7 | 100.4 | 100.9 | 101.1 | 100.9
2.8 | 1.6 | 3.5 | 78.2 | 99.46 | 100.3 | 100.9 | 101.1 | 101
2.9 | 1.5 | 3 | 67.05 | 99.375 | 100.28 | 100.9 | 101.1 | 101
3 | 1.4 | 2.6 | 54.0667 | 99.12 | 100.2 | 100.9 | 101.1 | 101
3.1 | 1.3 | 2.45 | 41.8 | 98.7667 | 100.18 | 100.9 | 101.2 | 101
3.2 | 1.3 | 2.3 | 30.3 | 98.4333 | 100.1 | 100.9 | 101.2 | 101.1
3.3 | 1.2 | 2.1 | 20.4 | 97.9333 | 100 | 101 | 101.25 | 101.1
3.4 | 1.2 | 2 | 12.85 | 97.1 | 99.925 | 101 | 101.3 | 101.1
3.5 | 1.1 | 1.9 | 8.8 | 95.6 | 99.9 | 101.025 | 101.32 | 101.1667
3.6 | 1.1 | 1.8 | 6.8 | 93.1 | 99.8 | 101.1 | 101.4 | 101.2
3.7 | 1 | 1.7 | 5.55 | 88.5 | 99.7 | 101.2 | 101.5 | 101.3
3.8 | 1 | 1.6 | 4.75 | 79.8 | 99.65 | 101.22 | 101.5 | 101.3833
3.9 | 0.93333 | 1.5 | 4.2 | 69.0667 | 99.5 | 101.3 | 101.6 | 101.4
4 | 0.9 | 1.5 | 3.7667 | 57.225 | 99.375 | 101.4 | 101.7 | 101.5
4.1 | 0.9 | 1.4 | 3.5 | 44.8 | 99.125 | 101.5 | 101.8 | 101.6
4.2 | 0.85 | 1.4 | 3.3 | 32.825 | 98.9 | 101.5667 | 101.9 | 101.68
4.3 | 0.8 | 1.3 | 3.1 | 22.6 | 98.425 | 101.6 | 102 | 101.7667
4.4 | 0.8 | 1.3 | 3 | 15.1 | 97.6 | 101.7 | 102.1 | 101.8333
4.5 | 0.8 | 1.2 | 2.8 | 10.65 | 96.45 | 101.7 | 102.2 | 101.9
4.6 | 0.8 | 1.2 | 2.7 | 8.2 | 94.3667 | 101.775 | 102.2 | 102.0167
4.7 | 0.7 | 1.1 | 2.6 | 6.8 | 90.7 | 101.8 | 102.3 | 102.1
4.8 | 0.7 | 1.1 | 2.4333 | 5.9 | 83.4 | 101.9 | 102.4 | 102.2
4.9 | 0.7 | 1.1 | 2.4 | 5.25 | 73.5 | 101.9 | 102.4 | 102.2667
5 | 0.7 | 1 | 2.3 | 4.8 | 62.1333 | 101.9333 | 102.5 | 102.3
5.1 | 0.7 | 1 | 2.2 | 4.5 | 49.6333 | 102 | 102.5 | 102.3
5.2 | 0.6 | 1 | 2.1 | 4.2333 | 37.6667 | 102.075 | 102.6 | 102.4
5.3 | 0.6 | 0.9 | 2 | 4 | 26.8333 | 102.1 | 102.6 | 102.44
5.4 | 0.6 | 0.9 | 1.9 | 3.8 | 17.8 | 102.1 | 102.62 | 102.5
5.5 | 0.6 | 0.9 | 1.9 | 3.7 | 12.42 | 102.1 | 102.7 | 102.5
5.6 | 0.6 | 0.9 | 1.8 | 3.5 | 9.5 | 102.1 | 102.7 | 102.5
5.7 | 0.6 | 0.8 | 1.8 | 3.4 | 7.9 | 102.1 | 102.8 | 102.6
5.8 | 0.6 | 0.8 | 1.7 | 3.2333 | 6.85 | 102.2 | 102.8 | 102.6
5.9 | 0.6 | 0.8 | 1.7 | 3.1333 | 6.1667 | 102.2 | 102.8 | 102.7
6 | 0.6 | 0.8 | 1.6 | 3 | 5.7333 | 102.2 | 102.8 | 102.7
6.1 | 0.6 | 0.8 | 1.5 | 2.9 | 5.4 | 102.2 | 102.8286 | 102.75
6.2 | 0.6 | 0.8 | 1.5 | 2.8 | 5.05 | 102.2 | 102.9 | 102.8
6.3 | 0.5 | 0.8 | 1.5 | 2.7 | 4.8333 | 102.14 | 102.9 | 102.8
6.4 | 0.5 | 0.7 | 1.4 | 2.6 | 4.6 | 102.1 | 102.9 | 102.8
6.5 | 0.5 | 0.7 | 1.4 | 2.5 | 4.45 | 102.1 | 103 | 102.8
6.6 |  | 0.7 | 1.3 | 2.4 | 4.26 | 102.1 | 103 | 102.825
6.7 |  | 0.7 | 1.3 | 2.325 | 4.1 | 102.1 | 103 | 102.9167
6.8 |  | 0.7 | 1.3 | 2.3 | 3.975 | 102.1 | 103.08 | 103
6.9 |  | 0.7 | 1.2 | 2.2 | 3.8 | 102.1 | 103.1 | 103
7 |  | 0.7 | 1.2 | 2.1 | 3.7 | 102.1 | 103.1 | 103
7.1 |  |  | 1.2 | 2.1 | 3.5333 | 102.1 | 103.1 | 103.06
7.2 |  |  | 1.15 | 2 | 3.4667 | 102.1 | 103.1 | 103.1
7.3 |  |  | 1.1 | 1.925 | 3.3 | 102 | 103.1 | 103.1
7.4 |  |  | 1.1 | 1.9 | 3.2 | 102 | 103.16 | 103.1
7.5 |  |  | 1.1 | 1.85 | 3.1 | 102 | 103.2 | 103.2
7.6 |  |  | 1.05 | 1.8 | 3 | 102 | 103.2 | 103.2
7.7 |  |  | 1 | 1.8 | 2.9 | 101.9 | 103.2 | 103.3
7.8 |  |  | 1 | 1.7 | 2.8 | 101.9 | 103.2 | 103.3
7.9 |  |  | 1 | 1.7 | 2.8 | 101.9 | 103.26 | 103.3
8 |  |  | 1 | 1.6 | 2.7 | 101.8 | 103.3 | 103.3
8.1 |  |  |  | 1.6 | 2.6 | 101.8 | 103.3 | 103.3
8.2 |  |  |  | 1.6 | 2.5 | 101.7 | 103.3 | 103.3
8.3 |  |  |  | 1.5 | 2.5 | 101.7 | 103.3 | 103.4
8.4 |  |  |  | 1.5 | 2.4 | 101.6 | 103.3 | 103.4
8.5 |  |  |  | 1.5 | 2.4 | 101.4333 | 103.38 | 103.4
8.6 |  |  |  | 1.425 | 2.3 | 101.325 | 103.4 | 103.5
8.7 |  |  |  | 1.4 | 2.2 | 101.3 | 103.4 | 103.5
8.8 |  |  |  | 1.4 | 2.175 | 101.125 | 103.4 | 103.5
8.9 |  |  |  | 1.4 | 2.1 | 100.9 | 103.4 | 103.5
9 |  |  |  | 1.3 | 2.1 | 100.8 | 103.4 | 103.6
9.1 |  |  |  |  | 2 | 100.55 | 103.4 | 103.6
9.2 |  |  |  |  | 2 | 100.3 | 103.4 | 103.64
9.3 |  |  |  |  | 1.9 | 99.925 | 103.4 | 103.7
9.4 |  |  |  |  | 1.9 | 99.5 | 103.4 | 103.7
9.5 |  |  |  |  | 1.9 | 98.6333 | 103.4 | 103.7
9.6 |  |  |  |  | 1.8 | 97.4 | 103.4 | 103.7
9.7 |  |  |  |  | 1.8 | 95.2 | 103.45 | 103.8
9.8 |  |  |  |  | 1.7 | 91.2 | 103.5 | 103.8
9.9 |  |  |  |  | 1.7 | 83.6 | 103.5 | 103.8
10 |  |  |  |  | 1.7 | 74.4 | 103.5 | 103.8
10.1 | [5 empty] | 62.8 | 103.5 | 103.8
10.2 | [5 empty] | 50.725 | 103.5 | 103.82
10.3 | [5 empty] | 39.5 | 103.5 | 103.9
10.4 | [5 empty] | 28.45 | 103.5 | 104
10.5 | [5 empty] | 19.775 | 103.5 | 104
10.6 | [5 empty] | 14.9 | 103.5 | 104
10.7 | [5 empty] | 12.2 | 103.5 | 104.06
10.8 | [5 empty] | 10.65 | 103.5 | 104.1
10.9 | [5 empty] | 9.6 | 103.5 | 104.1
11 | [5 empty] | 8.95 | 103.5 | 104.1
11.1 | [5 empty] | 8.4333 | 103.5 | 104.1
11.2 | [5 empty] | 8.05 | 103.5 | 104.2
11.3 | [5 empty] | 7.8 | 103.5 | 104.2
11.4 | [5 empty] | 7.5 | 103.5 | 104.2
11.5 | [5 empty] | 7.2333 | 103.5 | 104.2
11.6 | [5 empty] | 7 | 103.5 | 104.2
11.7 | [5 empty] | 6.825 | 103.5 | 104.2333
11.8 | [5 empty] | 6.6667 | 103.5 | 104.3
11.9 | [5 empty] | 6.4333 | 103.46 | 104.3
12 | [5 empty] | 6.3 | 103.4 | 104.3
12.1 | [5 empty] | 6.1667 | 103.4 | 104.4
12.2 | [5 empty] | 6 | 103.4 | 104.4
12.3 | [5 empty] | 5.8 | 103.3 | 104.4
12.4 | [5 empty] | 5.7 | 103.3 | 104.4
12.5 | [5 empty] | 5.575 | 103.3 | 104.4167
12.6 | [5 empty] | 5.4 | 103.3 | 104.5
12.7 | [5 empty] | 5.2333 | 103.2 | 104.5
12.8 | [5 empty] | 5.1 | 103.2 | 104.5
12.9 | [5 empty] | 5 | 103.2 | 104.5
13 | [5 empty] | 4.9 | 103.2 | 104.5
13.1 | [5 empty] | 4.7667 | 103.2 | 104.5
13.2 | [5 empty] | 4.64 | 103.1 | 104.6
13.3 | [5 empty] | 4.5 | 103.1 | 104.6
13.4 | [5 empty] | 4.4 | 103.0333 | 104.6167
13.5 | [5 empty] | 4.3 | 103 | 104.7
13.6 | [5 empty] | 4.2333 | 102.9 | 104.7
13.7 | [5 empty] | 4.1667 | 102.8 | 104.7
13.8 | [5 empty] | 4.1 | 102.7 | 104.8
13.9 | [5 empty] | 4 | 102.7 | 104.8
14 | [5 empty] | 3.9 | 102.5667 | 104.8
14.1 | [5 empty] | 3.8 | 102.38 | 104.8
14.2 | [5 empty] | 3.7 | 102.0667 | 104.8
14.3 | [5 empty] | 3.65 | 101.92 | 104.8667
14.4 | [5 empty] | 3.6 | 101.55 | 104.9
14.5 | [5 empty] | 3.5 | 101.075 | 104.9
14.6 | [5 empty] | 3.425 | 100.28 | 104.9
14.7 | [5 empty] | 3.4 | 99.1 | 104.9
14.8 | [5 empty] | 3.3 | 96.675 | 104.9
14.9 | [5 empty] | 3.2333 | 92.12 | 104.9
15 | [5 empty] | 3.2 | 84.825 | 104.9
15.1 | [5 empty] | 3.1 | 74.5 | 104.9
15.2 | [6 empty] | 62.8 | 104.9
15.3 | [6 empty] | 50.7 | 105
15.4 | [6 empty] | 39.125 | 105
15.5 | [6 empty] | 29.025 | 105
15.6 | [6 empty] | 20.9 | 105
15.7 | [6 empty] | 16.1 | 105
15.8 | [6 empty] | 13.5 | 105
15.9 | [6 empty] | 12.075 | 105
16 | [6 empty] | 11.18 | 105
16.1 | [6 empty] | 10.52 | 105
16.2 | [6 empty] | 10.05 | 105
16.3 | [6 empty] | 9.6667 | 105
16.4 | [6 empty] | 9.3 | 105
16.5 | [6 empty] | 9.05 | 105
16.6 | [6 empty] | 8.85 | 105
16.7 | [6 empty] | 8.62 | 105
16.8 | [6 empty] | 8.5 | 105
16.9 | [6 empty] | 8.275 | 105
17 | [6 empty] | 8.06 | 104.94
17.1 | [6 empty] | 7.86 | 104.9
17.2 | [6 empty] | 7.725 | 104.9
17.3 | [6 empty] | 7.6 | 104.9
17.4 | [6 empty] | 7.4 | 104.8833
17.5 | [6 empty] | 7.2 | 104.8
17.6 | [6 empty] | 7.1 | 104.8
17.7 | [6 empty] | 6.925 | 104.8
17.8 | [6 empty] | 6.8 | 104.8
17.9 | [6 empty] | 6.6 | 104.7
18 | [6 empty] | 6.5 | 104.6333
18.1 | [6 empty] | 6.36 | 104.6
18.2 | [6 empty] | 6.25 | 104.6
18.3 | [6 empty] | 6.1 | 104.6
18.4 | [6 empty] | 6 | 104.6
18.5 | [6 empty] | 5.9 | 104.5
18.6 | [6 empty] | 5.75 | 104.45
18.7 | [6 empty] | 5.62 | 104.35
18.8 | [6 empty] | 5.5333 | 104.28
18.9 | [6 empty] | 5.4 | 104.16
19 | [6 empty] | 5.35 | 104.0833
19.1 | [6 empty] | 5.2 | 104
19.2 | [6 empty] | 5.1 | 103.8286
19.3 | [6 empty] | 5 | 103.6667
19.4 | [6 empty] | 4.9 | 103.3
19.5 | [6 empty] | 4.82 | 103
19.6 | [6 empty] | 4.72 | 102.5333
19.7 | [6 empty] | 4.62 | 101.7
19.8 | [6 empty] | 4.6 | 100.2667
19.9 | [6 empty] | 4.5 | 97.65
20 | [6 empty] | 4.4 | 92.55
20.1 | [6 empty] | 4.3 | 84.8
20.2 | [7 empty] | 73.6
20.3 | [7 empty] | 62.0714
20.4 | [7 empty] | 50.04
20.5 | [7 empty] | 38.1143
20.6 | [7 empty] | 28.12
20.7 | [7 empty] | 20.82
20.8 | [7 empty] | 16.5
20.9 | [7 empty] | 14.18
21 | [7 empty] | 13
21.1 | [7 empty] | 12.075
21.2 | [7 empty] | 11.42
21.3 | [7 empty] | 10.94
21.4 | [7 empty] | 10.58
21.5 | [7 empty] | 10.25
21.6 | [7 empty] | 9.96
21.7 | [7 empty] | 9.7667
21.8 | [7 empty] | 9.5
21.9 | [7 empty] | 9.3333
22 | [7 empty] | 9.0667
22.1 | [7 empty] | 8.8833
22.2 | [7 empty] | 8.7
22.3 | [7 empty] | 8.6
22.4 | [7 empty] | 8.3333
22.5 | [7 empty] | 8.1167
22.6 | [7 empty] | 7.94
22.7 | [7 empty] | 7.8
22.8 | [7 empty] | 7.65
22.9 | [7 empty] | 7.4667
23 | [7 empty] | 7.3
23.1 | [7 empty] | 7.1333
23.2 | [7 empty] | 7
23.3 | [7 empty] | 6.8333
23.4 | [7 empty] | 6.68
23.5 | [7 empty] | 6.575
23.6 | [7 empty] | 6.375
23.7 | [7 empty] | 6.2
23.8 | [7 empty] | 6.1
23.9 | [7 empty] | 5.98
24 | [7 empty] | 5.875
24.1 | [7 empty] | 5.7
24.2 | [7 empty] | 5.6333
24.3 | [7 empty] | 5.5
24.4 | [7 empty] | 5.3667
24.5 | [7 empty] | 5.28
24.6 | [7 empty] | 5.2
24.7 | [7 empty] | 5.1
24.8 | [7 empty] | 5
24.9 | [7 empty] | 4.9
25 | [7 empty] | 4.8
25.1 | [7 empty] | 4.7
25.2 | [7 empty] | 4.6

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-27.9 | [7 empty] | 4.3
-27.8 | [7 empty] | 4.4
-27.7 | [7 empty] | 4.4571
-27.6 | [7 empty] | 4.5
-27.5 | [7 empty] | 4.5
-27.4 | [7 empty] | 4.6
-27.3 | [7 empty] | 4.65
-27.2 | [7 empty] | 4.7333
-27.1 | [7 empty] | 4.8
-27 | [7 empty] | 4.9
-26.9 | [7 empty] | 4.94
-26.8 | [7 empty] | 5.0167
-26.7 | [7 empty] | 5.1167
-26.6 | [7 empty] | 5.2
-26.5 | [7 empty] | 5.3
-26.4 | [7 empty] | 5.4
-26.3 | [7 empty] | 5.5167
-26.2 | [7 empty] | 5.6
-26.1 | [7 empty] | 5.7
-26 | [7 empty] | 5.8
-25.9 | [7 empty] | 5.9
-25.8 | [7 empty] | 6
-25.7 | [7 empty] | 6.1333
-25.6 | [7 empty] | 6.2
-25.5 | [7 empty] | 6.3667
-25.4 | [7 empty] | 6.5
-25.3 | [7 empty] | 6.6167
-25.2 | [7 empty] | 6.8
-25.1 | [7 empty] | 6.9
-25 | [7 empty] | 7.0167
-24.9 | [7 empty] | 7.1833
-24.8 | [7 empty] | 7.3
-24.7 | [7 empty] | 7.44
-24.6 | [7 empty] | 7.6
-24.5 | [7 empty] | 7.7667
-24.4 | [7 empty] | 7.92
-24.3 | [7 empty] | 8.12
-24.2 | [7 empty] | 8.2857
-24.1 | [7 empty] | 8.46
-24 | [7 empty] | 8.64
-23.9 | [7 empty] | 8.8167
-23.8 | [7 empty] | 9.0667
-23.7 | [7 empty] | 9.25
-23.6 | [7 empty] | 9.44
-23.5 | [7 empty] | 9.7167
-23.4 | [7 empty] | 9.8714
-23.3 | [7 empty] | 10.1333
-23.2 | [7 empty] | 10.3
-23.1 | [7 empty] | 10.6
-23 | [7 empty] | 10.8333
-22.9 | [7 empty] | 11.0571
-22.8 | [7 empty] | 11.3
-22.7 | [7 empty] | 11.55
-22.6 | [7 empty] | 11.84
-22.5 | [7 empty] | 12.1286
-22.4 | [7 empty] | 12.46
-22.3 | [7 empty] | 12.7
-22.2 | [7 empty] | 13.1
-22.1 | [7 empty] | 13.5333
-22 | [7 empty] | 14
-21.9 | [7 empty] | 14.6
-21.8 | [7 empty] | 15.3333
-21.7 | [7 empty] | 16.3333
-21.6 | [7 empty] | 17.62
-21.5 | [7 empty] | 20
-21.4 | [7 empty] | 23.7
-21.3 | [7 empty] | 30.15
-21.2 | [7 empty] | 38.65
-21.1 | [7 empty] | 49.2833
-21 | [7 empty] | 59.8
-20.9 | [7 empty] | 70.2429
-20.8 | [7 empty] | 79.18
-20.7 | [6 empty] | 5.36 | 86.6667
-20.6 | [6 empty] | 5.4429 | 91.8
-20.5 | [6 empty] | 5.52 | 94.6333
-20.4 | [6 empty] | 5.6 | 96.1333
-20.3 | [6 empty] | 5.7 | 97.175
-20.2 | [6 empty] | 5.85 | 97.9667
-20.1 | [6 empty] | 5.92 | 98.55
-20 | [6 empty] | 6.0286 | 99
-19.9 | [6 empty] | 6.1833 | 99.4
-19.8 | [6 empty] | 6.25 | 99.7167
-19.7 | [6 empty] | 6.4167 | 99.9429
-19.6 | [6 empty] | 6.5833 | 100.1429
-19.5 | [6 empty] | 6.66 | 100.3
-19.4 | [6 empty] | 6.8 | 100.5667
-19.3 | [6 empty] | 6.95 | 100.76
-19.2 | [6 empty] | 7.1 | 100.92
-19.1 | [6 empty] | 7.2 | 101.1
-19 | [6 empty] | 7.4 | 101.2
-18.9 | [6 empty] | 7.5 | 101.36
-18.8 | [6 empty] | 7.7333 | 101.5
-18.7 | [6 empty] | 7.9 | 101.6167
-18.6 | [6 empty] | 8.05 | 101.7
-18.5 | [6 empty] | 8.18 | 101.85
-18.4 | [6 empty] | 8.44 | 101.9167
-18.3 | [6 empty] | 8.5167 | 102
-18.2 | [6 empty] | 8.7667 | 102.1
-18.1 | [6 empty] | 8.975 | 102.2
-18 | [6 empty] | 9.16 | 102.3667
-17.9 | [6 empty] | 9.38 | 102.4333
-17.8 | [6 empty] | 9.58 | 102.5
-17.7 | [6 empty] | 9.86 | 102.5
-17.6 | [6 empty] | 10.0667 | 102.6
-17.5 | [6 empty] | 10.3 | 102.6
-17.4 | [6 empty] | 10.6 | 102.75
-17.3 | [6 empty] | 10.7833 | 102.8
-17.2 | [6 empty] | 11.1333 | 102.8
-17.1 | [6 empty] | 11.4333 | 102.9
-17 | [6 empty] | 11.7333 | 102.9
-16.9 | [6 empty] | 12.15 | 103
-16.8 | [6 empty] | 12.58 | 103
-16.7 | [6 empty] | 13.1 | 103.0667
-16.6 | [6 empty] | 13.7 | 103.1
-16.5 | [6 empty] | 14.5286 | 103.1
-16.4 | [6 empty] | 15.7 | 103.2
-16.3 | [6 empty] | 17.5 | 103.2
-16.2 | [6 empty] | 20.4714 | 103.2
-16.1 | [6 empty] | 25.65 | 103.3
-16 | [6 empty] | 33.22 | 103.3
-15.9 | [6 empty] | 42.6143 | 103.3
-15.8 | [6 empty] | 53.05 | 103.34
-15.7 | [6 empty] | 63.7 | 103.4
-15.6 | [6 empty] | 73.3429 | 103.4
-15.5 | [6 empty] | 82.4 | 103.4
-15.4 | [5 empty] | 4 | 88.4 | 103.4833
-15.3 | [5 empty] | 4.1 | 92.2286 | 103.5
-15.2 | [5 empty] | 4.1333 | 94.3714 | 103.5
-15.1 | [5 empty] | 4.22 | 95.7 | 103.5
-15 | [5 empty] | 4.3 | 96.68 | 103.5
-14.9 | [5 empty] | 4.4 | 97.36 | 103.5
-14.8 | [5 empty] | 4.5 | 97.8 | 103.5
-14.7 | [5 empty] | 4.6 | 98.3 | 103.45
-14.6 | [5 empty] | 4.65 | 98.58 | 103.4
-14.5 | [5 empty] | 4.8 | 98.85 | 103.4
-14.4 | [5 empty] | 4.9 | 99.1667 | 103.4
-14.3 | [5 empty] | 5 | 99.3 | 103.4
-14.2 | [5 empty] | 5.1 | 99.4833 | 103.4
-14.1 | [5 empty] | 5.2 | 99.6429 | 103.4
-14 | [5 empty] | 5.32 | 99.8167 | 103.4
-13.9 | [5 empty] | 5.5 | 99.9714 | 103.4
-13.8 | [5 empty] | 5.6 | 100.1333 | 103.4
-13.7 | [5 empty] | 5.7 | 100.3333 | 103.4
-13.6 | [5 empty] | 5.875 | 100.4 | 103.4
-13.5 | [5 empty] | 6 | 100.5667 | 103.4
-13.4 | [5 empty] | 6.2 | 100.65 | 103.4
-13.3 | [5 empty] | 6.3 | 100.72 | 103.4
-13.2 | [5 empty] | 6.45 | 100.82 | 103.4
-13.1 | [5 empty] | 6.65 | 100.94 | 103.4
-13 | [5 empty] | 6.8 | 101 | 103.4
-12.9 | [5 empty] | 6.975 | 101.1 | 103.4
-12.8 | [5 empty] | 7.2 | 101.2 | 103.4
-12.7 | [5 empty] | 7.4 | 101.3 | 103.4
-12.6 | [5 empty] | 7.575 | 101.3 | 103.4
-12.5 | [5 empty] | 7.775 | 101.4 | 103.4
-12.4 | [5 empty] | 8 | 101.4 | 103.4
-12.3 | [5 empty] | 8.275 | 101.5 | 103.4
-12.2 | [5 empty] | 8.45 | 101.6 | 103.4
-12.1 | [5 empty] | 8.8 | 101.7 | 103.4
-12 | [5 empty] | 9 | 101.7 | 103.4
-11.9 | [5 empty] | 9.3333 | 101.75 | 103.4
-11.8 | [5 empty] | 9.6 | 101.8 | 103.32
-11.7 | [5 empty] | 10 | 101.9 | 103.4
-11.6 | [5 empty] | 10.3667 | 101.9 | 103.4
-11.5 | [5 empty] | 10.7667 | 101.9 | 103.4
-11.4 | [5 empty] | 11.4 | 102 | 103.3
-11.3 | [5 empty] | 12.125 | 102 | 103.3
-11.2 | [5 empty] | 13 | 102.0167 | 103.3
-11.1 | [5 empty] | 14.5 | 102.1 | 103.3
-11 | [5 empty] | 16.65 | 102.1 | 103.3
-10.9 | [5 empty] | 20.3 | 102.1 | 103.3
-10.8 | [5 empty] | 26.7333 | 102.1 | 103.2667
-10.7 | [5 empty] | 34.9333 | 102.1 | 103.2
-10.6 | [5 empty] | 45.425 | 102.1333 | 103.2
-10.5 | [5 empty] | 56.2 | 102.2 | 103.2
-10.4 | [5 empty] | 66.8 | 102.2 | 103.2
-10.3 | [5 empty] | 76.4667 | 102.2 | 103.2
-10.2 |  |  |  |  | 2.1 | 84.45 | 102.3 | 103.1286
-10.1 |  |  |  |  | 2.2 | 89.7 | 102.3 | 103.1
-10 |  |  |  |  | 2.2 | 92.7 | 102.3 | 103.1
-9.9 |  |  |  |  | 2.3 | 94.5 | 102.3 | 103.1
-9.8 |  |  |  |  | 2.3 | 95.8 | 102.3 | 103.02
-9.7 |  |  |  |  | 2.4 | 96.4667 | 102.3 | 103
-9.6 |  |  |  |  | 2.4 | 97.1667 | 102.3 | 103
-9.5 |  |  |  |  | 2.5 | 97.6333 | 102.3 | 103
-9.4 |  |  |  |  | 2.5 | 98.025 | 102.3 | 103
-9.3 |  |  |  |  | 2.6 | 98.35 | 102.3 | 103
-9.2 |  |  |  |  | 2.7 | 98.6667 | 102.3 | 103
-9.1 |  |  |  | 1.7 | 2.8 | 98.925 | 102.3 | 103
-9 |  |  |  | 1.7 | 2.8 | 99.1 | 102.3 | 103
-8.9 |  |  |  | 1.8 | 2.9 | 99.35 | 102.4 | 103
-8.8 |  |  |  | 1.8 | 3 | 99.5 | 102.4 | 103
-8.7 |  |  |  | 1.9 | 3.05 | 99.8 | 102.4 | 103
-8.6 |  |  |  | 1.9 | 3.18 | 99.9 | 102.4 | 103
-8.5 |  |  |  | 2 | 3.22 | 100 | 102.4 | 103
-8.4 |  |  |  | 2 | 3.35 | 100.1 | 102.4 | 102.9
-8.3 |  |  |  | 2 | 3.4 | 100.275 | 102.4 | 102.9
-8.2 |  |  |  | 2.1 | 3.525 | 100.425 | 102.4 | 102.9
-8.1 |  |  | 1.3 | 2.2 | 3.625 | 100.5 | 102.4 | 102.9
-8 |  |  | 1.3 | 2.2 | 3.8 | 100.6 | 102.4 | 102.8
-7.9 |  |  | 1.3 | 2.3 | 3.8667 | 100.7 | 102.4 | 102.8
-7.8 |  |  | 1.4 | 2.4 | 4 | 100.8 | 102.4 | 102.8
-7.7 |  |  | 1.4 | 2.4 | 4.1 | 100.9 | 102.4 | 102.8
-7.6 |  |  | 1.4 | 2.5 | 4.25 | 100.95 | 102.4 | 102.75
-7.5 |  |  | 1.5 | 2.6 | 4.42 | 101 | 102.4 | 102.7
-7.4 |  |  | 1.5 | 2.7 | 4.6 | 101.1 | 102.4 | 102.7
-7.3 |  |  | 1.5 | 2.7 | 4.75 | 101.1 | 102.4 | 102.7
-7.2 |  |  | 1.6 | 2.8 | 4.9 | 101.2 | 102.4 | 102.7
-7.1 |  | 0.9 | 1.6 | 2.9 | 5.1 | 101.2 | 102.4 | 102.7
-7 |  | 0.9 | 1.7 | 3 | 5.3 | 101.3 | 102.36 | 102.6167
-6.9 |  | 0.9 | 1.7 | 3.1 | 5.5667 | 101.3 | 102.3 | 102.6
-6.8 |  | 0.9 | 1.8 | 3.2 | 5.7333 | 101.3333 | 102.3 | 102.6
-6.7 |  | 0.9 | 1.8 | 3.35 | 6.025 | 101.4 | 102.3 | 102.5333
-6.6 |  | 0.9 | 1.9 | 3.5 | 6.275 | 101.44 | 102.3 | 102.5
-6.5 | 0.7 | 1 | 1.9 | 3.6 | 6.6333 | 101.5 | 102.3 | 102.5
-6.4 | 0.7 | 1 | 2 | 3.8 | 6.95 | 101.5 | 102.3 | 102.5
-6.3 | 0.7 | 1 | 2.1 | 3.9 | 7.275 | 101.5 | 102.3 | 102.4833
-6.2 | 0.7 | 1 | 2.1 | 4.1 | 7.8 | 101.5 | 102.3 | 102.4
-6.1 | 0.7 | 1.1 | 2.2 | 4.3 | 8.425 | 101.6 | 102.28 | 102.4
-6 | 0.7 | 1.1 | 2.3 | 4.4 | 9.2 | 101.6 | 102.2 | 102.4
-5.9 | 0.8 | 1.1 | 2.4 | 4.65 | 10.25 | 101.6 | 102.2 | 102.4
-5.8 | 0.8 | 1.15 | 2.4 | 4.8 | 12 | 101.6 | 102.2 | 102.4
-5.7 | 0.8 | 1.2 | 2.5 | 5.0667 | 14.475 | 101.625 | 102.2 | 102.3667
-5.6 | 0.8 | 1.2 | 2.6 | 5.3 | 19.05 | 101.7 | 102.2 | 102.3
-5.5 | 0.8 | 1.3 | 2.8 | 5.55 | 26 | 101.675 | 102.1 | 102.3
-5.4 | 0.8 | 1.3 | 2.9 | 5.8667 | 36.0667 | 101.6 | 102.1 | 102.2
-5.3 | 0.9 | 1.3 | 3 | 6.15 | 47.3 | 101.6 | 102.1 | 102.2
-5.2 | 0.9 | 1.4 | 3.1 | 6.55 | 58.625 | 101.6 | 102 | 102.2
-5.1 | 0.9 | 1.5 | 3.3 | 7.1 | 69.2333 | 101.6 | 102 | 102.1
-5 | 1 | 1.5 | 3.4 | 7.75 | 78.725 | 101.6 | 101.9 | 102.0143
-4.9 | 1 | 1.6 | 3.6 | 8.6 | 86.3 | 101.6 | 101.9 | 102
-4.8 | 1 | 1.6 | 3.7 | 9.85 | 90.9 | 101.5 | 101.8 | 101.95
-4.7 | 1 | 1.7 | 3.95 | 11.7 | 93.6333 | 101.5 | 101.8 | 101.8857
-4.6 | 1.1 | 1.8 | 4.1 | 14.9 | 95.1 | 101.5 | 101.72 | 101.8
-4.5 | 1.1 | 1.8 | 4.35 | 20.6 | 96.1 | 101.4 | 101.7 | 101.7333
-4.4 | 1.1 | 1.9 | 4.55 | 28.8 | 96.84 | 101.4 | 101.6 | 101.6667
-4.3 | 1.2 | 2 | 4.8 | 39.15 | 97.3333 | 101.3 | 101.5 | 101.58
-4.2 | 1.2333 | 2.1 | 5.2 | 50.6 | 97.7333 | 101.3 | 101.48 | 101.5
-4.1 | 1.3 | 2.1667 | 5.65 | 61.2 | 98.08 | 101.175 | 101.4 | 101.4
-4 | 1.3 | 2.3 | 6.15 | 72.1 | 98.24 | 101.1 | 101.3 | 101.4
-3.9 | 1.4 | 2.4 | 6.9 | 81.25 | 98.5 | 101.1 | 101.2667 | 101.3
-3.8 | 1.5 | 2.5 | 7.8 | 88.1 | 98.7 | 101 | 101.2 | 101.2857
-3.7 | 1.5 | 2.7 | 9.4 | 92 | 98.8 | 101 | 101.2 | 101.2
-3.6 | 1.6 | 2.8 | 11.8 | 94.2 | 98.9 | 101 | 101.1167 | 101.2
-3.5 | 1.7 | 3 | 15.6 | 95.5 | 99.0667 | 101 | 101.1 | 101.1667
-3.4 | 1.8 | 3.2 | 22.5 | 96.4 | 99.26 | 100.9 | 101.05 | 101.1
-3.3 | 1.9 | 3.4 | 32.1 | 97 | 99.4 | 100.9 | 101 | 101
-3.2 | 2 | 3.7 | 42.8 | 97.6 | 99.5 | 100.9 | 101 | 101
-3.1 | 2.1 | 4 | 54.55 | 98 | 99.6 | 100.9 | 100.9833 | 101
-3 | 2.2 | 4.4 | 66.5 | 98.3 | 99.675 | 100.8 | 100.9 | 101
-2.9 | 2.3 | 4.95 | 76.7333 | 98.55 | 99.825 | 100.8 | 100.9 | 100.9
-2.8 | 2.6 | 5.7 | 84.9 | 98.7333 | 99.9 | 100.8 | 100.9 | 100.9
-2.7 | 2.75 | 7 | 90.6 | 98.9 | 100.025 | 100.8 | 100.9 | 100.9
-2.6 | 3.1 | 8.7 | 93.9 | 99.2 | 100.1 | 100.8 | 100.9 | 100.8
-2.5 | 3.4 | 11.9 | 95.5 | 99.3 | 100.2 | 100.8 | 100.9 | 100.8
-2.4 | 3.8667 | 17.2 | 96.8 | 99.5 | 100.25 | 100.8 | 100.8 | 100.8
-2.3 | 4.55 | 25.75 | 97.55 | 99.6 | 100.3 | 100.8 | 100.8 | 100.8
-2.2 | 5.5 | 36.7 | 98.1 | 99.7 | 100.3 | 100.7 | 100.8 | 100.8
-2.1 | 7.3 | 48.6 | 98.45 | 99.8 | 100.4 | 100.7 | 100.7 | 100.8
-2 | 9.75 | 60.85 | 98.75 | 99.9 | 100.4 | 100.7 | 100.7 | 100.7
-1.9 | 14.75 | 72 | 99.1 | 99.9 | 100.4 | 100.6 | 100.7 | 100.7
-1.8 | 22.1 | 82.2 | 99.3 | 100 | 100.4 | 100.6 | 100.6286 | 100.65
-1.7 | 32.65 | 89.3 | 99.4 | 100.1 | 100.4 | 100.6 | 100.6 | 100.6
-1.6 | 43.7 | 93.7 | 99.55 | 100.1 | 100.4 | 100.5 | 100.6 | 100.5167
-1.5 | 56.7 | 95.8 | 99.7 | 100.1 | 100.375 | 100.5 | 100.5 | 100.5
-1.4 | 69.4 | 97.4 | 99.8 | 100.2 | 100.3 | 100.5 | 100.5 | 100.5
-1.3 | 79.3 | 98.25 | 99.9 | 100.2 | 100.3 | 100.4 | 100.4 | 100.4
-1.2 | 87.9 | 98.75 | 99.9 | 100.2 | 100.225 | 100.35 | 100.4 | 100.3714
-1.1 | 92.8 | 99.1 | 100 | 100.2 | 100.2 | 100.3 | 100.3 | 100.3
-1 | 95.75 | 99.4 | 100 | 100.2 | 100.2 | 100.3 | 100.3 | 100.24
-0.9 | 97.4 | 99.6 | 100 | 100.1 | 100.2 | 100.2 | 100.2 | 100.2
-0.8 | 98.4 | 99.7 | 100 | 100.1 | 100.2 | 100.2 | 100.2 | 100.2
-0.7 | 99.05 | 99.8 | 100 | 100 | 100.2 | 100.18 | 100.2 | 100.2
-0.6 | 99.4 | 99.9 | 100 | 100 | 100.2 | 100.1 | 100.1 | 100.1
-0.5 | 99.7 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.1
-0.4 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1 | 100.1
-0.3 | 99.9 | 99.9 | 100 | 100 | 100 | 100 | 100.0833 | 100.02
-0.2 | 100 | 99.9333 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100.0571 | 100.0833
0.4 | 99.8 | 99.9 | 100 | 100 | 100.075 | 100 | 100.1 | 100.1
0.5 | 99.5667 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1 | 100.1
0.6 | 99.4 | 99.8 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.1667
0.7 | 99.05 | 99.8 | 100 | 100.1 | 100.1 | 100.1 | 100.1714 | 100.2
0.8 | 98.2 | 99.7 | 99.9 | 100.1 | 100.15 | 100.1 | 100.2 | 100.2
0.9 | 97.3 | 99.5 | 99.9 | 100.1 | 100.2 | 100.2 | 100.2 | 100.2
1 | 95.6333 | 99.4 | 99.9 | 100.1 | 100.2 | 100.2 | 100.3 | 100.2167
1.1 | 92.7 | 99.2 | 99.8 | 100.1 | 100.2 | 100.26 | 100.3 | 100.3
1.2 | 88 | 98.8 | 99.8 | 100.1 | 100.2 | 100.3 | 100.4 | 100.3333
1.3 | 79.3 | 98.1 | 99.7333 | 100.1 | 100.2 | 100.4 | 100.4 | 100.4
1.4 | 69.3 | 97.2 | 99.65 | 100.1 | 100.275 | 100.4 | 100.5 | 100.5
1.5 | 56.6 | 95.7 | 99.6 | 100.1 | 100.3 | 100.5 | 100.5167 | 100.5
1.6 | 43.7 | 93.5 | 99.5 | 100 | 100.3 | 100.5 | 100.6 | 100.5
1.7 | 32.65 | 89.2 | 99.3 | 99.9 | 100.3 | 100.5 | 100.6 | 100.6
1.8 | 22 | 82.2 | 99.2 | 99.9 | 100.3 | 100.6 | 100.7 | 100.7
1.9 | 14.5 | 72.3 | 98.9 | 99.85 | 100.3 | 100.6 | 100.7 | 100.7
2 | 9.75 | 60.5 | 98.75 | 99.75 | 100.3 | 100.6 | 100.7 | 100.7
2.1 | 7.3 | 48.7 | 98.5 | 99.7 | 100.3 | 100.6 | 100.7 | 100.78
2.2 | 5.6 | 36.2 | 98.1 | 99.6 | 100.2 | 100.6 | 100.8 | 100.8
2.3 | 4.5667 | 25.9 | 97.5 | 99.4667 | 100.2 | 100.6 | 100.8 | 100.8
2.4 | 3.8667 | 17 | 96.7 | 99.3 | 100.1 | 100.7 | 100.8 | 100.8
2.5 | 3.4 | 11.7 | 95.4667 | 99.2 | 100.05 | 100.7 | 100.9 | 100.8
2.6 | 3 | 8.65 | 93.6 | 99 | 99.95 | 100.6 | 100.9 | 100.8714
2.7 | 2.75 | 6.9 | 90.4 | 98.9 | 99.8 | 100.675 | 100.9 | 100.9
2.8 | 2.6 | 5.7 | 84.6 | 98.6 | 99.775 | 100.7 | 100.9 | 100.9
2.9 | 2.4 | 5 | 76.7 | 98.4 | 99.7 | 100.7 | 100.9 | 100.9
3 | 2.2333 | 4.4 | 66.2667 | 98.2 | 99.6 | 100.7 | 100.9 | 101
3.1 | 2.1 | 3.95 | 54.8 | 97.8 | 99.5 | 100.7 | 100.96 | 101
3.2 | 2 | 3.6667 | 42.85 | 97.4 | 99.4 | 100.7 | 101 | 101
3.3 | 1.9 | 3.4 | 31.85 | 96.8 | 99.3 | 100.7 | 101 | 101
3.4 | 1.8 | 3.2 | 22.3 | 96.3 | 99.2 | 100.7 | 101.02 | 101.1
3.5 | 1.7 | 3 | 15.5 | 95.4 | 99 | 100.775 | 101.1 | 101.1333
3.6 | 1.6 | 2.8 | 11.7 | 93.9 | 98.75 | 100.8 | 101.1 | 101.2
3.7 | 1.5 | 2.65 | 9.2333 | 91.7 | 98.6 | 100.8 | 101.2 | 101.2
3.8 | 1.5 | 2.5 | 7.8 | 87.7 | 98.5 | 100.8 | 101.2 | 101.2571
3.9 | 1.4 | 2.4 | 6.9 | 81.4 | 98.325 | 100.9 | 101.25 | 101.3
4 | 1.3 | 2.25 | 6.15 | 71.9 | 98.075 | 100.9 | 101.3 | 101.3833
4.1 | 1.3 | 2.1 | 5.6 | 61.1 | 97.825 | 101 | 101.3714 | 101.4
4.2 | 1.2333 | 2 | 5.2 | 50.65 | 97.5 | 101 | 101.46 | 101.4667
4.3 | 1.2 | 1.9 | 4.8 | 39.15 | 97.1 | 101.1 | 101.5 | 101.55
4.4 | 1.2 | 1.9 | 4.5333 | 28.6 | 96.62 | 101.1 | 101.6 | 101.6333
4.5 | 1.1 | 1.8 | 4.35 | 20.5 | 95.9333 | 101.2 | 101.6833 | 101.8
4.6 | 1.1 | 1.7 | 4.2 | 14.85 | 94.9 | 101.25 | 101.8 | 101.8
4.7 | 1.1 | 1.6667 | 4 | 11.6 | 93.3 | 101.3 | 101.8 | 101.9
4.8 | 1 | 1.6 | 3.7 | 9.7 | 90.85 | 101.3 | 101.8 | 101.9167
4.9 | 1 | 1.5 | 3.6 | 8.5 | 86.15 | 101.3 | 101.9 | 102
5 | 1 | 1.5 | 3.4 | 7.75 | 78.875 | 101.3 | 101.9 | 102.0667
5.1 | 0.9 | 1.4 | 3.25 | 7.1 | 69.3 | 101.4 | 102 | 102.1
5.2 | 0.9 | 1.4 | 3.1 | 6.55 | 58.8333 | 101.4 | 102 | 102.2
5.3 | 0.9 | 1.3 | 3 | 6.2 | 47.35 | 101.4 | 102.1 | 102.2
5.4 | 0.9 | 1.3 | 2.85 | 5.85 | 36.325 | 101.4 | 102.1 | 102.2
5.5 | 0.8 | 1.3 | 2.8 | 5.5 | 26.2 | 101.4 | 102.1 | 102.3
5.6 | 0.8 | 1.2 | 2.6 | 5.3 | 19.15 | 101.4 | 102.16 | 102.3
5.7 | 0.8 | 1.2 | 2.55 | 5.05 | 14.5 | 101.4 | 102.2 | 102.3333
5.8 | 0.8 | 1.1 | 2.5 | 4.8 | 11.9667 | 101.4 | 102.2 | 102.4
5.9 | 0.8 | 1.1 | 2.4 | 4.65 | 10.3 | 101.34 | 102.2 | 102.4
6 | 0.8 | 1.1 | 2.3 | 4.4 | 9.1 | 101.3 | 102.2 | 102.4
6.1 | 0.7 | 1.1 | 2.2 | 4.2 | 8.4667 | 101.3 | 102.24 | 102.4
6.2 | 0.7 | 1 | 2.1 | 4.1 | 7.8333 | 101.3 | 102.3 | 102.4
6.3 | 0.7 | 1 | 2.1 | 3.9 | 7.3333 | 101.3 | 102.3 | 102.45
6.4 | 0.7 | 1 | 2 | 3.8 | 6.95 | 101.3 | 102.3 | 102.5
6.5 | 0.7 | 1 | 2 | 3.6 | 6.5667 | 101.2 | 102.3 | 102.5
6.6 |  | 1 | 1.9 | 3.5 | 6.325 | 101.125 | 102.3 | 102.5
6.7 |  | 0.9 | 1.8 | 3.4 | 6.05 | 101.1 | 102.3 | 102.5
6.8 |  | 0.9 | 1.8 | 3.2 | 5.825 | 101.1 | 102.3 | 102.6
6.9 |  | 0.9 | 1.7 | 3.1 | 5.5 | 101 | 102.3 | 102.6
7 |  | 0.9 | 1.6333 | 3 | 5.3 | 101 | 102.34 | 102.6
7.1 |  | 0.9 | 1.6 | 2.9 | 5.125 | 101 | 102.4 | 102.7
7.2 |  |  | 1.6 | 2.8333 | 4.98 | 100.9 | 102.4 | 102.7
7.3 |  |  | 1.5 | 2.75 | 4.75 | 100.82 | 102.4 | 102.7
7.4 |  |  | 1.5 | 2.7 | 4.6 | 100.8 | 102.4 | 102.7
7.5 |  |  | 1.5 | 2.6 | 4.4 | 100.7 | 102.4 | 102.7
7.6 |  |  | 1.4 | 2.5 | 4.25 | 100.7 | 102.4 | 102.8
7.7 |  |  | 1.4 | 2.4 | 4.1 | 100.6 | 102.4 | 102.8
7.8 |  |  | 1.3 | 2.4 | 4 | 100.5 | 102.4 | 102.8
7.9 |  |  | 1.3 | 2.3 | 3.825 | 100.45 | 102.4 | 102.8
8 |  |  | 1.3 | 2.2 | 3.8 | 100.3 | 102.4 | 102.8
8.1 |  |  | 1.3 | 2.2 | 3.625 | 100.2 | 102.4 | 102.9
8.2 |  |  |  | 2.1 | 3.525 | 100.1 | 102.4 | 102.9
8.3 |  |  |  | 2.1 | 3.4 | 99.9667 | 102.4 | 102.9
8.4 |  |  |  | 2 | 3.3 | 99.875 | 102.4 | 102.9
8.5 |  |  |  | 2 | 3.2 | 99.725 | 102.4 | 103
8.6 |  |  |  | 1.9 | 3.125 | 99.6 | 102.4 | 103
8.7 |  |  |  | 1.9 | 3 | 99.4 | 102.4 | 103
8.8 |  |  |  | 1.8 | 2.96 | 99.2 | 102.4 | 103
8.9 |  |  |  | 1.8 | 2.9 | 99.0333 | 102.4 | 103
9 |  |  |  | 1.7667 | 2.8 | 98.85 | 102.3167 | 103
9.1 |  |  |  | 1.7 | 2.7 | 98.625 | 102.3 | 103
9.2 |  |  |  |  | 2.7 | 98.3667 | 102.3 | 103
9.3 |  |  |  |  | 2.6 | 98.05 | 102.3 | 103
9.4 |  |  |  |  | 2.5 | 97.725 | 102.3 | 103
9.5 |  |  |  |  | 2.4667 | 97.3667 | 102.3 | 103
9.6 |  |  |  |  | 2.4 | 96.825 | 102.3 | 103
9.7 |  |  |  |  | 2.4 | 96.2 | 102.3 | 103
9.8 |  |  |  |  | 2.3 | 95.4 | 102.3 | 103
9.9 |  |  |  |  | 2.3 | 94.2 | 102.3 | 103.1
10 |  |  |  |  | 2.2 | 92.45 | 102.3 | 103.1
10.1 |  |  |  |  | 2.2 | 89.6 | 102.3 | 103.1
10.2 |  |  |  |  | 2.1 | 84.1 | 102.28 | 103.1
10.3 | [5 empty] | 76.4333 | 102.2 | 103.2
10.4 | [5 empty] | 67.1 | 102.2 | 103.2
10.5 | [5 empty] | 56.325 | 102.2 | 103.2
10.6 | [5 empty] | 45.7 | 102.15 | 103.2
10.7 | [5 empty] | 35.275 | 102.1 | 103.2
10.8 | [5 empty] | 26.8 | 102.1 | 103.3
10.9 | [5 empty] | 20.4 | 102.1 | 103.3
11 | [5 empty] | 16.56 | 102.1 | 103.3
11.1 | [5 empty] | 14.5 | 102.1 | 103.3
11.2 | [5 empty] | 13.075 | 102.04 | 103.3
11.3 | [5 empty] | 12.1 | 102 | 103.3
11.4 | [5 empty] | 11.4 | 102 | 103.34
11.5 | [5 empty] | 10.75 | 101.9 | 103.4
11.6 | [5 empty] | 10.4 | 101.9 | 103.4
11.7 | [5 empty] | 9.95 | 101.8667 | 103.4
11.8 | [5 empty] | 9.6 | 101.8 | 103.36
11.9 | [5 empty] | 9.275 | 101.7833 | 103.4
12 | [5 empty] | 9 | 101.7 | 103.4
12.1 | [5 empty] | 8.7333 | 101.6571 | 103.4
12.2 | [5 empty] | 8.5 | 101.6 | 103.4
12.3 | [5 empty] | 8.2 | 101.5 | 103.4
12.4 | [5 empty] | 8 | 101.4286 | 103.4
12.5 | [5 empty] | 7.775 | 101.3833 | 103.4
12.6 | [5 empty] | 7.575 | 101.3 | 103.4
12.7 | [5 empty] | 7.4 | 101.3 | 103.4
12.8 | [5 empty] | 7.2 | 101.2 | 103.4
12.9 | [5 empty] | 6.975 | 101.1 | 103.4
13 | [5 empty] | 6.8 | 101 | 103.4
13.1 | [5 empty] | 6.64 | 100.9 | 103.4
13.2 | [5 empty] | 6.4333 | 100.86 | 103.4
13.3 | [5 empty] | 6.3 | 100.75 | 103.4
13.4 | [5 empty] | 6.1 | 100.6 | 103.4
13.5 | [5 empty] | 6 | 100.5 | 103.4
13.6 | [5 empty] | 5.9 | 100.42 | 103.4
13.7 | [5 empty] | 5.75 | 100.3 | 103.4
13.8 | [5 empty] | 5.6 | 100.1571 | 103.4
13.9 | [5 empty] | 5.46 | 100 | 103.4
14 | [5 empty] | 5.3 | 99.8333 | 103.4
14.1 | [5 empty] | 5.2 | 99.6714 | 103.4
14.2 | [5 empty] | 5.1 | 99.5 | 103.4
14.3 | [5 empty] | 5 | 99.325 | 103.4
14.4 | [5 empty] | 4.9 | 99.12 | 103.4
14.5 | [5 empty] | 4.8 | 98.8 | 103.4
14.6 | [5 empty] | 4.7 | 98.6 | 103.4
14.7 | [5 empty] | 4.6 | 98.2333 | 103.5
14.8 | [5 empty] | 4.5 | 97.78 | 103.5
14.9 | [5 empty] | 4.35 | 97.35 | 103.5
15 | [5 empty] | 4.3 | 96.56 | 103.5
15.1 | [5 empty] | 4.2 | 95.625 | 103.5
15.2 | [5 empty] | 4.16 | 94.3667 | 103.4333
15.3 | [5 empty] | 4.025 | 92.3 | 103.5
15.4 | [5 empty] | 4 | 88.4 | 103.4167
15.5 | [6 empty] | 82.0333 | 103.4
15.6 | [6 empty] | 74.0429 | 103.4
15.7 | [6 empty] | 63.6 | 103.4
15.8 | [6 empty] | 53.26 | 103.38
15.9 | [6 empty] | 42.5571 | 103.3
16 | [6 empty] | 33.55 | 103.3
16.1 | [6 empty] | 25.56 | 103.25
16.2 | [6 empty] | 20.4 | 103.2
16.3 | [6 empty] | 17.48 | 103.2
16.4 | [6 empty] | 15.7333 | 103.2
16.5 | [6 empty] | 14.6 | 103.1
16.6 | [6 empty] | 13.7333 | 103.1
16.7 | [6 empty] | 13.1 | 103
16.8 | [6 empty] | 12.56 | 103
16.9 | [6 empty] | 12.1 | 103
17 | [6 empty] | 11.78 | 102.9
17.1 | [6 empty] | 11.38 | 102.9
17.2 | [6 empty] | 11.0833 | 102.8167
17.3 | [6 empty] | 10.8333 | 102.8
17.4 | [6 empty] | 10.52 | 102.7
17.5 | [6 empty] | 10.24 | 102.6167
17.6 | [6 empty] | 10.0571 | 102.6
17.7 | [6 empty] | 9.8 | 102.5
17.8 | [6 empty] | 9.6 | 102.5
17.9 | [6 empty] | 9.4 | 102.4
18 | [6 empty] | 9.18 | 102.3
18.1 | [6 empty] | 9 | 102.2
18.2 | [6 empty] | 8.8143 | 102.1
18.3 | [6 empty] | 8.55 | 102
18.4 | [6 empty] | 8.36 | 101.95
18.5 | [6 empty] | 8.175 | 101.8
18.6 | [6 empty] | 7.9833 | 101.7167
18.7 | [6 empty] | 7.8167 | 101.65
18.8 | [6 empty] | 7.6571 | 101.4833
18.9 | [6 empty] | 7.5 | 101.4
19 | [6 empty] | 7.4 | 101.2167
19.1 | [6 empty] | 7.2333 | 101.025
19.2 | [6 empty] | 7.1 | 100.9667
19.3 | [6 empty] | 6.9667 | 100.8
19.4 | [6 empty] | 6.7833 | 100.6
19.5 | [6 empty] | 6.68 | 100.3333
19.6 | [6 empty] | 6.5143 | 100.1429
19.7 | [6 empty] | 6.45 | 99.9167
19.8 | [6 empty] | 6.2714 | 99.65
19.9 | [6 empty] | 6.2 | 99.38
20 | [6 empty] | 6.05 | 99.05
20.1 | [6 empty] | 5.95 | 98.55
20.2 | [6 empty] | 5.8833 | 97.9286
20.3 | [6 empty] | 5.7167 | 97.2
20.4 | [6 empty] | 5.6333 | 96.06
20.5 | [6 empty] | 5.56 | 94.5667
20.6 | [6 empty] | 5.4625 | 91.8
20.7 | [7 empty] | 86.8
20.8 | [7 empty] | 79.64
20.9 | [7 empty] | 70.1
21 | [7 empty] | 59.46
21.1 | [7 empty] | 48.9
21.2 | [7 empty] | 38.7
21.3 | [7 empty] | 30.35
21.4 | [7 empty] | 24
21.5 | [7 empty] | 19.9429
21.6 | [7 empty] | 17.7
21.7 | [7 empty] | 16.25
21.8 | [7 empty] | 15.3
21.9 | [7 empty] | 14.5667
22 | [7 empty] | 14
22.1 | [7 empty] | 13.5
22.2 | [7 empty] | 13.1
22.3 | [7 empty] | 12.75
22.4 | [7 empty] | 12.42
22.5 | [7 empty] | 12.1
22.6 | [7 empty] | 11.85
22.7 | [7 empty] | 11.5667
22.8 | [7 empty] | 11.3286
22.9 | [7 empty] | 11.0571
23 | [7 empty] | 10.8
23.1 | [7 empty] | 10.54
23.2 | [7 empty] | 10.3333
23.3 | [7 empty] | 10.1333
23.4 | [7 empty] | 9.9
23.5 | [7 empty] | 9.65
23.6 | [7 empty] | 9.46
23.7 | [7 empty] | 9.28
23.8 | [7 empty] | 9.0286
23.9 | [7 empty] | 8.85
24 | [7 empty] | 8.68
24.1 | [7 empty] | 8.4857
24.2 | [7 empty] | 8.32
24.3 | [7 empty] | 8.15
24.4 | [7 empty] | 7.95
24.5 | [7 empty] | 7.7667
24.6 | [7 empty] | 7.6
24.7 | [7 empty] | 7.48
24.8 | [7 empty] | 7.3
24.9 | [7 empty] | 7.2
25 | [7 empty] | 7.0429
25.1 | [7 empty] | 6.9
25.2 | [7 empty] | 6.725
25.3 | [7 empty] | 6.6
25.4 | [7 empty] | 6.4667
25.5 | [7 empty] | 6.4
25.6 | [7 empty] | 6.2333
25.7 | [7 empty] | 6.1667
25.8 | [7 empty] | 6
25.9 | [7 empty] | 5.9
26 | [7 empty] | 5.7667
26.1 | [7 empty] | 5.7
26.2 | [7 empty] | 5.6
26.3 | [7 empty] | 5.5
26.4 | [7 empty] | 5.4
26.5 | [7 empty] | 5.3
26.6 | [7 empty] | 5.2167
26.7 | [7 empty] | 5.14
26.8 | [7 empty] | 5.0571
26.9 | [7 empty] | 4.98
27 | [7 empty] | 4.9
27.1 | [7 empty] | 4.8
27.2 | [7 empty] | 4.7
27.3 | [7 empty] | 4.68
27.4 | [7 empty] | 4.6
27.5 | [7 empty] | 4.52
27.6 | [7 empty] | 4.5
27.7 | [7 empty] | 4.4
27.8 | [7 empty] | 4.4
27.9 | [7 empty] | 4.3

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-26.9 | [7 empty] | 8.4167
-26.8 | [7 empty] | 8.5833
-26.7 | [7 empty] | 8.7
-26.6 | [7 empty] | 8.8
-26.5 | [7 empty] | 8.95
-26.4 | [7 empty] | 9.1167
-26.3 | [7 empty] | 9.3143
-26.2 | [7 empty] | 9.48
-26.1 | [7 empty] | 9.64
-26 | [7 empty] | 9.8143
-25.9 | [7 empty] | 9.9667
-25.8 | [7 empty] | 10.2
-25.7 | [7 empty] | 10.3667
-25.6 | [7 empty] | 10.5
-25.5 | [7 empty] | 10.7667
-25.4 | [7 empty] | 10.975
-25.3 | [7 empty] | 11.18
-25.2 | [7 empty] | 11.35
-25.1 | [7 empty] | 11.6167
-25 | [7 empty] | 11.7833
-24.9 | [7 empty] | 12.05
-24.8 | [7 empty] | 12.2333
-24.7 | [7 empty] | 12.5
-24.6 | [7 empty] | 12.7333
-24.5 | [7 empty] | 13.02
-24.4 | [7 empty] | 13.2667
-24.3 | [7 empty] | 13.5
-24.2 | [7 empty] | 13.8
-24.1 | [7 empty] | 14.1
-24 | [7 empty] | 14.4167
-23.9 | [7 empty] | 14.7167
-23.8 | [7 empty] | 15.0143
-23.7 | [7 empty] | 15.3
-23.6 | [7 empty] | 15.68
-23.5 | [7 empty] | 16
-23.4 | [7 empty] | 16.3333
-23.3 | [7 empty] | 16.7667
-23.2 | [7 empty] | 17.2
-23.1 | [7 empty] | 17.6667
-23 | [7 empty] | 18.18
-22.9 | [7 empty] | 18.76
-22.8 | [7 empty] | 19.55
-22.7 | [7 empty] | 20.55
-22.6 | [7 empty] | 22
-22.5 | [7 empty] | 24.26
-22.4 | [7 empty] | 27.94
-22.3 | [7 empty] | 33.7429
-22.2 | [7 empty] | 41.12
-22.1 | [7 empty] | 49.5571
-22 | [7 empty] | 58.2
-21.9 | [7 empty] | 67.1667
-21.8 | [7 empty] | 75.25
-21.7 | [7 empty] | 81.0667
-21.6 | [7 empty] | 85.1167
-21.5 | [7 empty] | 87.3667
-21.4 | [6 empty] | 7.4 | 88.9
-21.3 | [6 empty] | 7.5286 | 89.96
-21.2 | [6 empty] | 7.6833 | 90.775
-21.1 | [6 empty] | 7.8 | 91.45
-21 | [6 empty] | 8 | 91.98
-20.9 | [6 empty] | 8.1 | 92.4
-20.8 | [6 empty] | 8.2833 | 92.7429
-20.7 | [6 empty] | 8.44 | 93.2
-20.6 | [6 empty] | 8.6 | 93.54
-20.5 | [6 empty] | 8.6667 | 93.8
-20.4 | [6 empty] | 8.8667 | 94.2
-20.3 | [6 empty] | 9.1 | 94.3
-20.2 | [6 empty] | 9.16 | 94.6333
-20.1 | [6 empty] | 9.4167 | 94.8667
-20 | [6 empty] | 9.6 | 95.1
-19.9 | [6 empty] | 9.75 | 95.3
-19.8 | [6 empty] | 9.9167 | 95.6333
-19.7 | [6 empty] | 10.1 | 95.8286
-19.6 | [6 empty] | 10.3 | 95.9167
-19.5 | [6 empty] | 10.5 | 96.1667
-19.4 | [6 empty] | 10.8 | 96.35
-19.3 | [6 empty] | 10.9667 | 96.5167
-19.2 | [6 empty] | 11.1667 | 96.6833
-19.1 | [6 empty] | 11.4333 | 96.8667
-19 | [6 empty] | 11.7 | 97
-18.9 | [6 empty] | 11.9 | 97.15
-18.8 | [6 empty] | 12.1667 | 97.25
-18.7 | [6 empty] | 12.4 | 97.4333
-18.6 | [6 empty] | 12.725 | 97.6
-18.5 | [6 empty] | 13.025 | 97.8
-18.4 | [6 empty] | 13.28 | 97.9
-18.3 | [6 empty] | 13.54 | 98.0333
-18.2 | [6 empty] | 13.9167 | 98.1857
-18.1 | [6 empty] | 14.1667 | 98.3
-18 | [6 empty] | 14.55 | 98.4
-17.9 | [6 empty] | 14.92 | 98.5167
-17.8 | [6 empty] | 15.25 | 98.6
-17.7 | [6 empty] | 15.7 | 98.75
-17.6 | [6 empty] | 16.1143 | 98.82
-17.5 | [6 empty] | 16.62 | 98.9
-17.4 | [6 empty] | 17.24 | 99
-17.3 | [6 empty] | 18 | 99.1333
-17.2 | [6 empty] | 19 | 99.22
-17.1 | [6 empty] | 20.4143 | 99.3
-17 | [6 empty] | 22.45 | 99.35
-16.9 | [6 empty] | 25.96 | 99.5
-16.8 | [6 empty] | 31.65 | 99.6
-16.7 | [6 empty] | 39.1833 | 99.7
-16.6 | [6 empty] | 47.4 | 99.8
-16.5 | [6 empty] | 56.3 | 99.8571
-16.4 | [6 empty] | 65.88 | 99.9
-16.3 | [6 empty] | 74.1 | 99.9
-16.2 | [6 empty] | 80.1333 | 100
-16.1 | [6 empty] | 84.7 | 100.1
-16 | [5 empty] | 5.625 | 87.46 | 100.1
-15.9 | [5 empty] | 5.775 | 89.1667 | 100.1667
-15.8 | [5 empty] | 5.9 | 90.28 | 100.24
-15.7 | [5 empty] | 6 | 91.02 | 100.3
-15.6 | [5 empty] | 6.1 | 91.7 | 100.3833
-15.5 | [5 empty] | 6.2 | 92.325 | 100.4
-15.4 | [5 empty] | 6.3 | 92.75 | 100.4
-15.3 | [5 empty] | 6.5 | 93.1 | 100.4833
-15.2 | [5 empty] | 6.6 | 93.46 | 100.5
-15.1 | [5 empty] | 6.7 | 93.88 | 100.6
-15 | [5 empty] | 6.9 | 94.18 | 100.6
-14.9 | [5 empty] | 7 | 94.44 | 100.6333
-14.8 | [5 empty] | 7.1 | 94.725 | 100.7
-14.7 | [5 empty] | 7.3 | 94.95 | 100.7
-14.6 | [5 empty] | 7.44 | 95.22 | 100.8
-14.5 | [5 empty] | 7.6333 | 95.48 | 100.8
-14.4 | [5 empty] | 7.9 | 95.6571 | 100.9
-14.3 | [5 empty] | 8 | 95.9167 | 100.9
-14.2 | [5 empty] | 8.16 | 96.1167 | 100.9
-14.1 | [5 empty] | 8.3 | 96.3 | 101
-14 | [5 empty] | 8.5 | 96.4 | 101
-13.9 | [5 empty] | 8.7667 | 96.6 | 101
-13.8 | [5 empty] | 8.95 | 96.82 | 101
-13.7 | [5 empty] | 9.15 | 97 | 101
-13.6 | [5 empty] | 9.3 | 97.15 | 101.02
-13.5 | [5 empty] | 9.6 | 97.3 | 101.1
-13.4 | [5 empty] | 9.8 | 97.44 | 101.1
-13.3 | [5 empty] | 10.075 | 97.54 | 101.2
-13.2 | [5 empty] | 10.34 | 97.74 | 101.2
-13.1 | [5 empty] | 10.6 | 97.82 | 101.2
-13 | [5 empty] | 10.875 | 98 | 101.2714
-12.9 | [5 empty] | 11.125 | 98.1 | 101.3
-12.8 | [5 empty] | 11.46 | 98.2333 | 101.3
-12.7 | [5 empty] | 11.7 | 98.32 | 101.3
-12.6 | [5 empty] | 12.0667 | 98.4857 | 101.4
-12.5 | [5 empty] | 12.4 | 98.55 | 101.4
-12.4 | [5 empty] | 12.75 | 98.7 | 101.4
-12.3 | [5 empty] | 13.0667 | 98.8 | 101.4
-12.2 | [5 empty] | 13.5 | 98.9 | 101.4
-12.1 | [5 empty] | 14 | 99 | 101.4
-12 | [5 empty] | 14.55 | 99.1 | 101.44
-11.9 | [5 empty] | 15.125 | 99.14 | 101.5
-11.8 | [5 empty] | 15.9 | 99.22 | 101.5
-11.7 | [5 empty] | 16.9 | 99.325 | 101.5
-11.6 | [5 empty] | 18.175 | 99.4 | 101.5
-11.5 | [5 empty] | 20.4 | 99.5 | 101.6
-11.4 | [5 empty] | 23.825 | 99.6 | 101.6
-11.3 | [5 empty] | 29.55 | 99.6571 | 101.6
-11.2 | [5 empty] | 36.72 | 99.7 | 101.6
-11.1 | [5 empty] | 46.12 | 99.8 | 101.6
-11 | [5 empty] | 55.46 | 99.9 | 101.6
-10.9 | [5 empty] | 64.9 | 99.9333 | 101.6
-10.8 | [5 empty] | 73.55 | 100 | 101.6
-10.7 | [5 empty] | 80.7 | 100.1 | 101.6
-10.6 | [5 empty] | 85.4 | 100.1 | 101.6
-10.5 | [5 empty] | 88.5 | 100.1 | 101.6
-10.4 |  |  |  |  | 3.3 | 90.05 | 100.2 | 101.6
-10.3 |  |  |  |  | 3.32 | 91.3 | 100.2167 | 101.6
-10.2 |  |  |  |  | 3.4 | 92.25 | 100.3 | 101.6
-10.1 |  |  |  |  | 3.5 | 92.85 | 100.3 | 101.6
-10 |  |  |  |  | 3.575 | 93.4333 | 100.4 | 101.65
-9.9 |  |  |  |  | 3.675 | 93.925 | 100.4333 | 101.7
-9.8 |  |  |  |  | 3.7333 | 94.32 | 100.5 | 101.7
-9.7 |  |  |  |  | 3.9 | 94.7 | 100.5167 | 101.7286
-9.6 |  |  |  |  | 4 | 95.05 | 100.6 | 101.8
-9.5 |  |  |  |  | 4.025 | 95.4 | 100.65 | 101.8
-9.4 |  |  |  |  | 4.1 | 95.6667 | 100.7 | 101.8
-9.3 |  |  |  | 2.6667 | 4.25 | 96.05 | 100.7833 | 101.8
-9.2 |  |  |  | 2.7 | 4.4 | 96.3 | 100.8 | 101.8
-9.1 |  |  |  | 2.8 | 4.5 | 96.525 | 100.8 | 101.7
-9 |  |  |  | 2.9 | 4.6 | 96.775 | 100.8714 | 101.7
-8.9 |  |  |  | 2.9 | 4.72 | 96.975 | 100.9 | 101.7
-8.8 |  |  |  | 3 | 4.8667 | 97.22 | 100.9 | 101.7
-8.7 |  |  |  | 3.05 | 5 | 97.4 | 100.9167 | 101.7333
-8.6 |  |  |  | 3.2 | 5.1 | 97.6 | 101 | 101.8
-8.5 |  |  |  | 3.2 | 5.3 | 97.75 | 101 | 101.8
-8.4 |  |  |  | 3.3 | 5.425 | 97.9 | 101 | 101.8
-8.3 |  |  | 2 | 3.4 | 5.6 | 98.1 | 101.1 | 101.8
-8.2 |  |  | 2 | 3.5 | 5.75 | 98.2 | 101.1 | 101.8
-8.1 |  |  | 2.1 | 3.6 | 5.9 | 98.325 | 101.1 | 101.8
-8 |  |  | 2.1 | 3.7 | 6.15 | 98.5 | 101.1 | 101.8
-7.9 |  |  | 2.15 | 3.8 | 6.3 | 98.7 | 101.1333 | 101.7
-7.8 |  |  | 2.2 | 3.9 | 6.5667 | 98.7667 | 101.2 | 101.7
-7.7 |  |  | 2.3 | 4 | 6.7333 | 98.9 | 101.2 | 101.7
-7.6 |  |  | 2.3 | 4.1 | 7 | 99.1 | 101.2 | 101.7
-7.5 |  |  | 2.4 | 4.3 | 7.2 | 99.2 | 101.2 | 101.7
-7.4 |  |  | 2.4 | 4.4 | 7.45 | 99.3 | 101.2833 | 101.7
-7.3 |  |  | 2.5 | 4.5 | 7.7333 | 99.4 | 101.3 | 101.7
-7.2 |  |  | 2.6 | 4.7 | 7.975 | 99.5 | 101.3 | 101.7
-7.1 |  | 1.3 | 2.6 | 4.8 | 8.25 | 99.6 | 101.3 | 101.7
-7 |  | 1.4 | 2.7 | 5 | 8.62 | 99.7 | 101.3 | 101.7
-6.9 |  | 1.4 | 2.8 | 5.2 | 8.9 | 99.8 | 101.3 | 101.7
-6.8 |  | 1.4 | 2.9 | 5.3 | 9.2 | 99.8 | 101.3 | 101.7
-6.7 |  | 1.5 | 3 | 5.5 | 9.6 | 99.9 | 101.3 | 101.7
-6.6 | 1 | 1.5 | 3.1 | 5.7 | 10 | 99.95 | 101.3 | 101.7
-6.5 | 1 | 1.5 | 3.2 | 5.9 | 10.5 | 100 | 101.3 | 101.7
-6.4 | 1 | 1.6 | 3.3 | 6.1 | 11.14 | 100.1 | 101.3 | 101.7
-6.3 | 1.1 | 1.6 | 3.4333 | 6.4 | 11.86 | 100.2 | 101.3 | 101.7
-6.2 | 1.1 | 1.7 | 3.5 | 6.6 | 12.8 | 100.3 | 101.4 | 101.7
-6.1 | 1.1 | 1.7 | 3.7 | 6.85 | 14.225 | 100.3 | 101.4 | 101.7
-6 | 1.2 | 1.8 | 3.8 | 7.1 | 16.275 | 100.4 | 101.4 | 101.7
-5.9 | 1.2 | 1.8 | 3.9 | 7.4 | 19.45 | 100.4 | 101.4 | 101.6
-5.8 | 1.2 | 1.85 | 4.1 | 7.7 | 25.4 | 100.5 | 101.4 | 101.6
-5.7 | 1.2 | 1.9 | 4.2 | 8 | 33.2 | 100.5 | 101.4 | 101.6
-5.6 | 1.3 | 2 | 4.4 | 8.4 | 42.95 | 100.5 | 101.3 | 101.6
-5.5 | 1.3 | 2 | 4.5 | 8.8 | 52.45 | 100.6 | 101.3 | 101.5333
-5.4 | 1.3 | 2.1 | 4.6333 | 9.3 | 63.52 | 100.6 | 101.3 | 101.5
-5.3 | 1.4 | 2.2 | 4.9 | 9.9 | 72.6 | 100.6 | 101.3 | 101.5
-5.2 | 1.4 | 2.3 | 5.1 | 10.6 | 80.7 | 100.6 | 101.3 | 101.4714
-5.1 | 1.5 | 2.4 | 5.3 | 11.4 | 86.4 | 100.7 | 101.3 | 101.4
-5 | 1.5 | 2.4 | 5.5 | 12.75 | 90.1 | 100.6 | 101.2 | 101.4
-4.9 | 1.6 | 2.5 | 5.7 | 14.75 | 91.94 | 100.6 | 101.2 | 101.3
-4.8 | 1.6 | 2.6 | 5.9667 | 18.0333 | 93.175 | 100.65 | 101.2 | 101.3
-4.7 | 1.7 | 2.7 | 6.2 | 23.25 | 94.2 | 100.6 | 101.1 | 101.2
-4.6 | 1.7 | 2.8 | 6.5 | 31.2 | 94.92 | 100.6 | 101.1 | 101.2
-4.5 | 1.8 | 2.9 | 6.8667 | 40.7 | 95.4333 | 100.6 | 101.0333 | 101.1
-4.4 | 1.8 | 3.05 | 7.2 | 51.15 | 95.8 | 100.6 | 101 | 101
-4.3 | 1.9 | 3.2 | 7.65 | 61.1 | 96.2 | 100.6 | 100.9143 | 101
-4.2 | 2 | 3.4 | 8.15 | 71.2333 | 96.3667 | 100.6 | 100.9 | 100.9
-4.1 | 2 | 3.5 | 8.85 | 79.6 | 96.7333 | 100.6 | 100.8857 | 100.9
-4 | 2.1 | 3.6667 | 9.8 | 85.9 | 96.975 | 100.575 | 100.8 | 100.8
-3.9 | 2.3 | 3.8 | 11 | 89.9667 | 97.25 | 100.5 | 100.8 | 100.8
-3.8 | 2.3 | 4 | 12.9 | 92.1 | 97.525 | 100.5 | 100.8 | 100.8
-3.7 | 2.4 | 4.2 | 15.9 | 93.6 | 97.76 | 100.5 | 100.7 | 100.8
-3.6 | 2.5 | 4.4 | 21.3 | 94.6333 | 97.9 | 100.5 | 100.7 | 100.72
-3.5 | 2.7 | 4.6667 | 29.2 | 95.25 | 98.1 | 100.5 | 100.7 | 100.7
-3.4 | 2.8 | 4.9 | 38.3333 | 95.9 | 98.3 | 100.5 | 100.7 | 100.7
-3.3 | 3 | 5.2 | 49.4 | 96.4 | 98.5 | 100.5 | 100.7 | 100.7
-3.2 | 3.1 | 5.6 | 60.6 | 96.7 | 98.66 | 100.5 | 100.7 | 100.7
-3.1 | 3.25 | 6.05 | 70.9 | 97.0667 | 98.8 | 100.5 | 100.6 | 100.7
-3 | 3.4 | 6.8 | 80.05 | 97.4 | 98.9 | 100.5 | 100.7 | 100.7
-2.9 | 3.6 | 7.6 | 86.95 | 97.6 | 99 | 100.5 | 100.7 | 100.7
-2.8 | 3.85 | 8.8 | 91.1 | 97.9 | 99.1333 | 100.5 | 100.7 | 100.7
-2.7 | 4.1 | 10.75 | 93.4 | 98.2 | 99.3 | 100.6 | 100.7 | 100.7
-2.6 | 4.5333 | 13.9 | 94.9 | 98.4 | 99.5 | 100.6 | 100.6333 | 100.6
-2.5 | 5 | 19.4 | 96 | 98.6 | 99.6 | 100.6 | 100.6 | 100.6
-2.4 | 5.7 | 27.5 | 96.7667 | 98.8 | 99.7 | 100.6 | 100.6 | 100.6
-2.3 | 6.7 | 37.5 | 97.4667 | 99 | 99.725 | 100.6 | 100.6 | 100.6
-2.2 | 8.2 | 49.4 | 97.8667 | 99.1 | 99.8 | 100.6 | 100.6 | 100.6
-2.1 | 10.6 | 61.0667 | 98.2 | 99.2 | 99.9 | 100.6 | 100.6 | 100.6
-2 | 14.5 | 71.6 | 98.45 | 99.3 | 99.96 | 100.5 | 100.5 | 100.5
-1.9 | 21.1 | 81.2 | 98.6 | 99.5 | 100 | 100.5 | 100.5 | 100.5
-1.8 | 30.9 | 88.4 | 98.8 | 99.5 | 100.1 | 100.5 | 100.5 | 100.5
-1.7 | 41.5 | 92.7 | 99 | 99.6 | 100.1 | 100.5 | 100.5 | 100.46
-1.6 | 54.05 | 95.15 | 99.1 | 99.7 | 100.1 | 100.5 | 100.4 | 100.4
-1.5 | 66 | 96.6 | 99.2667 | 99.7 | 100.1 | 100.4 | 100.4 | 100.4
-1.4 | 76.6 | 97.6 | 99.4 | 99.7 | 100.1 | 100.34 | 100.3 | 100.4
-1.3 | 85.6 | 98.25 | 99.55 | 99.8 | 100.1 | 100.3 | 100.3 | 100.3
-1.2 | 91.2 | 98.8333 | 99.6 | 99.9 | 100.1 | 100.3 | 100.3 | 100.3
-1.1 | 94.7 | 99.1 | 99.7 | 99.9 | 100.1 | 100.3 | 100.2 | 100.2333
-1 | 96.6 | 99.3 | 99.7667 | 100 | 100.1 | 100.3 | 100.2 | 100.2
-0.9 | 97.8 | 99.5 | 99.8 | 100 | 100.1 | 100.2 | 100.2 | 100.1167
-0.8 | 98.5 | 99.6 | 99.9 | 100 | 100.1 | 100.2 | 100.2 | 100.1
-0.7 | 99.1 | 99.7 | 99.9 | 100 | 100.1 | 100.2 | 100.16 | 100.1
-0.6 | 99.35 | 99.8 | 99.95 | 100 | 100.1 | 100.2 | 100.1 | 100.1
-0.5 | 99.6 | 99.9 | 100 | 100 | 100.1 | 100.2 | 100.1 | 100.1
-0.4 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.1
-0.3 | 99.9 | 100 | 100 | 100 | 100 | 100.1 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100 | 100.1 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.4 | 99.7 | 99.9 | 100 | 99.9 | 100 | 100 | 100.1 | 100.1
0.5 | 99.5 | 99.8 | 99.9 | 99.9 | 100 | 100.1 | 100.1 | 100.1
0.6 | 99.25 | 99.7 | 99.9 | 99.9333 | 100 | 100.1 | 100.1 | 100.1
0.7 | 99 | 99.6 | 99.9 | 100 | 100 | 100.1 | 100.16 | 100.1
0.8 | 98.4 | 99.5 | 99.85 | 99.9 | 100 | 100.1 | 100.2 | 100.1
0.9 | 97.6 | 99.35 | 99.75 | 99.9 | 100.1 | 100.1 | 100.2 | 100.1167
1 | 96.5 | 99.2 | 99.7 | 99.9 | 100.1 | 100.16 | 100.2 | 100.2
1.1 | 94.3 | 99 | 99.6 | 99.8 | 100.1 | 100.2 | 100.2 | 100.2333
1.2 | 90.95 | 98.6333 | 99.5 | 99.8 | 100.05 | 100.3 | 100.3 | 100.3
1.3 | 85.0333 | 98.15 | 99.45 | 99.8 | 100 | 100.3 | 100.3 | 100.3
1.4 | 76.4333 | 97.5 | 99.4 | 99.7 | 100 | 100.3 | 100.3 | 100.4
1.5 | 65.75 | 96.45 | 99.2667 | 99.7 | 100 | 100.4 | 100.4 | 100.4
1.6 | 54.05 | 95 | 99.1 | 99.6 | 100.1 | 100.4 | 100.4 | 100.4
1.7 | 42.05 | 92.6 | 98.9667 | 99.5 | 100.08 | 100.4 | 100.5 | 100.46
1.8 | 31 | 88.3 | 98.8 | 99.5 | 100 | 100.4 | 100.5 | 100.5
1.9 | 21.3 | 81.7 | 98.6 | 99.4 | 100 | 100.4 | 100.5 | 100.5
2 | 14.6 | 71.6 | 98.35 | 99.3 | 99.9 | 100.4 | 100.5143 | 100.5
2.1 | 10.5 | 60.3 | 98.05 | 99.1 | 99.85 | 100.4 | 100.6 | 100.6
2.2 | 8.2 | 48.9667 | 97.65 | 99 | 99.8 | 100.4 | 100.6 | 100.6
2.3 | 6.7 | 37.2 | 97.15 | 98.8 | 99.7 | 100.4 | 100.6 | 100.6
2.4 | 5.6 | 27.45 | 96.6333 | 98.6 | 99.575 | 100.4 | 100.6 | 100.6
2.5 | 5.05 | 19.35 | 95.9 | 98.5 | 99.5 | 100.4 | 100.6 | 100.6
2.6 | 4.5667 | 13.9 | 94.8 | 98.25 | 99.4 | 100.4 | 100.64 | 100.6
2.7 | 4.15 | 10.65 | 93.2 | 98.0333 | 99.3 | 100.4 | 100.7 | 100.7
2.8 | 3.85 | 8.8 | 90.8 | 97.8 | 99.1 | 100.4 | 100.7 | 100.7
2.9 | 3.6 | 7.5 | 86.85 | 97.5667 | 99 | 100.4 | 100.7 | 100.7
3 | 3.4 | 6.7 | 80 | 97.3 | 98.875 | 100.4 | 100.7 | 100.7
3.1 | 3.25 | 6.05 | 70.7 | 97 | 98.8 | 100.4 | 100.6 | 100.7
3.2 | 3.05 | 5.55 | 60.4 | 96.6 | 98.65 | 100.4 | 100.7 | 100.7
3.3 | 2.9 | 5.2 | 49.9 | 96.2 | 98.4 | 100.3 | 100.7 | 100.7
3.4 | 2.8 | 4.85 | 38.4333 | 95.8 | 98.2 | 100.3 | 100.7 | 100.7
3.5 | 2.7 | 4.6 | 28.8 | 95.15 | 98 | 100.3 | 100.7 | 100.7
3.6 | 2.6 | 4.35 | 21.1333 | 94.4 | 97.8 | 100.3 | 100.7 | 100.72
3.7 | 2.5 | 4.2 | 15.9 | 93.4 | 97.6 | 100.3 | 100.7 | 100.8
3.8 | 2.3 | 4 | 12.9 | 92 | 97.3333 | 100.3 | 100.8 | 100.8
3.9 | 2.25 | 3.8 | 11 | 89.7 | 97.1 | 100.4 | 100.8 | 100.8
4 | 2.1333 | 3.65 | 9.8 | 85.8 | 96.875 | 100.4 | 100.8 | 100.8
4.1 | 2.05 | 3.5 | 8.9 | 79.9 | 96.625 | 100.3 | 100.9 | 100.9
4.2 | 2 | 3.3 | 8.2 | 71 | 96.3 | 100.3 | 100.9 | 100.9
4.3 | 1.95 | 3.2 | 7.7 | 61.6 | 96.05 | 100.3 | 100.9167 | 101
4.4 | 1.8 | 3.1 | 7.2 | 51.2 | 95.6333 | 100.3 | 101 | 101
4.5 | 1.8 | 2.95 | 6.8667 | 40.8 | 95.2 | 100.375 | 101.05 | 101.1
4.6 | 1.7 | 2.8 | 6.5 | 31.2333 | 94.7 | 100.4 | 101.1 | 101.2
4.7 | 1.7 | 2.7 | 6.25 | 23.25 | 94.025 | 100.4 | 101.1 | 101.2
4.8 | 1.6 | 2.6 | 5.9667 | 18 | 93.025 | 100.4 | 101.2 | 101.3
4.9 | 1.6 | 2.5 | 5.7 | 14.75 | 91.625 | 100.4 | 101.2 | 101.3
5 | 1.5 | 2.45 | 5.5 | 12.8 | 89.74 | 100.4 | 101.2 | 101.4
5.1 | 1.5 | 2.3 | 5.3 | 11.5 | 86.28 | 100.4 | 101.3 | 101.4
5.2 | 1.4 | 2.3 | 5.05 | 10.6 | 80.7 | 100.4 | 101.3 | 101.4714
5.3 | 1.4 | 2.2 | 4.9 | 9.8 | 72.7 | 100.4 | 101.3 | 101.5
5.4 | 1.3 | 2.1 | 4.7333 | 9.3 | 63.25 | 100.4 | 101.3 | 101.5
5.5 | 1.3 | 2.1 | 4.5 | 8.7667 | 52.4 | 100.375 | 101.3 | 101.5333
5.6 | 1.3 | 2 | 4.4 | 8.4 | 43.025 | 100.3 | 101.3 | 101.6
5.7 | 1.2 | 2 | 4.2 | 8 | 33 | 100.3 | 101.4 | 101.6
5.8 | 1.2 | 1.9 | 4.0667 | 7.7 | 25.4667 | 100.2333 | 101.4 | 101.6
5.9 | 1.2 | 1.8 | 4 | 7.4 | 19.7 | 100.2 | 101.4 | 101.6
6 | 1.1 | 1.8 | 3.8 | 7.2 | 16.275 | 100.1 | 101.4 | 101.7
6.1 | 1.1 | 1.7 | 3.6667 | 6.9 | 14.2 | 100 | 101.4 | 101.7
6.2 | 1.1 | 1.7 | 3.5 | 6.6 | 12.825 | 99.975 | 101.3857 | 101.7
6.3 | 1.1 | 1.6 | 3.4 | 6.4 | 11.825 | 99.9 | 101.3 | 101.7
6.4 | 1.1 | 1.6 | 3.3 | 6.2 | 11.22 | 99.84 | 101.3 | 101.7
6.5 | 1 | 1.5 | 3.2333 | 5.9 | 10.6 | 99.8 | 101.3 | 101.7
6.6 | 1 | 1.5 | 3.1 | 5.7333 | 10.1 | 99.7 | 101.3 | 101.7
6.7 |  | 1.4 | 3 | 5.5 | 9.6 | 99.6 | 101.3 | 101.7
6.8 |  | 1.4 | 2.9 | 5.4 | 9.2 | 99.5 | 101.3 | 101.7
6.9 |  | 1.4 | 2.8 | 5.2 | 8.9 | 99.46 | 101.3 | 101.7
7 |  | 1.4 | 2.8 | 5 | 8.62 | 99.4 | 101.3 | 101.7
7.1 |  | 1.3 | 2.7 | 4.8 | 8.26 | 99.3 | 101.3 | 101.7
7.2 |  |  | 2.6 | 4.7 | 8.025 | 99.125 | 101.3 | 101.7
7.3 |  |  | 2.5 | 4.6 | 7.7333 | 99 | 101.3 | 101.7
7.4 |  |  | 2.5 | 4.4 | 7.475 | 98.95 | 101.2667 | 101.7
7.5 |  |  | 2.4 | 4.3 | 7.25 | 98.9 | 101.2 | 101.7
7.6 |  |  | 2.4 | 4.1333 | 7 | 98.7 | 101.2 | 101.7
7.7 |  |  | 2.3 | 4 | 6.775 | 98.62 | 101.2 | 101.7
7.8 |  |  | 2.2 | 3.9 | 6.525 | 98.46 | 101.2 | 101.7
7.9 |  |  | 2.2 | 3.8 | 6.4 | 98.325 | 101.1333 | 101.7
8 |  |  | 2.1 | 3.7 | 6.2 | 98.275 | 101.1 | 101.8
8.1 |  |  | 2.1 | 3.6 | 6 | 98.1 | 101.1 | 101.8
8.2 |  |  | 2 | 3.5 | 5.74 | 97.975 | 101.1 | 101.8
8.3 |  |  | 2 | 3.45 | 5.6 | 97.8 | 101.1 | 101.8
8.4 |  |  |  | 3.3 | 5.425 | 97.6 | 101.0143 | 101.8
8.5 |  |  |  | 3.3 | 5.3 | 97.4333 | 101 | 101.8
8.6 |  |  |  | 3.2 | 5.2 | 97.2 | 101 | 101.8
8.7 |  |  |  | 3.1 | 5 | 97.1 | 100.9167 | 101.7333
8.8 |  |  |  | 3 | 4.9 | 96.8 | 100.9 | 101.7
8.9 |  |  |  | 2.9 | 4.7 | 96.6667 | 100.9 | 101.7
9 |  |  |  | 2.9 | 4.6 | 96.375 | 100.8667 | 101.7
9.1 |  |  |  | 2.8 | 4.5 | 96.125 | 100.8 | 101.7
9.2 |  |  |  | 2.8 | 4.4 | 95.875 | 100.8 | 101.8
9.3 |  |  |  | 2.7 | 4.2333 | 95.625 | 100.7714 | 101.8
9.4 |  |  |  | 2.6 | 4.1 | 95.4 | 100.7 | 101.8
9.5 |  |  |  |  | 4 | 95.025 | 100.65 | 101.8
9.6 |  |  |  |  | 3.9 | 94.7 | 100.6 | 101.8
9.7 |  |  |  |  | 3.8 | 94.4 | 100.6 | 101.7286
9.8 |  |  |  |  | 3.7 | 94 | 100.5 | 101.7
9.9 |  |  |  |  | 3.675 | 93.55 | 100.42 | 101.7
10 |  |  |  |  | 3.6 | 93.1 | 100.4 | 101.65
10.1 |  |  |  |  | 3.5 | 92.4 | 100.3667 | 101.6
10.2 |  |  |  |  | 3.4 | 91.86 | 100.3 | 101.6
10.3 |  |  |  |  | 3.375 | 90.975 | 100.2 | 101.6
10.4 |  |  |  |  | 3.3 | 89.8 | 100.2 | 101.6
10.5 | [5 empty] | 87.85 | 100.18 | 101.6
10.6 | [5 empty] | 84.8 | 100.1 | 101.6
10.7 | [5 empty] | 79.9 | 100.1 | 101.6
10.8 | [5 empty] | 73.5 | 100 | 101.6
10.9 | [5 empty] | 65 | 99.9333 | 101.6
11 | [5 empty] | 55.65 | 99.9 | 101.6
11.1 | [5 empty] | 46.1333 | 99.8 | 101.6
11.2 | [5 empty] | 37 | 99.7 | 101.6
11.3 | [5 empty] | 29.6 | 99.65 | 101.6
11.4 | [5 empty] | 23.6 | 99.6 | 101.6
11.5 | [5 empty] | 20.2 | 99.5 | 101.6
11.6 | [5 empty] | 18.15 | 99.4 | 101.5
11.7 | [5 empty] | 16.8 | 99.32 | 101.5
11.8 | [5 empty] | 15.8 | 99.2167 | 101.5
11.9 | [5 empty] | 15.05 | 99.1333 | 101.5
12 | [5 empty] | 14.4 | 99.1 | 101.44
12.1 | [5 empty] | 13.95 | 99 | 101.4
12.2 | [5 empty] | 13.45 | 98.9 | 101.4
12.3 | [5 empty] | 13.05 | 98.8286 | 101.4
12.4 | [5 empty] | 12.65 | 98.7 | 101.4
12.5 | [5 empty] | 12.28 | 98.5429 | 101.4
12.6 | [5 empty] | 12 | 98.4714 | 101.4
12.7 | [5 empty] | 11.7 | 98.3 | 101.3
12.8 | [5 empty] | 11.45 | 98.22 | 101.3
12.9 | [5 empty] | 11.1 | 98.1 | 101.3
13 | [5 empty] | 10.8 | 98 | 101.2714
13.1 | [5 empty] | 10.525 | 97.82 | 101.2
13.2 | [5 empty] | 10.275 | 97.72 | 101.2
13.3 | [5 empty] | 10.025 | 97.54 | 101.2
13.4 | [5 empty] | 9.775 | 97.44 | 101.1
13.5 | [5 empty] | 9.525 | 97.3 | 101.1
13.6 | [5 empty] | 9.28 | 97.14 | 101.02
13.7 | [5 empty] | 9.1333 | 97 | 101
13.8 | [5 empty] | 8.9 | 96.8 | 101
13.9 | [5 empty] | 8.7333 | 96.6 | 101
14 | [5 empty] | 8.46 | 96.4 | 101
14.1 | [5 empty] | 8.3 | 96.2833 | 101
14.2 | [5 empty] | 8.1 | 96.1 | 100.9
14.3 | [5 empty] | 7.975 | 95.9167 | 100.9
14.4 | [5 empty] | 7.8 | 95.65 | 100.9
14.5 | [5 empty] | 7.6 | 95.48 | 100.8
14.6 | [5 empty] | 7.45 | 95.22 | 100.8
14.7 | [5 empty] | 7.325 | 94.95 | 100.7
14.8 | [5 empty] | 7.175 | 94.725 | 100.7
14.9 | [5 empty] | 7 | 94.44 | 100.6333
15 | [5 empty] | 6.9 | 94.1667 | 100.6
15.1 | [5 empty] | 6.7 | 93.88 | 100.6
15.2 | [5 empty] | 6.575 | 93.46 | 100.5
15.3 | [5 empty] | 6.4 | 93.1 | 100.4833
15.4 | [5 empty] | 6.3 | 92.7667 | 100.4
15.5 | [5 empty] | 6.2 | 92.28 | 100.4
15.6 | [5 empty] | 6.1 | 91.7 | 100.3833
15.7 | [5 empty] | 5.96 | 91.02 | 100.3
15.8 | [5 empty] | 5.82 | 90.28 | 100.24
15.9 | [5 empty] | 5.8 | 89.0333 | 100.1667
16 | [6 empty] | 87.4 | 100.1
16.1 | [6 empty] | 84.7 | 100.1
16.2 | [6 empty] | 80.3667 | 100
16.3 | [6 empty] | 73.5 | 99.9
16.4 | [6 empty] | 65.92 | 99.9
16.5 | [6 empty] | 56.7 | 99.8571
16.6 | [6 empty] | 47.4 | 99.8
16.7 | [6 empty] | 38.7 | 99.7
16.8 | [6 empty] | 31.9 | 99.6
16.9 | [6 empty] | 26 | 99.5
17 | [6 empty] | 22.55 | 99.35
17.1 | [6 empty] | 20.3833 | 99.3
17.2 | [6 empty] | 19 | 99.22
17.3 | [6 empty] | 18.04 | 99.1333
17.4 | [6 empty] | 17.28 | 99
17.5 | [6 empty] | 16.62 | 98.9
17.6 | [6 empty] | 16.2 | 98.82
17.7 | [6 empty] | 15.7 | 98.75
17.8 | [6 empty] | 15.2 | 98.6
17.9 | [6 empty] | 14.9 | 98.5167
18 | [6 empty] | 14.5667 | 98.4
18.1 | [6 empty] | 14.1667 | 98.3
18.2 | [6 empty] | 13.9286 | 98.1857
18.3 | [6 empty] | 13.6 | 98.0333
18.4 | [6 empty] | 13.28 | 97.9
18.5 | [6 empty] | 13.025 | 97.8
18.6 | [6 empty] | 12.72 | 97.6
18.7 | [6 empty] | 12.45 | 97.4333
18.8 | [6 empty] | 12.1667 | 97.25
18.9 | [6 empty] | 11.8857 | 97.15
19 | [6 empty] | 11.7 | 97
19.1 | [6 empty] | 11.4333 | 96.8667
19.2 | [6 empty] | 11.1667 | 96.6833
19.3 | [6 empty] | 10.9667 | 96.5167
19.4 | [6 empty] | 10.8 | 96.35
19.5 | [6 empty] | 10.5 | 96.1667
19.6 | [6 empty] | 10.3 | 95.9167
19.7 | [6 empty] | 10.1 | 95.8286
19.8 | [6 empty] | 9.9 | 95.6333
19.9 | [6 empty] | 9.7429 | 95.3
20 | [6 empty] | 9.5833 | 95.1
20.1 | [6 empty] | 9.4167 | 94.8667
20.2 | [6 empty] | 9.2 | 94.6333
20.3 | [6 empty] | 9.0833 | 94.3
20.4 | [6 empty] | 8.9167 | 94.2
20.5 | [6 empty] | 8.7333 | 93.8
20.6 | [6 empty] | 8.6 | 93.54
20.7 | [6 empty] | 8.4333 | 93.2
20.8 | [6 empty] | 8.2667 | 92.7429
20.9 | [6 empty] | 8.1 | 92.4
21 | [6 empty] | 8 | 91.98
21.1 | [6 empty] | 7.8 | 91.45
21.2 | [6 empty] | 7.6833 | 90.775
21.3 | [6 empty] | 7.5167 | 89.96
21.4 | [6 empty] | 7.4 | 88.9
21.5 | [7 empty] | 87.3667
21.6 | [7 empty] | 85.1167
21.7 | [7 empty] | 81.0667
21.8 | [7 empty] | 75.25
21.9 | [7 empty] | 67.1667
22 | [7 empty] | 58.2
22.1 | [7 empty] | 49.5571
22.2 | [7 empty] | 41.12
22.3 | [7 empty] | 33.7429
22.4 | [7 empty] | 27.94
22.5 | [7 empty] | 24.26
22.6 | [7 empty] | 22
22.7 | [7 empty] | 20.55
22.8 | [7 empty] | 19.55
22.9 | [7 empty] | 18.76
23 | [7 empty] | 18.18
23.1 | [7 empty] | 17.6667
23.2 | [7 empty] | 17.2
23.3 | [7 empty] | 16.7667
23.4 | [7 empty] | 16.3333
23.5 | [7 empty] | 16
23.6 | [7 empty] | 15.68
23.7 | [7 empty] | 15.3
23.8 | [7 empty] | 15.0143
23.9 | [7 empty] | 14.7167
24 | [7 empty] | 14.4167
24.1 | [7 empty] | 14.1
24.2 | [7 empty] | 13.8
24.3 | [7 empty] | 13.5
24.4 | [7 empty] | 13.2667
24.5 | [7 empty] | 13.02
24.6 | [7 empty] | 12.7333
24.7 | [7 empty] | 12.5
24.8 | [7 empty] | 12.2333
24.9 | [7 empty] | 12.05
25 | [7 empty] | 11.7833
25.1 | [7 empty] | 11.6167
25.2 | [7 empty] | 11.35
25.3 | [7 empty] | 11.18
25.4 | [7 empty] | 10.975
25.5 | [7 empty] | 10.7667
25.6 | [7 empty] | 10.5
25.7 | [7 empty] | 10.3667
25.8 | [7 empty] | 10.2
25.9 | [7 empty] | 9.9667
26 | [7 empty] | 9.8143
26.1 | [7 empty] | 9.64
26.2 | [7 empty] | 9.48
26.3 | [7 empty] | 9.3143
26.4 | [7 empty] | 9.1167
26.5 | [7 empty] | 8.95
26.6 | [7 empty] | 8.8
26.7 | [7 empty] | 8.7
26.8 | [7 empty] | 8.5833
26.9 | [7 empty] | 8.4167

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 12.4
-28.8 | [7 empty] | 12.6286
-28.7 | [7 empty] | 12.7625
-28.6 | [7 empty] | 12.975
-28.5 | [7 empty] | 13.1571
-28.4 | [7 empty] | 13.3714
-28.3 | [7 empty] | 13.6286
-28.2 | [7 empty] | 13.7625
-28.1 | [7 empty] | 13.9556
-28 | [7 empty] | 14.1
-27.9 | [7 empty] | 14.35
-27.8 | [7 empty] | 14.6
-27.7 | [7 empty] | 14.775
-27.6 | [7 empty] | 15.0143
-27.5 | [7 empty] | 15.2143
-27.4 | [7 empty] | 15.4
-27.3 | [7 empty] | 15.6286
-27.2 | [7 empty] | 15.9143
-27.1 | [7 empty] | 16.175
-27 | [7 empty] | 16.4143
-26.9 | [7 empty] | 16.6
-26.8 | [7 empty] | 16.85
-26.7 | [7 empty] | 17.2
-26.6 | [7 empty] | 17.375
-26.5 | [7 empty] | 17.7
-26.4 | [7 empty] | 17.95
-26.3 | [7 empty] | 18.2
-26.2 | [7 empty] | 18.5
-26.1 | [7 empty] | 18.825
-26 | [7 empty] | 19.0778
-25.9 | [7 empty] | 19.4444
-25.8 | [7 empty] | 19.7125
-25.7 | [7 empty] | 20.025
-25.6 | [7 empty] | 20.45
-25.5 | [7 empty] | 20.7286
-25.4 | [7 empty] | 21.05
-25.3 | [7 empty] | 21.5
-25.2 | [7 empty] | 21.9
-25.1 | [7 empty] | 22.3
-25 | [7 empty] | 22.9143
-24.9 | [7 empty] | 23.4857
-24.8 | [7 empty] | 24.25
-24.7 | [7 empty] | 25.0429
-24.6 | [7 empty] | 26.3
-24.5 | [7 empty] | 28.3
-24.4 | [7 empty] | 31
-24.3 | [7 empty] | 35.6
-24.2 | [7 empty] | 41.1125
-24.1 | [7 empty] | 47.2143
-24 | [7 empty] | 54.0714
-23.9 | [7 empty] | 60.3556
-23.8 | [7 empty] | 66.6
-23.7 | [7 empty] | 71.3
-23.6 | [7 empty] | 74.95
-23.5 | [7 empty] | 77.4
-23.4 | [7 empty] | 78.9
-23.3 | [7 empty] | 80.025
-23.2 | [7 empty] | 80.8
-23.1 | [7 empty] | 81.35
-23 | [7 empty] | 81.9125
-22.9 | [6 empty] | 11 | 82.5375
-22.8 | [6 empty] | 11.2 | 82.9375
-22.7 | [6 empty] | 11.42 | 83.3125
-22.6 | [6 empty] | 11.525 | 83.625
-22.5 | [6 empty] | 11.7333 | 84.05
-22.4 | [6 empty] | 11.9 | 84.375
-22.3 | [6 empty] | 12.025 | 84.75
-22.2 | [6 empty] | 12.25 | 85.05
-22.1 | [6 empty] | 12.46 | 85.4
-22 | [6 empty] | 12.675 | 85.6571
-21.9 | [6 empty] | 12.85 | 85.9429
-21.8 | [6 empty] | 13.05 | 86.225
-21.7 | [6 empty] | 13.275 | 86.5125
-21.6 | [6 empty] | 13.46 | 86.825
-21.5 | [6 empty] | 13.66 | 87.075
-21.4 | [6 empty] | 13.94 | 87.35
-21.3 | [6 empty] | 14.1333 | 87.6
-21.2 | [6 empty] | 14.4 | 87.775
-21.1 | [6 empty] | 14.6333 | 88.1
-21 | [6 empty] | 14.8667 | 88.35
-20.9 | [6 empty] | 15.0667 | 88.6
-20.8 | [6 empty] | 15.3333 | 88.85
-20.7 | [6 empty] | 15.5833 | 89.1
-20.6 | [6 empty] | 15.85 | 89.325
-20.5 | [6 empty] | 16.1333 | 89.5143
-20.4 | [6 empty] | 16.4429 | 89.6875
-20.3 | [6 empty] | 16.72 | 89.9125
-20.2 | [6 empty] | 16.98 | 90.075
-20.1 | [6 empty] | 17.3333 | 90.325
-20 | [6 empty] | 17.6333 | 90.4875
-19.9 | [6 empty] | 17.9 | 90.7125
-19.8 | [6 empty] | 18.2 | 90.875
-19.7 | [6 empty] | 18.5333 | 91.1429
-19.6 | [6 empty] | 18.8 | 91.3
-19.5 | [6 empty] | 19.1857 | 91.425
-19.4 | [6 empty] | 19.6 | 91.6
-19.3 | [6 empty] | 19.9667 | 91.8714
-19.2 | [6 empty] | 20.35 | 92
-19.1 | [6 empty] | 20.78 | 92.0875
-19 | [6 empty] | 21.35 | 92.3222
-18.9 | [6 empty] | 21.9333 | 92.525
-18.8 | [6 empty] | 22.56 | 92.65
-18.7 | [6 empty] | 23.45 | 92.7
-18.6 | [6 empty] | 24.66 | 92.7889
-18.5 | [6 empty] | 26.45 | 93.025
-18.4 | [6 empty] | 29 | 93.25
-18.3 | [6 empty] | 33.2 | 93.3875
-18.2 | [6 empty] | 38.65 | 93.5111
-18.1 | [6 empty] | 45.45 | 93.625
-18 | [6 empty] | 52.5 | 93.75
-17.9 | [6 empty] | 59.6 | 93.9714
-17.8 | [6 empty] | 66.32 | 94.1
-17.7 | [6 empty] | 71.85 | 94.1625
-17.6 | [6 empty] | 76.08 | 94.375
-17.5 | [6 empty] | 78.7857 | 94.4125
-17.4 | [6 empty] | 80.5 | 94.5
-17.3 | [6 empty] | 81.7 | 94.6
-17.2 | [6 empty] | 82.58 | 94.825
-17.1 | [6 empty] | 83.32 | 94.95
-17 | [5 empty] | 9.0333 | 83.9333 | 95
-16.9 | [5 empty] | 9.2333 | 84.45 | 95.1
-16.8 | [5 empty] | 9.375 | 85 | 95.1875
-16.7 | [5 empty] | 9.56 | 85.35 | 95.3167
-16.6 | [5 empty] | 9.7 | 85.75 | 95.4833
-16.5 | [5 empty] | 9.875 | 86.08 | 95.6
-16.4 | [5 empty] | 10 | 86.5167 | 95.675
-16.3 | [5 empty] | 10.225 | 86.8 | 95.7222
-16.2 | [5 empty] | 10.4 | 87.0667 | 95.8375
-16.1 | [5 empty] | 10.6 | 87.4 | 95.9
-16 | [5 empty] | 10.72 | 87.78 | 95.9875
-15.9 | [5 empty] | 10.95 | 88.1 | 96.1
-15.8 | [5 empty] | 11.15 | 88.3667 | 96.2333
-15.7 | [5 empty] | 11.38 | 88.6 | 96.3
-15.6 | [5 empty] | 11.525 | 88.9333 | 96.3889
-15.5 | [5 empty] | 11.8 | 89.2167 | 96.5
-15.4 | [5 empty] | 12 | 89.3833 | 96.5429
-15.3 | [5 empty] | 12.2 | 89.72 | 96.6
-15.2 | [5 empty] | 12.4333 | 89.9333 | 96.7
-15.1 | [5 empty] | 12.62 | 90.1833 | 96.85
-15 | [5 empty] | 12.85 | 90.475 | 97
-14.9 | [5 empty] | 13.1 | 90.6429 | 97.025
-14.8 | [5 empty] | 13.3 | 90.9167 | 97.1571
-14.7 | [5 empty] | 13.65 | 91.0833 | 97.2
-14.6 | [5 empty] | 13.9 | 91.28 | 97.3
-14.5 | [5 empty] | 14.175 | 91.4667 | 97.3
-14.4 | [5 empty] | 14.375 | 91.7333 | 97.375
-14.3 | [5 empty] | 14.68 | 92 | 97.5
-14.2 | [5 empty] | 15 | 92.1 | 97.525
-14.1 | [5 empty] | 15.225 | 92.2333 | 97.6
-14 | [5 empty] | 15.65 | 92.4833 | 97.6714
-13.9 | [5 empty] | 15.9 | 92.66 | 97.7
-13.8 | [5 empty] | 16.25 | 92.8429 | 97.8
-13.7 | [5 empty] | 16.62 | 93.02 | 97.8625
-13.6 | [5 empty] | 16.9 | 93.3 | 98
-13.5 | [5 empty] | 17.3 | 93.4 | 98
-13.4 | [5 empty] | 17.68 | 93.55 | 98.075
-13.3 | [5 empty] | 18.1 | 93.74 | 98.1
-13.2 | [5 empty] | 18.4667 | 93.9167 | 98.15
-13.1 | [5 empty] | 18.9667 | 94.0833 | 98.275
-13 | [5 empty] | 19.525 | 94.2667 | 98.3
-12.9 | [5 empty] | 20.075 | 94.4 | 98.4
-12.8 | [5 empty] | 20.825 | 94.5 | 98.5
-12.7 | [5 empty] | 21.9 | 94.6667 | 98.5
-12.6 | [5 empty] | 23.225 | 94.825 | 98.5375
-12.5 | [5 empty] | 25.1667 | 95 | 98.6
-12.4 | [5 empty] | 28.54 | 95.0833 | 98.6778
-12.3 | [5 empty] | 32.85 | 95.18 | 98.7
-12.2 | [5 empty] | 39.4 | 95.38 | 98.8
-12.1 | [5 empty] | 46.7 | 95.58 | 98.8444
-12 | [5 empty] | 54.6 | 95.6667 | 98.9
-11.9 | [5 empty] | 62.8 | 95.72 | 99
-11.8 | [5 empty] | 69.8 | 95.8 | 99
-11.7 | [5 empty] | 75.24 | 95.98 | 99.0778
-11.6 | [5 empty] | 79.7333 | 96.08 | 99.1
-11.5 | [5 empty] | 82.475 | 96.2429 | 99.1
-11.4 | [5 empty] | 84.25 | 96.4 | 99.2
-11.3 | [5 empty] | 85.5 | 96.5 | 99.2
-11.2 | [5 empty] | 86.475 | 96.6 | 99.2889
-11.1 | [5 empty] | 87.1333 | 96.6571 | 99.3
-11 | [5 empty] | 87.6 | 96.8 | 99.3625
-10.9 |  |  |  |  | 5.7 | 88.3 | 96.86 | 99.4
-10.8 |  |  |  |  | 5.8 | 88.7 | 97 | 99.5
-10.7 |  |  |  |  | 5.9 | 89.15 | 97.08 | 99.5
-10.6 |  |  |  |  | 6 | 89.6 | 97.2 | 99.5
-10.5 |  |  |  |  | 6.175 | 89.95 | 97.3 | 99.5
-10.4 |  |  |  |  | 6.3 | 90.4 | 97.4 | 99.625
-10.3 |  |  |  |  | 6.4 | 90.68 | 97.4667 | 99.7
-10.2 |  |  |  |  | 6.6 | 91.1 | 97.5167 | 99.7
-10.1 |  |  |  |  | 6.725 | 91.38 | 97.6833 | 99.8
-10 |  |  |  |  | 6.875 | 91.75 | 97.7429 | 99.8
-9.9 |  |  |  |  | 7 | 92.05 | 97.9 | 99.8625
-9.8 |  |  |  |  | 7.15 | 92.3 | 98 | 99.9
-9.7 |  |  |  | 4.7 | 7.325 | 92.575 | 98.0714 | 99.9125
-9.6 |  |  |  | 4.74 | 7.45 | 92.84 | 98.1333 | 100
-9.5 |  |  |  | 4.9 | 7.6 | 93.1 | 98.2 | 100
-9.4 |  |  |  | 5 | 7.85 | 93.2667 | 98.3 | 100
-9.3 |  |  |  | 5.1 | 8 | 93.65 | 98.35 | 100.1
-9.2 |  |  |  | 5.2667 | 8.2 | 93.9 | 98.4 | 100.1
-9.1 |  |  |  | 5.4 | 8.3 | 94.075 | 98.5 | 100.1833
-9 |  |  |  | 5.4 | 8.6 | 94.2333 | 98.6 | 100.2
-8.9 |  |  |  | 5.5 | 8.74 | 94.525 | 98.6 | 100.2
-8.8 |  |  |  | 5.7 | 8.975 | 94.725 | 98.7 | 100.2
-8.7 |  |  |  | 5.8 | 9.2 | 94.9667 | 98.8667 | 100.2
-8.6 |  |  |  | 5.9 | 9.44 | 95.1 | 98.9 | 100.3
-8.5 |  |  | 3.5 | 6.0667 | 9.675 | 95.25 | 98.9167 | 100.3
-8.4 |  |  | 3.6 | 6.2 | 9.9 | 95.5 | 99 | 100.3
-8.3 |  |  | 3.7 | 6.4 | 10.15 | 95.65 | 99.05 | 100.3
-8.2 |  |  | 3.8 | 6.6 | 10.44 | 95.8 | 99.1 | 100.4
-8.1 |  |  | 3.8 | 6.75 | 10.72 | 95.925 | 99.125 | 100.4
-8 |  |  | 3.9667 | 6.85 | 11 | 96.125 | 99.2 | 100.4778
-7.9 |  |  | 4 | 7.1 | 11.25 | 96.375 | 99.3 | 100.5
-7.8 |  |  | 4.1 | 7.225 | 11.55 | 96.6 | 99.3 | 100.525
-7.7 |  |  | 4.2 | 7.4 | 11.86 | 96.7 | 99.38 | 100.6
-7.6 |  |  | 4.3 | 7.625 | 12.2 | 96.8 | 99.4 | 100.6
-7.5 |  |  | 4.5 | 7.775 | 12.5667 | 96.85 | 99.4 | 100.6
-7.4 |  | 2.2 | 4.6 | 8 | 12.9 | 97.1 | 99.6 | 100.6
-7.3 |  | 2.3667 | 4.7 | 8.2 | 13.2 | 97.25 | 99.6 | 100.6
-7.2 |  | 2.4 | 4.8 | 8.5 | 13.65 | 97.4 | 99.7 | 100.6
-7.1 |  | 2.5 | 4.9 | 8.64 | 14.1 | 97.5 | 99.7 | 100.6
-7 |  | 2.5 | 5 | 8.9 | 14.6 | 97.675 | 99.75 | 100.6
-6.9 |  | 2.6 | 5.2 | 9.2 | 15.3 | 97.825 | 99.8 | 100.6
-6.8 |  | 2.6 | 5.3 | 9.35 | 16.1 | 97.9 | 99.8 | 100.6
-6.7 | 1.7 | 2.7 | 5.5 | 9.68 | 17.14 | 98 | 99.9 | 100.6
-6.6 | 1.8 | 2.8 | 5.6 | 9.9 | 18.575 | 98.1 | 99.9 | 100.6
-6.5 | 1.8 | 2.8 | 5.8 | 10.2 | 20.7333 | 98.2 | 99.9 | 100.6
-6.4 | 1.9 | 2.9 | 6 | 10.5 | 23.825 | 98.3 | 100 | 100.6556
-6.3 | 1.9 | 2.9 | 6.1 | 10.86 | 28.825 | 98.5 | 100 | 100.7
-6.2 | 1.9 | 3 | 6.2 | 11.225 | 35.9 | 98.6 | 100 | 100.7
-6.1 | 2 | 3.1 | 6.4 | 11.56 | 44.7 | 98.625 | 100 | 100.75
-6 | 2 | 3.2 | 6.6 | 12.025 | 53.15 | 98.7 | 100 | 100.8
-5.9 | 2.1 | 3.25 | 6.8 | 12.45 | 62.95 | 98.72 | 100.0167 | 100.8
-5.8 | 2.1 | 3.4 | 7 | 12.9333 | 71.55 | 98.8 | 100.1 | 100.7
-5.7 | 2.2 | 3.5 | 7.3 | 13.48 | 78.425 | 98.9 | 100.1 | 100.7
-5.6 | 2.2 | 3.5 | 7.5 | 14.3 | 83.825 | 98.9 | 100.1 | 100.7
-5.5 | 2.2667 | 3.6 | 7.7 | 15.4 | 87.225 | 99 | 100.1 | 100.7
-5.4 | 2.35 | 3.75 | 7.95 | 16.65 | 89.4 | 99.04 | 100.1 | 100.7
-5.3 | 2.4 | 3.9 | 8.1333 | 18.8 | 90.8 | 99.2 | 100.2 | 100.7
-5.2 | 2.5 | 4 | 8.4667 | 21.975 | 91.8 | 99.2 | 100.2 | 100.7
-5.1 | 2.5 | 4.1 | 8.7 | 27.025 | 92.6 | 99.3 | 100.2 | 100.7
-5 | 2.6 | 4.2 | 8.9 | 34.35 | 93.16 | 99.3 | 100.1857 | 100.7
-4.9 | 2.7 | 4.35 | 9.3 | 42.8 | 93.65 | 99.3 | 100.1 | 100.6625
-4.8 | 2.8 | 4.45 | 9.6667 | 51.9 | 94.225 | 99.3 | 100.1 | 100.6
-4.7 | 2.8 | 4.6 | 10.1 | 61.44 | 94.46 | 99.3 | 100.1 | 100.6
-4.6 | 2.9 | 4.8 | 10.5 | 70.72 | 94.72 | 99.4 | 100.1 | 100.5571
-4.5 | 2.9333 | 5 | 11.1333 | 78 | 95.1 | 99.4 | 100.2 | 100.5
-4.4 | 3 | 5.1 | 11.9 | 84.1 | 95.4 | 99.4 | 100.2 | 100.5
-4.3 | 3.2 | 5.2333 | 12.9333 | 87.9 | 95.7 | 99.5 | 100.2 | 100.5
-4.2 | 3.3 | 5.4 | 14.25 | 90.35 | 96 | 99.5 | 100.1 | 100.5
-4.1 | 3.4 | 5.6 | 16.3 | 91.8 | 96.25 | 99.6 | 100.1 | 100.5
-4 | 3.5 | 5.75 | 19.5 | 92.875 | 96.475 | 99.6 | 100.1 | 100.5
-3.9 | 3.6 | 6.1 | 24.55 | 93.8 | 96.625 | 99.625 | 100.1 | 100.5
-3.8 | 3.75 | 6.25 | 31.1667 | 94.45 | 96.9 | 99.7 | 100.1 | 100.5
-3.7 | 3.9 | 6.55 | 40.1 | 94.875 | 97.2 | 99.7 | 100.1 | 100.5
-3.6 | 4 | 6.9 | 50.8333 | 95.3333 | 97.35 | 99.7667 | 100.2 | 100.5
-3.5 | 4.1 | 7.2 | 60.7 | 95.7 | 97.6 | 99.8 | 100.2 | 100.5
-3.4 | 4.3 | 7.5667 | 70.5 | 96.1 | 97.75 | 99.85 | 100.2 | 100.5
-3.3 | 4.5 | 8.1667 | 78.7 | 96.4 | 98 | 99.9 | 100.2 | 100.5
-3.2 | 4.7 | 8.8667 | 85.1 | 96.7 | 98.125 | 100 | 100.2 | 100.5
-3.1 | 4.8 | 9.7 | 89.35 | 96.9 | 98.275 | 100 | 100.2 | 100.5
-3 | 5 | 11 | 92 | 97.2333 | 98.525 | 100.1 | 100.2 | 100.5
-2.9 | 5.3 | 13.1 | 93.7 | 97.525 | 98.65 | 100.1 | 100.2 | 100.525
-2.8 | 5.7667 | 16.3 | 94.95 | 97.75 | 98.9 | 100.1 | 100.2 | 100.55
-2.7 | 6.2 | 21.8 | 95.85 | 98 | 99 | 100.1 | 100.3 | 100.5
-2.6 | 6.9667 | 28.85 | 96.3667 | 98.1 | 99.14 | 100.1 | 100.3 | 100.5
-2.5 | 7.8 | 38.5 | 97 | 98.3 | 99.2 | 100.1 | 100.3 | 100.5
-2.4 | 9.1 | 49.45 | 97.4 | 98.55 | 99.3 | 100.2 | 100.3 | 100.5
-2.3 | 11.2 | 60.7 | 97.7 | 98.7 | 99.4 | 100.2 | 100.2667 | 100.5
-2.2 | 14.55 | 71 | 97.9 | 98.925 | 99.5 | 100.2 | 100.2 | 100.5
-2.1 | 19.9 | 80.3 | 98.2 | 99 | 99.6 | 100.2 | 100.25 | 100.5
-2 | 27.65 | 87.45 | 98.4 | 99.175 | 99.6 | 100.2 | 100.3 | 100.475
-1.9 | 37.6 | 91.5667 | 98.5 | 99.3 | 99.7 | 100.1 | 100.3 | 100.4
-1.8 | 48.3 | 94.2 | 98.8 | 99.4 | 99.7 | 100.1 | 100.3 | 100.4
-1.7 | 60.2 | 96 | 98.9 | 99.5 | 99.85 | 100.1 | 100.3 | 100.4
-1.6 | 70.5 | 97 | 99.05 | 99.5 | 99.9 | 100.1 | 100.2 | 100.325
-1.5 | 80.4 | 97.8 | 99.2 | 99.6 | 99.9 | 100.1 | 100.2 | 100.3
-1.4 | 87.2 | 98.35 | 99.3 | 99.68 | 99.9 | 100.1 | 100.2 | 100.3
-1.3 | 92.2 | 98.85 | 99.4 | 99.7 | 100 | 100.1 | 100.1 | 100.3
-1.2 | 94.8 | 99.05 | 99.5333 | 99.7 | 100 | 100.1 | 100.1 | 100.2143
-1.1 | 96.55 | 99.2 | 99.7 | 99.8 | 100 | 100.1 | 100.1 | 100.2
-1 | 97.9 | 99.4 | 99.7 | 99.8 | 100 | 100.1 | 100.1 | 100.2
-0.9 | 98.45 | 99.45 | 99.8 | 99.9 | 100 | 100.1 | 100.1 | 100.1375
-0.8 | 99 | 99.6 | 99.8 | 99.95 | 100 | 100.2 | 100.1 | 100.1
-0.7 | 99.3 | 99.7 | 99.9 | 100 | 100 | 100.1333 | 100.1 | 100.1
-0.6 | 99.5 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1
-0.5 | 99.7 | 99.85 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1
-0.4 | 99.7667 | 99.9 | 100 | 100 | 100 | 100.075 | 100 | 100.1
-0.3 | 99.9 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100.1
-0.2 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100.1
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0125
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0125
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100.1
0.3 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100 | 100 | 100.1
0.4 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1
0.5 | 99.6667 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1
0.6 | 99.5 | 99.7 | 99.8 | 100 | 100 | 100.1 | 100 | 100.1
0.7 | 99.2 | 99.7 | 99.8 | 99.9 | 100 | 100.1 | 100.1 | 100.1
0.8 | 98.85 | 99.6 | 99.7 | 99.9 | 100 | 100.1 | 100.1 | 100.1
0.9 | 98.3 | 99.4 | 99.7 | 99.8 | 100 | 100.1 | 100.1 | 100.1375
1 | 97.5667 | 99.3 | 99.65 | 99.8 | 99.98 | 100.1 | 100.1 | 100.2
1.1 | 96.45 | 99.1 | 99.6 | 99.8 | 99.9 | 100.1 | 100.1 | 100.2
1.2 | 94.7 | 98.95 | 99.5 | 99.775 | 99.9 | 100.1 | 100.1 | 100.2143
1.3 | 92 | 98.7 | 99.5 | 99.7 | 99.9 | 100.12 | 100.1 | 100.3
1.4 | 87.1 | 98.1 | 99.4 | 99.675 | 99.9 | 100.2 | 100.2 | 100.3
1.5 | 80.4 | 97.6 | 99.2 | 99.6 | 99.825 | 100.2 | 100.2 | 100.3
1.6 | 70.65 | 96.8 | 99 | 99.5 | 99.8 | 100.2 | 100.2 | 100.325
1.7 | 59.8 | 95.8 | 98.9 | 99.4 | 99.74 | 100.2 | 100.3 | 100.4
1.8 | 48.4 | 94 | 98.6667 | 99.4 | 99.7 | 100.2 | 100.3 | 100.4
1.9 | 37.7 | 91.4333 | 98.5 | 99.2 | 99.65 | 100.2 | 100.3 | 100.4
2 | 27.8 | 87.15 | 98.3 | 99.175 | 99.6 | 100.2 | 100.3 | 100.475
2.1 | 19.9 | 80.25 | 98.1 | 98.9 | 99.55 | 100.2 | 100.2 | 100.5
2.2 | 14.4667 | 70.3 | 97.85 | 98.825 | 99.4 | 100.2 | 100.2 | 100.5
2.3 | 11.1 | 60.65 | 97.7 | 98.6 | 99.325 | 100.2 | 100.2429 | 100.5
2.4 | 9.05 | 49.45 | 97.3 | 98.5 | 99.275 | 100.2 | 100.3 | 100.5
2.5 | 7.8 | 38.4667 | 96.95 | 98.28 | 99.1 | 100.1 | 100.3 | 100.5
2.6 | 6.9333 | 28.85 | 96.3667 | 98.15 | 99 | 100.1 | 100.3 | 100.5
2.7 | 6.3 | 21.7 | 95.7 | 98 | 98.9 | 100.1 | 100.3 | 100.5
2.8 | 5.8 | 16.2 | 94.7 | 97.72 | 98.8 | 100 | 100.2167 | 100.55
2.9 | 5.35 | 13.1 | 93.5 | 97.425 | 98.6 | 100 | 100.2 | 100.525
3 | 5.1 | 11.1 | 91.9 | 97.175 | 98.46 | 99.9 | 100.2 | 100.5
3.1 | 4.9 | 9.7 | 89.25 | 96.95 | 98.2 | 99.9 | 100.2 | 100.5
3.2 | 4.7 | 8.7667 | 85.3 | 96.6 | 97.975 | 99.825 | 100.2 | 100.5
3.3 | 4.5 | 8.15 | 78.7 | 96.34 | 97.84 | 99.8 | 100.2 | 100.5
3.4 | 4.3 | 7.5667 | 70.5 | 96 | 97.7 | 99.7 | 100.2 | 100.5
3.5 | 4.2 | 7.1 | 60.6 | 95.7 | 97.5 | 99.7 | 100.2 | 100.5
3.6 | 4 | 6.9 | 50.3 | 95.3 | 97.25 | 99.6 | 100.2 | 100.5
3.7 | 3.9 | 6.55 | 40.3 | 94.85 | 97 | 99.6 | 100.1167 | 100.5
3.8 | 3.7 | 6.3 | 31.3 | 94.35 | 96.76 | 99.5 | 100.1 | 100.5
3.9 | 3.6 | 6.1 | 24.4 | 93.675 | 96.425 | 99.5 | 100.1 | 100.5
4 | 3.5 | 5.85 | 19.5 | 92.8 | 96.275 | 99.425 | 100.1 | 100.5
4.1 | 3.4 | 5.5667 | 16.25 | 91.65 | 96.025 | 99.375 | 100.1 | 100.5
4.2 | 3.3 | 5.5 | 14.2667 | 90.2 | 95.9 | 99.3 | 100.1 | 100.5
4.3 | 3.2 | 5.3 | 12.9 | 87.7333 | 95.6 | 99.3 | 100.1857 | 100.5
4.4 | 3.1 | 5.0333 | 11.9 | 84.1 | 95.2 | 99.3 | 100.2 | 100.5
4.5 | 3 | 4.9 | 11.15 | 78.2 | 94.95 | 99.3 | 100.2 | 100.5
4.6 | 2.9 | 4.8 | 10.55 | 71 | 94.65 | 99.3 | 100.1167 | 100.5571
4.7 | 2.8 | 4.6 | 10.1 | 61.45 | 94.2 | 99.2 | 100.1 | 100.6
4.8 | 2.7 | 4.5 | 9.7 | 52.2 | 94.025 | 99.1 | 100.1 | 100.6
4.9 | 2.7 | 4.3 | 9.3667 | 42.55 | 93.4667 | 99.1 | 100.1 | 100.6625
5 | 2.6 | 4.2 | 9 | 34.2 | 92.9667 | 99.1 | 100.1667 | 100.7
5.1 | 2.5 | 4.05 | 8.7 | 27.325 | 92.4 | 99.025 | 100.2 | 100.7
5.2 | 2.4 | 4 | 8.4667 | 22.075 | 91.6 | 99.1 | 100.2 | 100.7
5.3 | 2.4 | 3.9 | 8.2 | 18.75 | 90.6 | 99 | 100.2 | 100.7
5.4 | 2.3 | 3.7333 | 8 | 16.725 | 89.05 | 99 | 100.1 | 100.7
5.5 | 2.2 | 3.7 | 7.7 | 15.34 | 87.02 | 98.9 | 100.1 | 100.7
5.6 | 2.2 | 3.5 | 7.4667 | 14.32 | 83.3 | 98.7333 | 100.1 | 100.7
5.7 | 2.2 | 3.4667 | 7.2 | 13.5 | 78.125 | 98.7 | 100.1 | 100.7
5.8 | 2.1 | 3.4 | 7.05 | 12.9333 | 71 | 98.6 | 100.1 | 100.7
5.9 | 2.1 | 3.3 | 6.8 | 12.4 | 62.525 | 98.6 | 100.0429 | 100.8
6 | 2 | 3.2 | 6.6 | 12 | 53.35 | 98.5 | 100 | 100.8
6.1 | 2 | 3.1 | 6.4 | 11.65 | 44.1667 | 98.42 | 100 | 100.75
6.2 | 2 | 3.1 | 6.2 | 11.24 | 35.8667 | 98.3 | 100 | 100.7
6.3 | 1.9 | 3 | 6.1 | 10.875 | 28.95 | 98.2 | 100 | 100.7
6.4 | 1.9 | 2.9 | 6 | 10.55 | 24.025 | 98.0333 | 99.9667 | 100.6556
6.5 | 1.8 | 2.8 | 5.9 | 10.3 | 20.7333 | 98 | 99.9 | 100.6
6.6 | 1.8 | 2.7 | 5.6667 | 10 | 18.575 | 97.9 | 99.9 | 100.6
6.7 | 1.7 | 2.7 | 5.4 | 9.7 | 17.125 | 97.76 | 99.8667 | 100.6
6.8 |  | 2.6 | 5.3 | 9.3667 | 16.1 | 97.7 | 99.8 | 100.6
6.9 |  | 2.6 | 5.2 | 9.15 | 15.35 | 97.5 | 99.8 | 100.6
7 |  | 2.5 | 5.05 | 8.95 | 14.7 | 97.4 | 99.7 | 100.6
7.1 |  | 2.45 | 4.9 | 8.76 | 14.2 | 97.3 | 99.7 | 100.6
7.2 |  | 2.4 | 4.8 | 8.55 | 13.7 | 97.2 | 99.7 | 100.6
7.3 |  | 2.35 | 4.7 | 8.26 | 13.3333 | 97.0333 | 99.6 | 100.6
7.4 |  | 2.2 | 4.6 | 8 | 12.95 | 96.95 | 99.6 | 100.6
7.5 |  |  | 4.5 | 7.825 | 12.6 | 96.8 | 99.4 | 100.6
7.6 |  |  | 4.3 | 7.66 | 12.2333 | 96.56 | 99.4 | 100.6
7.7 |  |  | 4.2 | 7.5 | 11.9 | 96.42 | 99.3 | 100.6
7.8 |  |  | 4.1 | 7.275 | 11.575 | 96.275 | 99.3 | 100.525
7.9 |  |  | 4 | 7.1 | 11.3 | 96.2 | 99.3 | 100.5
8 |  |  | 3.9 | 6.9 | 10.96 | 96 | 99.2 | 100.4778
8.1 |  |  | 3.9 | 6.8 | 10.7 | 95.7333 | 99.16 | 100.4
8.2 |  |  | 3.8 | 6.575 | 10.46 | 95.5 | 99.1 | 100.4
8.3 |  |  | 3.7 | 6.4333 | 10.2 | 95.36 | 99 | 100.3
8.4 |  |  | 3.6 | 6.2 | 9.95 | 95.2 | 99 | 100.3
8.5 |  |  | 3.5333 | 6.1 | 9.725 | 95.1 | 98.9 | 100.3
8.6 |  |  |  | 6.02 | 9.44 | 94.875 | 98.9 | 100.3
8.7 |  |  |  | 5.8667 | 9.2 | 94.64 | 98.8 | 100.2
8.8 |  |  |  | 5.7 | 9 | 94.5 | 98.72 | 100.2
8.9 |  |  |  | 5.6 | 8.725 | 94.3 | 98.6333 | 100.2
9 |  |  |  | 5.525 | 8.6 | 94.1 | 98.54 | 100.2
9.1 |  |  |  | 5.4 | 8.4 | 93.8 | 98.45 | 100.1833
9.2 |  |  |  | 5.225 | 8.22 | 93.55 | 98.4 | 100.1
9.3 |  |  |  | 5.1 | 8 | 93.3 | 98.3833 | 100.1
9.4 |  |  |  | 5.0667 | 7.8667 | 93.075 | 98.3 | 100
9.5 |  |  |  | 4.975 | 7.6 | 92.9 | 98.2 | 100
9.6 |  |  |  | 4.9 | 7.45 | 92.55 | 98.1 | 100
9.7 |  |  |  | 4.8 | 7.32 | 92.275 | 98 | 99.9125
9.8 |  |  |  |  | 7.2 | 92.025 | 98 | 99.9
9.9 |  |  |  |  | 7.05 | 91.68 | 97.9 | 99.8625
10 |  |  |  |  | 6.82 | 91.4 | 97.7714 | 99.8
10.1 |  |  |  |  | 6.7 | 91.15 | 97.6 | 99.8
10.2 |  |  |  |  | 6.6 | 90.7 | 97.55 | 99.7
10.3 |  |  |  |  | 6.45 | 90.4 | 97.4857 | 99.7
10.4 |  |  |  |  | 6.3 | 90.1 | 97.4 | 99.625
10.5 |  |  |  |  | 6.16 | 89.72 | 97.3 | 99.5
10.6 |  |  |  |  | 6 | 89.32 | 97.2 | 99.5
10.7 |  |  |  |  | 5.9 | 88.925 | 97.1 | 99.5
10.8 |  |  |  |  | 5.8 | 88.475 | 97 | 99.5
10.9 |  |  |  |  | 5.65 | 88.05 | 96.8833 | 99.4
11 | [5 empty] | 87.44 | 96.8 | 99.3625
11.1 | [5 empty] | 86.8 | 96.6833 | 99.3
11.2 | [5 empty] | 86.1 | 96.6 | 99.2889
11.3 | [5 empty] | 85.16 | 96.44 | 99.2
11.4 | [5 empty] | 84 | 96.35 | 99.2
11.5 | [5 empty] | 82.2 | 96.2714 | 99.1
11.6 | [5 empty] | 79.5667 | 96.1 | 99.1
11.7 | [5 empty] | 75.1 | 96 | 99.0778
11.8 | [5 empty] | 69.35 | 95.8167 | 99
11.9 | [5 empty] | 62.52 | 95.75 | 99
12 | [5 empty] | 54.8 | 95.7 | 98.9
12.1 | [5 empty] | 46.6 | 95.6 | 98.8444
12.2 | [5 empty] | 39.3 | 95.4167 | 98.8
12.3 | [5 empty] | 33.25 | 95.2 | 98.7
12.4 | [5 empty] | 28.3 | 95.1 | 98.6778
12.5 | [5 empty] | 25.15 | 94.95 | 98.6
12.6 | [5 empty] | 23.15 | 94.86 | 98.5375
12.7 | [5 empty] | 21.825 | 94.6857 | 98.5
12.8 | [5 empty] | 20.9333 | 94.5333 | 98.5
12.9 | [5 empty] | 20 | 94.3667 | 98.4
13 | [5 empty] | 19.4333 | 94.3 | 98.3
13.1 | [5 empty] | 18.9 | 94.1 | 98.275
13.2 | [5 empty] | 18.4 | 93.95 | 98.15
13.3 | [5 empty] | 18 | 93.7667 | 98.1
13.4 | [5 empty] | 17.6 | 93.5833 | 98.075
13.5 | [5 empty] | 17.2 | 93.4286 | 98
13.6 | [5 empty] | 16.92 | 93.2333 | 98
13.7 | [5 empty] | 16.525 | 93.04 | 97.8625
13.8 | [5 empty] | 16.25 | 92.8 | 97.8
13.9 | [5 empty] | 15.925 | 92.7 | 97.7
14 | [5 empty] | 15.6 | 92.4 | 97.6714
14.1 | [5 empty] | 15.28 | 92.2429 | 97.6
14.2 | [5 empty] | 14.9 | 92.1 | 97.525
14.3 | [5 empty] | 14.7 | 91.9333 | 97.5
14.4 | [5 empty] | 14.45 | 91.72 | 97.375
14.5 | [5 empty] | 14.1 | 91.5 | 97.3
14.6 | [5 empty] | 13.86 | 91.3 | 97.3
14.7 | [5 empty] | 13.6 | 91.04 | 97.2
14.8 | [5 empty] | 13.325 | 90.8429 | 97.1571
14.9 | [5 empty] | 13.05 | 90.6667 | 97.025
15 | [5 empty] | 12.775 | 90.5 | 97
15.1 | [5 empty] | 12.525 | 90.2167 | 96.85
15.2 | [5 empty] | 12.375 | 89.9429 | 96.7
15.3 | [5 empty] | 12.125 | 89.6667 | 96.6
15.4 | [5 empty] | 11.875 | 89.4167 | 96.5429
15.5 | [5 empty] | 11.7 | 89.2 | 96.5
15.6 | [5 empty] | 11.5 | 88.8833 | 96.3889
15.7 | [5 empty] | 11.26 | 88.6571 | 96.3
15.8 | [5 empty] | 11.1 | 88.36 | 96.2333
15.9 | [5 empty] | 10.9 | 88.16 | 96.1
16 | [5 empty] | 10.6333 | 87.7667 | 95.9875
16.1 | [5 empty] | 10.5 | 87.46 | 95.9
16.2 | [5 empty] | 10.325 | 87.1 | 95.8375
16.3 | [5 empty] | 10.075 | 86.74 | 95.7222
16.4 | [5 empty] | 10 | 86.5 | 95.675
16.5 | [5 empty] | 9.825 | 86.16 | 95.6
16.6 | [5 empty] | 9.65 | 85.7667 | 95.4833
16.7 | [5 empty] | 9.425 | 85.35 | 95.3167
16.8 | [5 empty] | 9.3 | 84.9333 | 95.1875
16.9 | [5 empty] | 9.2 | 84.5 | 95.1
17 | [5 empty] | 9.1 | 83.9286 | 95
17.1 | [6 empty] | 83.34 | 94.95
17.2 | [6 empty] | 82.6 | 94.825
17.3 | [6 empty] | 81.7 | 94.6
17.4 | [6 empty] | 80.6 | 94.5
17.5 | [6 empty] | 78.6667 | 94.4125
17.6 | [6 empty] | 75.94 | 94.375
17.7 | [6 empty] | 71.94 | 94.1625
17.8 | [6 empty] | 66.58 | 94.1
17.9 | [6 empty] | 59.8167 | 93.9714
18 | [6 empty] | 52.4667 | 93.75
18.1 | [6 empty] | 45.28 | 93.625
18.2 | [6 empty] | 38.72 | 93.5111
18.3 | [6 empty] | 33.3 | 93.3875
18.4 | [6 empty] | 29.1 | 93.25
18.5 | [6 empty] | 26.48 | 93.025
18.6 | [6 empty] | 24.6 | 92.7889
18.7 | [6 empty] | 23.5 | 92.7
18.8 | [6 empty] | 22.5667 | 92.65
18.9 | [6 empty] | 21.9571 | 92.525
19 | [6 empty] | 21.3 | 92.3222
19.1 | [6 empty] | 20.775 | 92.0875
19.2 | [6 empty] | 20.3667 | 92
19.3 | [6 empty] | 19.95 | 91.8714
19.4 | [6 empty] | 19.55 | 91.6
19.5 | [6 empty] | 19.2 | 91.425
19.6 | [6 empty] | 18.8667 | 91.3
19.7 | [6 empty] | 18.5 | 91.1429
19.8 | [6 empty] | 18.1571 | 90.875
19.9 | [6 empty] | 17.88 | 90.7125
20 | [6 empty] | 17.56 | 90.4875
20.1 | [6 empty] | 17.3333 | 90.325
20.2 | [6 empty] | 17.0333 | 90.075
20.3 | [6 empty] | 16.7 | 89.9125
20.4 | [6 empty] | 16.4 | 89.6875
20.5 | [6 empty] | 16.2 | 89.5143
20.6 | [6 empty] | 15.8667 | 89.325
20.7 | [6 empty] | 15.6167 | 89.1
20.8 | [6 empty] | 15.35 | 88.85
20.9 | [6 empty] | 15.06 | 88.6
21 | [6 empty] | 14.9 | 88.35
21.1 | [6 empty] | 14.6 | 88.1
21.2 | [6 empty] | 14.42 | 87.775
21.3 | [6 empty] | 14.1667 | 87.6
21.4 | [6 empty] | 13.8667 | 87.35
21.5 | [6 empty] | 13.68 | 87.075
21.6 | [6 empty] | 13.48 | 86.825
21.7 | [6 empty] | 13.3 | 86.5125
21.8 | [6 empty] | 13.08 | 86.225
21.9 | [6 empty] | 12.8667 | 85.9429
22 | [6 empty] | 12.6 | 85.6571
22.1 | [6 empty] | 12.38 | 85.4
22.2 | [6 empty] | 12.18 | 85.05
22.3 | [6 empty] | 12.06 | 84.75
22.4 | [6 empty] | 11.9333 | 84.375
22.5 | [6 empty] | 11.72 | 84.05
22.6 | [6 empty] | 11.55 | 83.625
22.7 | [6 empty] | 11.35 | 83.3125
22.8 | [6 empty] | 11.2333 | 82.9375
22.9 | [6 empty] | 11 | 82.5375
23 | [7 empty] | 81.9125
23.1 | [7 empty] | 81.35
23.2 | [7 empty] | 80.8
23.3 | [7 empty] | 80.025
23.4 | [7 empty] | 78.9
23.5 | [7 empty] | 77.4
23.6 | [7 empty] | 74.95
23.7 | [7 empty] | 71.3
23.8 | [7 empty] | 66.6
23.9 | [7 empty] | 60.3556
24 | [7 empty] | 54.0714
24.1 | [7 empty] | 47.2143
24.2 | [7 empty] | 41.1125
24.3 | [7 empty] | 35.6
24.4 | [7 empty] | 31
24.5 | [7 empty] | 28.3
24.6 | [7 empty] | 26.3
24.7 | [7 empty] | 25.0429
24.8 | [7 empty] | 24.25
24.9 | [7 empty] | 23.4857
25 | [7 empty] | 22.9143
25.1 | [7 empty] | 22.3
25.2 | [7 empty] | 21.9
25.3 | [7 empty] | 21.5
25.4 | [7 empty] | 21.05
25.5 | [7 empty] | 20.7286
25.6 | [7 empty] | 20.45
25.7 | [7 empty] | 20.025
25.8 | [7 empty] | 19.7125
25.9 | [7 empty] | 19.4444
26 | [7 empty] | 19.0778
26.1 | [7 empty] | 18.825
26.2 | [7 empty] | 18.5
26.3 | [7 empty] | 18.2
26.4 | [7 empty] | 17.95
26.5 | [7 empty] | 17.7
26.6 | [7 empty] | 17.375
26.7 | [7 empty] | 17.2
26.8 | [7 empty] | 16.85
26.9 | [7 empty] | 16.6
27 | [7 empty] | 16.4143
27.1 | [7 empty] | 16.175
27.2 | [7 empty] | 15.9143
27.3 | [7 empty] | 15.6286
27.4 | [7 empty] | 15.4
27.5 | [7 empty] | 15.2143
27.6 | [7 empty] | 15.0143
27.7 | [7 empty] | 14.775
27.8 | [7 empty] | 14.6
27.9 | [7 empty] | 14.35
28 | [7 empty] | 14.1
28.1 | [7 empty] | 13.9556
28.2 | [7 empty] | 13.7625
28.3 | [7 empty] | 13.6286
28.4 | [7 empty] | 13.3714
28.5 | [7 empty] | 13.1571
28.6 | [7 empty] | 12.975
28.7 | [7 empty] | 12.7625
28.8 | [7 empty] | 12.6286
28.9 | [7 empty] | 12.4

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-30.9 | [7 empty] | 15.6875
-30.8 | [7 empty] | 15.9222
-30.7 | [7 empty] | 16.125
-30.6 | [7 empty] | 16.3
-30.5 | [7 empty] | 16.475
-30.4 | [7 empty] | 16.6
-30.3 | [7 empty] | 16.8429
-30.2 | [7 empty] | 16.9857
-30.1 | [7 empty] | 17.2571
-30 | [7 empty] | 17.4714
-29.9 | [7 empty] | 17.625
-29.8 | [7 empty] | 17.875
-29.7 | [7 empty] | 18.125
-29.6 | [7 empty] | 18.2875
-29.5 | [7 empty] | 18.525
-29.4 | [7 empty] | 18.7
-29.3 | [7 empty] | 18.95
-29.2 | [7 empty] | 19.2
-29.1 | [7 empty] | 19.4286
-29 | [7 empty] | 19.6429
-28.9 | [7 empty] | 19.9143
-28.8 | [7 empty] | 20.1333
-28.7 | [7 empty] | 20.425
-28.6 | [7 empty] | 20.725
-28.5 | [7 empty] | 20.975
-28.4 | [7 empty] | 21.225
-28.3 | [7 empty] | 21.475
-28.2 | [7 empty] | 21.8125
-28.1 | [7 empty] | 22.05
-28 | [7 empty] | 22.3556
-27.9 | [7 empty] | 22.6
-27.8 | [7 empty] | 22.9
-27.7 | [7 empty] | 23.225
-27.6 | [7 empty] | 23.5
-27.5 | [7 empty] | 23.8667
-27.4 | [7 empty] | 24.2125
-27.3 | [7 empty] | 24.5875
-27.2 | [7 empty] | 25
-27.1 | [7 empty] | 25.4125
-27 | [7 empty] | 25.8222
-26.9 | [7 empty] | 26.4286
-26.8 | [7 empty] | 27.1
-26.7 | [7 empty] | 28.1
-26.6 | [7 empty] | 29.325
-26.5 | [7 empty] | 31.1
-26.4 | [7 empty] | 34.025
-26.3 | [7 empty] | 37.9
-26.2 | [7 empty] | 42.6
-26.1 | [7 empty] | 48
-26 | [7 empty] | 53.3556
-25.9 | [7 empty] | 58.4
-25.8 | [7 empty] | 63.075
-25.7 | [7 empty] | 67
-25.6 | [7 empty] | 69.7375
-25.5 | [7 empty] | 71.6125
-25.4 | [7 empty] | 72.8875
-25.3 | [7 empty] | 73.9714
-25.2 | [7 empty] | 74.7125
-25.1 | [7 empty] | 75.3375
-25 | [7 empty] | 75.8
-24.9 | [7 empty] | 76.35
-24.8 | [7 empty] | 76.65
-24.7 | [7 empty] | 77.1
-24.6 | [7 empty] | 77.4857
-24.5 | [7 empty] | 77.8429
-24.4 | [6 empty] | 14.6857 | 78.1
-24.3 | [6 empty] | 14.9 | 78.4125
-24.2 | [6 empty] | 15 | 78.7875
-24.1 | [6 empty] | 15.26 | 79.075
-24 | [6 empty] | 15.38 | 79.42
-23.9 | [6 empty] | 15.5667 | 79.8
-23.8 | [6 empty] | 15.8 | 80.075
-23.7 | [6 empty] | 15.9571 | 80.325
-23.6 | [6 empty] | 16.2 | 80.7
-23.5 | [6 empty] | 16.45 | 80.9667
-23.4 | [6 empty] | 16.6167 | 81.25
-23.3 | [6 empty] | 16.8 | 81.5
-23.2 | [6 empty] | 17.0667 | 81.8333
-23.1 | [6 empty] | 17.24 | 82.05
-23 | [6 empty] | 17.54 | 82.3
-22.9 | [6 empty] | 17.74 | 82.475
-22.8 | [6 empty] | 18 | 82.7875
-22.7 | [6 empty] | 18.2667 | 83.1
-22.6 | [6 empty] | 18.4667 | 83.3429
-22.5 | [6 empty] | 18.7 | 83.5714
-22.4 | [6 empty] | 18.9667 | 83.725
-22.3 | [6 empty] | 19.22 | 83.9857
-22.2 | [6 empty] | 19.4714 | 84.25
-22.1 | [6 empty] | 19.78 | 84.5
-22 | [6 empty] | 20.06 | 84.75
-21.9 | [6 empty] | 20.3 | 85
-21.8 | [6 empty] | 20.6 | 85.175
-21.7 | [6 empty] | 20.9429 | 85.4125
-21.6 | [6 empty] | 21.1333 | 85.575
-21.5 | [6 empty] | 21.4 | 85.7833
-21.4 | [6 empty] | 21.7667 | 86.025
-21.3 | [6 empty] | 22.0571 | 86.2375
-21.2 | [6 empty] | 22.3667 | 86.425
-21.1 | [6 empty] | 22.72 | 86.675
-21 | [6 empty] | 23 | 86.8222
-20.9 | [6 empty] | 23.45 | 87
-20.8 | [6 empty] | 23.75 | 87.15
-20.7 | [6 empty] | 24.2333 | 87.3778
-20.6 | [6 empty] | 24.7333 | 87.49
-20.5 | [6 empty] | 25.2 | 87.725
-20.4 | [6 empty] | 25.7 | 87.9375
-20.3 | [6 empty] | 26.4333 | 88.0625
-20.2 | [6 empty] | 27.44 | 88.2286
-20.1 | [6 empty] | 28.8667 | 88.45
-20 | [6 empty] | 30.8 | 88.5857
-19.9 | [6 empty] | 33.6667 | 88.8
-19.8 | [6 empty] | 37.7 | 88.9286
-19.7 | [6 empty] | 42.75 | 89.0625
-19.6 | [6 empty] | 48.5833 | 89.1875
-19.5 | [6 empty] | 54.55 | 89.4125
-19.4 | [6 empty] | 60.45 | 89.5667
-19.3 | [6 empty] | 65.6167 | 89.75
-19.2 | [6 empty] | 69.5 | 89.95
-19.1 | [6 empty] | 72.7667 | 90.1
-19 | [6 empty] | 74.7714 | 90.2143
-18.9 | [6 empty] | 76.1 | 90.4143
-18.8 | [6 empty] | 77.2333 | 90.5875
-18.7 | [6 empty] | 78.0667 | 90.725
-18.6 | [6 empty] | 78.76 | 90.85
-18.5 | [6 empty] | 79.3 | 90.975
-18.4 | [6 empty] | 79.8333 | 91.1125
-18.3 | [6 empty] | 80.3 | 91.2375
-18.2 | [6 empty] | 80.68 | 91.3714
-18.1 | [6 empty] | 81.04 | 91.5
-18 | [5 empty] | 12 | 81.44 | 91.625
-17.9 | [5 empty] | 12.14 | 81.84 | 91.8
-17.8 | [5 empty] | 12.275 | 82.18 | 91.975
-17.7 | [5 empty] | 12.525 | 82.5 | 92.1143
-17.6 | [5 empty] | 12.675 | 82.8333 | 92.2444
-17.5 | [5 empty] | 12.9 | 83.1667 | 92.3625
-17.4 | [5 empty] | 13 | 83.5 | 92.4875
-17.3 | [5 empty] | 13.2 | 83.7667 | 92.6286
-17.2 | [5 empty] | 13.46 | 84.04 | 92.7
-17.1 | [5 empty] | 13.68 | 84.4333 | 92.8143
-17 | [5 empty] | 13.8 | 84.6833 | 92.9571
-16.9 | [5 empty] | 14 | 84.9 | 93.1
-16.8 | [5 empty] | 14.275 | 85.2333 | 93.1222
-16.7 | [5 empty] | 14.475 | 85.4333 | 93.2375
-16.6 | [5 empty] | 14.7 | 85.6833 | 93.3625
-16.5 | [5 empty] | 14.9 | 86.02 | 93.4875
-16.4 | [5 empty] | 15.125 | 86.3143 | 93.6111
-16.3 | [5 empty] | 15.375 | 86.4667 | 93.7286
-16.2 | [5 empty] | 15.625 | 86.6667 | 93.8833
-16.1 | [5 empty] | 15.875 | 86.9 | 94
-16 | [5 empty] | 16.125 | 87.1667 | 94.0429
-15.9 | [5 empty] | 16.375 | 87.4333 | 94.1857
-15.8 | [5 empty] | 16.62 | 87.64 | 94.3125
-15.7 | [5 empty] | 16.95 | 87.9 | 94.4333
-15.6 | [5 empty] | 17.2 | 88.1167 | 94.55
-15.5 | [5 empty] | 17.5333 | 88.2833 | 94.6
-15.4 | [5 empty] | 17.9 | 88.5 | 94.725
-15.3 | [5 empty] | 18.14 | 88.8143 | 94.85
-15.2 | [5 empty] | 18.36 | 89.0333 | 94.975
-15.1 | [5 empty] | 18.65 | 89.2333 | 95
-15 | [5 empty] | 18.975 | 89.4 | 95.15
-14.9 | [5 empty] | 19.35 | 89.5667 | 95.2
-14.8 | [5 empty] | 19.65 | 89.8333 | 95.3125
-14.7 | [5 empty] | 19.925 | 90.0833 | 95.4375
-14.6 | [5 empty] | 20.25 | 90.16 | 95.5
-14.5 | [5 empty] | 20.625 | 90.4 | 95.5875
-14.4 | [5 empty] | 20.9 | 90.62 | 95.7
-14.3 | [5 empty] | 21.3 | 90.8 | 95.725
-14.2 | [5 empty] | 21.8 | 90.9571 | 95.8
-14.1 | [5 empty] | 22.3 | 91.0286 | 95.8875
-14 | [5 empty] | 22.85 | 91.28 | 96.0125
-13.9 | [5 empty] | 23.4 | 91.45 | 96.1
-13.8 | [5 empty] | 24.1 | 91.7143 | 96.1625
-13.7 | [5 empty] | 25.2 | 91.8667 | 96.2
-13.6 | [5 empty] | 26.6 | 92.0333 | 96.3
-13.5 | [5 empty] | 28.55 | 92.2 | 96.4
-13.4 | [5 empty] | 31.4 | 92.3667 | 96.4571
-13.3 | [5 empty] | 35.525 | 92.44 | 96.5875
-13.2 | [5 empty] | 41 | 92.5167 | 96.6125
-13.1 | [5 empty] | 47.75 | 92.7667 | 96.7333
-13 | [5 empty] | 54.7 | 92.95 | 96.8
-12.9 | [5 empty] | 60.8 | 93.14 | 96.875
-12.8 | [5 empty] | 67.8 | 93.2 | 97
-12.7 | [5 empty] | 72.76 | 93.4667 | 97.1
-12.6 | [5 empty] | 76.55 | 93.6 | 97.1375
-12.5 | [5 empty] | 79.15 | 93.7 | 97.2
-12.4 | [5 empty] | 80.8 | 93.8667 | 97.2875
-12.3 | [5 empty] | 82.1 | 94 | 97.3111
-12.2 | [5 empty] | 83.1 | 94.1167 | 97.425
-12.1 | [5 empty] | 83.82 | 94.2833 | 97.5
-12 | [5 empty] | 84.525 | 94.3333 | 97.5
-11.9 | [5 empty] | 84.975 | 94.5 | 97.6
-11.8 | [5 empty] | 85.55 | 94.5667 | 97.6857
-11.7 | [5 empty] | 86.05 | 94.8 | 97.8
-11.6 | [5 empty] | 86.375 | 94.8833 | 97.8333
-11.5 | [5 empty] | 86.75 | 94.9 | 97.9
-11.4 |  |  |  |  | 7.8 | 87.2 | 95.04 | 97.9
-11.3 |  |  |  |  | 8.075 | 87.46 | 95.2333 | 98
-11.2 |  |  |  |  | 8.125 | 87.7 | 95.4 | 98.1
-11.1 |  |  |  |  | 8.36 | 88.05 | 95.4571 | 98.1625
-11 |  |  |  |  | 8.56 | 88.55 | 95.6333 | 98.2875
-10.9 |  |  |  |  | 8.6 | 88.9 | 95.7 | 98.3111
-10.8 |  |  |  |  | 8.75 | 89.15 | 95.8143 | 98.4
-10.7 |  |  |  |  | 9 | 89.5 | 95.9667 | 98.45
-10.6 |  |  |  |  | 9.0667 | 89.775 | 96 | 98.5
-10.5 |  |  |  |  | 9.25 | 90.0667 | 96.15 | 98.6
-10.4 |  |  |  |  | 9.4 | 90.35 | 96.2 | 98.6
-10.3 |  |  |  |  | 9.6 | 90.6 | 96.3167 | 98.6
-10.2 |  |  |  |  | 9.85 | 90.85 | 96.4833 | 98.6778
-10.1 |  |  |  | 6.6 | 9.95 | 91.15 | 96.55 | 98.7
-10 |  |  |  | 6.64 | 10.175 | 91.3 | 96.6333 | 98.75
-9.9 |  |  |  | 6.78 | 10.425 | 91.625 | 96.7 | 98.875
-9.8 |  |  |  | 6.98 | 10.525 | 91.875 | 96.8667 | 98.9
-9.7 |  |  |  | 7.0333 | 10.775 | 92.12 | 96.9333 | 99
-9.6 |  |  |  | 7.175 | 10.96 | 92.3667 | 97.0167 | 99.0375
-9.5 |  |  |  | 7.325 | 11.2 | 92.6 | 97.1833 | 99.1
-9.4 |  |  |  | 7.5 | 11.42 | 92.75 | 97.26 | 99.1
-9.3 |  |  |  | 7.625 | 11.65 | 92.9 | 97.3 | 99.2
-9.2 |  |  |  | 7.8 | 11.875 | 93.15 | 97.4 | 99.2429
-9.1 |  |  |  | 7.96 | 12.1 | 93.4 | 97.4833 | 99.3
-9 |  |  |  | 8.1 | 12.36 | 93.575 | 97.5333 | 99.4
-8.9 |  |  |  | 8.2333 | 12.52 | 93.825 | 97.62 | 99.4375
-8.8 |  |  | 4.9 | 8.4 | 12.86 | 94.06 | 97.7 | 99.5
-8.7 |  |  | 5 | 8.55 | 13.1 | 94.22 | 97.8 | 99.5
-8.6 |  |  | 5.2 | 8.675 | 13.48 | 94.4 | 97.92 | 99.5
-8.5 |  |  | 5.3 | 8.94 | 13.775 | 94.7 | 98.0167 | 99.5
-8.4 |  |  | 5.4 | 9.2 | 14.02 | 94.85 | 98.1 | 99.5556
-8.3 |  |  | 5.45 | 9.4 | 14.3 | 94.98 | 98.2 | 99.6
-8.2 |  |  | 5.55 | 9.6 | 14.68 | 95.1667 | 98.3 | 99.6
-8.1 |  |  | 5.7 | 9.8 | 14.9667 | 95.325 | 98.3 | 99.6
-8 |  |  | 5.8 | 9.9 | 15.28 | 95.475 | 98.3667 | 99.6429
-7.9 |  |  | 5.95 | 10.1 | 15.6333 | 95.625 | 98.5 | 99.7
-7.8 |  |  | 6.1 | 10.325 | 16.05 | 95.8 | 98.5 | 99.8
-7.7 |  |  | 6.3 | 10.575 | 16.375 | 96 | 98.6 | 99.8
-7.6 |  |  | 6.4 | 10.8 | 16.95 | 96.1 | 98.68 | 99.8556
-7.5 |  | 3.4 | 6.5 | 11.1 | 17.54 | 96.3 | 98.7 | 99.9
-7.4 |  | 3.5 | 6.6 | 11.35 | 18.1 | 96.4 | 98.7 | 99.9
-7.3 |  | 3.5 | 6.8 | 11.575 | 19 | 96.5 | 98.8 | 100
-7.2 |  | 3.5 | 6.9 | 11.8 | 20 | 96.625 | 98.9 | 100
-7.1 |  | 3.6667 | 7 | 12.2 | 21.5 | 96.775 | 98.9 | 100
-7 |  | 3.8 | 7.25 | 12.4 | 23.65 | 96.9 | 99 | 100
-6.9 | 2.5 | 3.8 | 7.4 | 12.75 | 27 | 97 | 99.1 | 100.1
-6.8 | 2.5 | 3.9 | 7.5 | 13.1 | 31.9 | 97.2 | 99.1 | 100.1625
-6.7 | 2.5 | 3.9 | 7.7 | 13.4 | 38.3 | 97.3 | 99.2 | 100.2
-6.6 | 2.5 | 4 | 7.9 | 13.7333 | 46.2 | 97.4 | 99.2 | 100.2
-6.5 | 2.6 | 4.1 | 8.1 | 14.05 | 54.075 | 97.6 | 99.3 | 100.2
-6.4 | 2.6 | 4.2 | 8.3 | 14.5 | 62.5 | 97.775 | 99.34 | 100.2
-6.3 | 2.7 | 4.3 | 8.55 | 14.95 | 70.45 | 97.825 | 99.4 | 100.2
-6.2 | 2.7 | 4.4 | 8.75 | 15.4 | 77.45 | 97.9 | 99.4833 | 100.2
-6.1 | 2.8 | 4.4 | 8.9 | 16.1 | 82.4 | 98 | 99.5 | 100.3
-6 | 2.8 | 4.5 | 9 | 16.9 | 85.725 | 98 | 99.5 | 100.3
-5.9 | 2.8 | 4.6 | 9.3 | 17.9 | 88.075 | 98.1 | 99.5 | 100.3
-5.8 | 2.9 | 4.7 | 9.45 | 19.475 | 89.5 | 98.2 | 99.5 | 100.3
-5.7 | 3 | 4.8 | 9.9 | 21.7 | 90.58 | 98.275 | 99.5 | 100.3
-5.6 | 3.1 | 4.95 | 10.1 | 24.9 | 91.4 | 98.3667 | 99.5 | 100.2
-5.5 | 3.1 | 5.1 | 10.3 | 29.2 | 91.9 | 98.4 | 99.5 | 100.2125
-5.4 | 3.2 | 5.2 | 10.75 | 36.4 | 92.5 | 98.5 | 99.6 | 100.3
-5.3 | 3.2667 | 5.3 | 11 | 44.5333 | 92.9333 | 98.55 | 99.6 | 100.2
-5.2 | 3.4 | 5.5 | 11.3 | 53.1 | 93.4 | 98.6 | 99.6 | 100.2
-5.1 | 3.45 | 5.6667 | 11.7 | 61.85 | 93.64 | 98.64 | 99.6 | 100.2
-5 | 3.5 | 5.7667 | 12.15 | 70.5 | 93.95 | 98.7 | 99.6333 | 100.2
-4.9 | 3.7 | 5.9 | 12.6667 | 77.3 | 94.34 | 98.8 | 99.7 | 100.2
-4.8 | 3.75 | 6.1 | 13.2 | 83.1 | 94.6667 | 98.9 | 99.7 | 100.2
-4.7 | 3.9 | 6.2 | 14.1 | 86.72 | 94.9 | 98.95 | 99.7 | 100.2
-4.6 | 3.9 | 6.4 | 15 | 89.3 | 95.25 | 99 | 99.7 | 100.2875
-4.5 | 4 | 6.5 | 16.5667 | 90.8 | 95.4 | 99.1 | 99.7 | 100.2
-4.4 | 4.1 | 6.8 | 18.7333 | 92.1 | 95.64 | 99.1 | 99.7 | 100.2333
-4.3 | 4.2 | 7 | 21.8 | 92.7333 | 95.875 | 99.15 | 99.74 | 100.3
-4.2 | 4.3 | 7.2 | 26.6 | 93.55 | 96.175 | 99.2 | 99.8 | 100.3
-4.1 | 4.4 | 7.4 | 33.8 | 94.15 | 96.4 | 99.2 | 99.8 | 100.3
-4 | 4.5 | 7.8 | 41.7 | 94.55 | 96.55 | 99.3 | 99.8571 | 100.3
-3.9 | 4.7 | 7.95 | 51.5 | 94.8 | 96.7333 | 99.3 | 99.9 | 100.3
-3.8 | 4.8 | 8.2333 | 61.3 | 95.1 | 97.1 | 99.4 | 99.9 | 100.3
-3.7 | 4.9667 | 8.6333 | 69.5 | 95.56 | 97.25 | 99.4 | 99.9833 | 100.3
-3.6 | 5.2 | 9 | 77.8 | 95.8667 | 97.4 | 99.45 | 100 | 100.3
-3.5 | 5.4 | 9.7 | 84.1 | 96.2 | 97.575 | 99.5 | 100.06 | 100.3
-3.4 | 5.6 | 10.4333 | 88.4 | 96.4667 | 97.825 | 99.6 | 100.1 | 100.3
-3.3 | 5.7 | 11.4 | 91.3667 | 96.725 | 97.95 | 99.6 | 100.1 | 100.3
-3.2 | 6 | 12.85 | 93 | 97.02 | 98.2 | 99.6333 | 100.1 | 100.3
-3.1 | 6.2 | 15 | 94.4 | 97.175 | 98.35 | 99.7 | 100.1 | 100.3444
-3 | 6.6333 | 18.25 | 95.2 | 97.5 | 98.48 | 99.75 | 100.2 | 100.4
-2.9 | 7 | 23.2333 | 95.85 | 97.75 | 98.65 | 99.8 | 100.2 | 100.4
-2.8 | 7.4667 | 30.2 | 96.3333 | 97.9 | 98.8 | 99.85 | 100.2 | 100.4
-2.7 | 8.3 | 39.7 | 96.8 | 98.1 | 98.875 | 100 | 100.2 | 100.4
-2.6 | 9.5 | 49.8 | 97.1 | 98.3 | 99 | 100 | 100.2 | 100.4
-2.5 | 11.4 | 60.5 | 97.6 | 98.425 | 99.05 | 100 | 100.2 | 100.4
-2.4 | 13.9 | 70.9 | 97.8 | 98.575 | 99.2 | 100 | 100.2 | 100.4
-2.3 | 18.4 | 79.6 | 98 | 98.8 | 99.3 | 100.1 | 100.1 | 100.4
-2.2 | 24.25 | 85.8 | 98.2 | 98.96 | 99.4 | 100.1 | 100.1 | 100.4
-2.1 | 32.8 | 90.7 | 98.4 | 99.1 | 99.55 | 100.1 | 100.2 | 100.3
-2 | 43.15 | 93.6 | 98.5 | 99.2 | 99.6 | 100.1 | 100.2 | 100.3
-1.9 | 53.6 | 95.5 | 98.7 | 99.2 | 99.6 | 100.1 | 100.2 | 100.3
-1.8 | 64.85 | 96.6 | 98.8 | 99.38 | 99.7 | 100.1 | 100.2 | 100.3
-1.7 | 75.2 | 97.4 | 99.1 | 99.4 | 99.8 | 100.1 | 100.2 | 100.3
-1.6 | 82.9 | 98 | 99.3 | 99.5 | 99.8 | 100.1 | 100.2 | 100.3
-1.5 | 89.1 | 98.4 | 99.35 | 99.52 | 99.825 | 100.1 | 100.2 | 100.3
-1.4 | 92.9 | 98.85 | 99.6 | 99.6 | 99.9 | 100.1 | 100.2 | 100.2
-1.3 | 95.2 | 99.1 | 99.6 | 99.7 | 99.9 | 100.1 | 100.2 | 100.2
-1.2 | 96.9 | 99.3 | 99.7 | 99.7 | 100 | 100.1 | 100.1667 | 100.2
-1.1 | 97.8 | 99.4 | 99.8 | 99.8 | 100 | 100.1 | 100.1 | 100.1125
-1 | 98.45 | 99.5333 | 99.8 | 99.8 | 100 | 100.1 | 100.1 | 100.1
-0.9 | 98.95 | 99.7 | 99.9 | 99.8 | 100 | 100.05 | 100.1 | 100.1
-0.8 | 99.2 | 99.8 | 99.9 | 99.85 | 100 | 100.08 | 100.1 | 100.1
-0.7 | 99.4 | 99.8 | 99.9 | 99.9 | 100 | 100.1 | 100.1 | 100.1
-0.6 | 99.6 | 99.9 | 99.9 | 99.9 | 100 | 100.1 | 100.0667 | 100.05
-0.5 | 99.8 | 100 | 100 | 100 | 100.1 | 100.1 | 100 | 100
-0.4 | 99.9 | 100 | 100 | 100 | 100.1 | 100.1 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100 | 100.1 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100.1 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.9 | 100 | 100 | 100.1 | 100 | 100 | 100 | 100
0.4 | 99.8 | 100.1 | 100 | 100 | 100.1 | 100 | 100 | 100
0.5 | 99.7667 | 100 | 99.9 | 100 | 100.1 | 100 | 100 | 100
0.6 | 99.6 | 99.95 | 99.9 | 99.9667 | 100.1 | 100 | 100.1 | 100
0.7 | 99.2 | 99.8 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100.1
0.8 | 99.05 | 99.7 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100.1
0.9 | 98.7667 | 99.6 | 99.7 | 99.9 | 99.9 | 100 | 100.1 | 100.1
1 | 98.35 | 99.5 | 99.7 | 99.8 | 99.9 | 100.1 | 100.1 | 100.1
1.1 | 97.7 | 99.3 | 99.6 | 99.8 | 99.9 | 100.1 | 100.1 | 100.1625
1.2 | 96.5 | 99.15 | 99.5 | 99.7 | 99.9 | 100.1 | 100.1 | 100.2
1.3 | 95.1 | 98.8 | 99.4 | 99.6 | 99.8 | 100.1 | 100.2 | 100.2
1.4 | 92.7333 | 98.6 | 99.3 | 99.6 | 99.8 | 100.1 | 100.2 | 100.2
1.5 | 88.8 | 98.2667 | 99.2 | 99.5 | 99.8 | 100.04 | 100.2 | 100.25
1.6 | 82.5 | 97.8667 | 99.2 | 99.5 | 99.775 | 100 | 100.2 | 100.3
1.7 | 74.4 | 97.2 | 99 | 99.4 | 99.7 | 100 | 100.2 | 100.3
1.8 | 64.7 | 96.3 | 98.8 | 99.3 | 99.7 | 100 | 100.2 | 100.3
1.9 | 54.35 | 95.1 | 98.7 | 99.2 | 99.6 | 100 | 100.2 | 100.3
2 | 43 | 93.2 | 98.5 | 99.05 | 99.6 | 100 | 100.2 | 100.3
2.1 | 32.8667 | 90.4 | 98.3 | 98.9 | 99.5 | 100.075 | 100.2 | 100.3375
2.2 | 24.0667 | 85.5 | 98.1 | 98.8 | 99.4 | 100.1 | 100.12 | 100.4
2.3 | 18.1 | 78.8 | 97.9 | 98.55 | 99.3 | 100.1 | 100.1333 | 100.4
2.4 | 14.1 | 70.5 | 97.7333 | 98.5 | 99.2 | 100 | 100.2 | 100.4
2.5 | 11.3 | 60.3 | 97.3333 | 98.3 | 99.1 | 99.9 | 100.2 | 100.4
2.6 | 9.6 | 49.3333 | 97.0667 | 98.125 | 99 | 99.9 | 100.2 | 100.4
2.7 | 8.3 | 39.3 | 96.7 | 97.975 | 98.875 | 99.8 | 100.2 | 100.4
2.8 | 7.5 | 30.1 | 96.3 | 97.8 | 98.725 | 99.8 | 100.2 | 100.4
2.9 | 7.15 | 22.7 | 95.85 | 97.7 | 98.5 | 99.7 | 100.2 | 100.4
3 | 6.55 | 17.9 | 95 | 97.5 | 98.4 | 99.7 | 100.2 | 100.4
3.1 | 6.3 | 14.8 | 94.2 | 97.32 | 98.3 | 99.6 | 100.1 | 100.3
3.2 | 5.9 | 12.8 | 92.7 | 96.95 | 98.0667 | 99.6 | 100.1 | 100.3
3.3 | 5.8 | 11.3 | 91.0667 | 96.8 | 97.9 | 99.6 | 100.1 | 100.3
3.4 | 5.65 | 10.3667 | 88.3 | 96.5 | 97.76 | 99.6 | 100.1 | 100.375
3.5 | 5.4 | 9.5667 | 84.2 | 96.15 | 97.525 | 99.5 | 100.025 | 100.4
3.6 | 5.3 | 9 | 77.7 | 95.7667 | 97.325 | 99.4 | 100 | 100.3
3.7 | 5.0667 | 8.5667 | 69.8 | 95.46 | 97.18 | 99.28 | 99.95 | 100.3
3.8 | 4.9 | 8.2 | 61.35 | 95.18 | 97 | 99.125 | 99.9 | 100.3
3.9 | 4.7 | 7.9 | 51.1333 | 94.7 | 96.8 | 99.1 | 99.9 | 100.3
4 | 4.6 | 7.55 | 41.8 | 94.35 | 96.55 | 99 | 99.8333 | 100.3
4.1 | 4.5 | 7.3667 | 33.4 | 94.04 | 96.3 | 99 | 99.8 | 100.3
4.2 | 4.3 | 7.0667 | 26.65 | 93.4 | 96.06 | 99 | 99.8 | 100.3
4.3 | 4.3 | 6.9 | 21.8 | 92.65 | 95.825 | 99 | 99.8 | 100.3
4.4 | 4.2 | 6.7 | 18.6 | 92 | 95.575 | 99 | 99.7 | 100.2
4.5 | 4 | 6.5 | 16.2667 | 90.625 | 95.325 | 98.84 | 99.7 | 100.2375
4.6 | 3.9667 | 6.4 | 14.75 | 89 | 95.1 | 98.8 | 99.7 | 100.2375
4.7 | 3.8333 | 6.2 | 14 | 86.62 | 94.75 | 98.75 | 99.7 | 100.2
4.8 | 3.8 | 6 | 13.2 | 82.8 | 94.5 | 98.7 | 99.7 | 100.2
4.9 | 3.7 | 5.9 | 12.5667 | 77.3 | 94.3333 | 98.6 | 99.7 | 100.2
5 | 3.5667 | 5.65 | 12.1 | 69.9333 | 93.9667 | 98.525 | 99.6 | 100.2
5.1 | 3.5 | 5.5 | 11.75 | 61.7667 | 93.55 | 98.5 | 99.6 | 100.2
5.2 | 3.4333 | 5.4 | 11.3 | 52.9 | 93.15 | 98.44 | 99.6 | 100.2
5.3 | 3.3 | 5.2 | 11 | 44 | 92.8667 | 98.4 | 99.6 | 100.2
5.4 | 3.3 | 5.1 | 10.7 | 36.3 | 92.3333 | 98.26 | 99.6 | 100.3
5.5 | 3.2 | 5 | 10.4 | 29.4 | 91.8 | 98.2 | 99.5 | 100.2
5.6 | 3.1667 | 4.95 | 10.1 | 24.8 | 91.2 | 98.2 | 99.5 | 100.25
5.7 | 3.1 | 4.8 | 9.8 | 21.5 | 90.3 | 98.125 | 99.5 | 100.3
5.8 | 3 | 4.7 | 9.55 | 19.275 | 89.3 | 98.1 | 99.5 | 100.3
5.9 | 2.9 | 4.6 | 9.3 | 17.9 | 87.76 | 98 | 99.5 | 100.3
6 | 2.8 | 4.55 | 9.1 | 16.9 | 85.625 | 97.9 | 99.5 | 100.3
6.1 | 2.8 | 4.4 | 8.9 | 16.1333 | 82.16 | 97.8 | 99.5 | 100.3
6.2 | 2.8 | 4.3 | 8.7 | 15.5 | 77.15 | 97.7 | 99.45 | 100.2375
6.3 | 2.8 | 4.2 | 8.5 | 14.925 | 70.575 | 97.6 | 99.4 | 100.2
6.4 | 2.8 | 4.1 | 8.2 | 14.5 | 62.65 | 97.5 | 99.4 | 100.2
6.5 | 2.7 | 4 | 8.05 | 14.1 | 54.375 | 97.4 | 99.3 | 100.2
6.6 | 2.6 | 3.9 | 7.9 | 13.6667 | 45.72 | 97.3 | 99.2 | 100.2
6.7 | 2.5 | 3.9 | 7.75 | 13.35 | 38.32 | 97.2 | 99.2 | 100.2
6.8 | 2.5 | 3.8 | 7.6 | 13 | 31.8 | 97.0333 | 99.1 | 100.1167
6.9 | 2.5 | 3.8 | 7.4 | 12.775 | 27 | 96.9 | 99.0833 | 100.1
7 |  | 3.7 | 7.2 | 12.45 | 23.7 | 96.75 | 99 | 100.0429
7.1 |  | 3.6667 | 7 | 12.2 | 21.5 | 96.6 | 98.9 | 100
7.2 |  | 3.5 | 6.8 | 11.9 | 20.05 | 96.525 | 98.9 | 100
7.3 |  | 3.5 | 6.7 | 11.575 | 19 | 96.425 | 98.8 | 100
7.4 |  | 3.4 | 6.6 | 11.3 | 18.15 | 96.275 | 98.7 | 99.95
7.5 |  | 3.35 | 6.45 | 11.0333 | 17.5 | 96.125 | 98.7 | 99.9
7.6 |  |  | 6.4 | 10.8 | 16.95 | 95.9667 | 98.7 | 99.9
7.7 |  |  | 6.2667 | 10.6 | 16.525 | 95.8 | 98.6 | 99.8
7.8 |  |  | 6.1 | 10.4 | 16.05 | 95.6 | 98.56 | 99.7571
7.9 |  |  | 5.9 | 10.225 | 15.7333 | 95.5 | 98.5 | 99.7
8 |  |  | 5.8 | 10 | 15.34 | 95.3 | 98.4 | 99.6
8.1 |  |  | 5.7 | 9.75 | 15.0333 | 95.025 | 98.3 | 99.6
8.2 |  |  | 5.6333 | 9.5 | 14.7 | 94.875 | 98.3 | 99.6
8.3 |  |  | 5.5 | 9.28 | 14.35 | 94.7333 | 98.2 | 99.6
8.4 |  |  | 5.4 | 9.12 | 14.1 | 94.6 | 98.1 | 99.5
8.5 |  |  | 5.3 | 8.9 | 13.7 | 94.4 | 98.0667 | 99.5
8.6 |  |  | 5.2 | 8.7 | 13.425 | 94.3 | 97.9 | 99.5
8.7 |  |  | 5.1 | 8.6 | 13.2 | 94.075 | 97.8 | 99.5
8.8 |  |  | 5 | 8.4 | 12.8 | 93.9 | 97.7 | 99.5
8.9 |  |  |  | 8.225 | 12.62 | 93.68 | 97.6 | 99.4
9 |  |  |  | 8.1 | 12.425 | 93.4333 | 97.5 | 99.4
9.1 |  |  |  | 8 | 12.15 | 93.1 | 97.5 | 99.325
9.2 |  |  |  | 7.8 | 11.82 | 92.9 | 97.3833 | 99.2857
9.3 |  |  |  | 7.7 | 11.7 | 92.675 | 97.3 | 99.1556
9.4 |  |  |  | 7.575 | 11.4 | 92.5 | 97.225 | 99.1
9.5 |  |  |  | 7.4 | 11.2 | 92.3 | 97.15 | 99.1
9.6 |  |  |  | 7.2 | 10.9 | 92.075 | 97.08 | 99.0875
9.7 |  |  |  | 7.1 | 10.7 | 91.925 | 97 | 99
9.8 |  |  |  | 6.925 | 10.6 | 91.6 | 96.8333 | 98.9444
9.9 |  |  |  | 6.8 | 10.4 | 91.5 | 96.7667 | 98.825
10 |  |  |  | 6.8 | 10.225 | 91.25 | 96.6 | 98.7
10.1 |  |  |  | 6.65 | 9.975 | 91.1 | 96.5167 | 98.7
10.2 |  |  |  |  | 9.84 | 90.65 | 96.45 | 98.6375
10.3 |  |  |  |  | 9.7 | 90.36 | 96.3 | 98.6
10.4 |  |  |  |  | 9.46 | 90.125 | 96.2 | 98.6
10.5 |  |  |  |  | 9.2333 | 89.775 | 96.12 | 98.6
10.6 |  |  |  |  | 9.1 | 89.6 | 96.0333 | 98.525
10.7 |  |  |  |  | 8.9 | 89.3 | 95.94 | 98.5
10.8 |  |  |  |  | 8.8 | 89 | 95.7833 | 98.375
10.9 |  |  |  |  | 8.65 | 88.6333 | 95.7 | 98.3
11 |  |  |  |  | 8.5 | 88.3 | 95.6 | 98.2375
11.1 |  |  |  |  | 8.3333 | 87.9 | 95.52 | 98.2
11.2 |  |  |  |  | 8.1333 | 87.5 | 95.3667 | 98.1
11.3 |  |  |  |  | 8.075 | 87.25 | 95.2 | 98
11.4 |  |  |  |  | 7.9 | 86.9 | 95 | 97.925
11.5 | [5 empty] | 86.4 | 94.9429 | 97.9
11.6 | [5 empty] | 86.025 | 94.86 | 97.8875
11.7 | [5 empty] | 85.625 | 94.7857 | 97.7625
11.8 | [5 empty] | 85.15 | 94.6333 | 97.6375
11.9 | [5 empty] | 84.675 | 94.46 | 97.6
12 | [5 empty] | 84.25 | 94.3 | 97.5429
12.1 | [5 empty] | 83.6333 | 94.25 | 97.5
12.2 | [5 empty] | 82.9 | 94.0833 | 97.4
12.3 | [5 empty] | 82 | 93.9857 | 97.35
12.4 | [5 empty] | 80.7 | 93.8333 | 97.3
12.5 | [5 empty] | 78.75 | 93.6667 | 97.2125
12.6 | [5 empty] | 76.3 | 93.6 | 97.1875
12.7 | [5 empty] | 72.55 | 93.4333 | 97.0667
12.8 | [5 empty] | 67.5 | 93.2714 | 97
12.9 | [5 empty] | 61.38 | 93.1 | 96.925
13 | [5 empty] | 54.45 | 92.9167 | 96.8111
13.1 | [5 empty] | 47.725 | 92.75 | 96.7
13.2 | [5 empty] | 41.45 | 92.5833 | 96.6625
13.3 | [5 empty] | 35.95 | 92.4 | 96.5375
13.4 | [5 empty] | 31.525 | 92.3333 | 96.5
13.5 | [5 empty] | 28.45 | 92.1667 | 96.4
13.6 | [5 empty] | 26.45 | 92 | 96.35
13.7 | [5 empty] | 25.1 | 91.8429 | 96.2333
13.8 | [5 empty] | 24.2 | 91.6714 | 96.1125
13.9 | [5 empty] | 23.45 | 91.5167 | 96.0875
14 | [5 empty] | 22.8 | 91.25 | 96
14.1 | [5 empty] | 22.3 | 91.0857 | 95.8889
14.2 | [5 empty] | 21.8333 | 90.9333 | 95.8
14.3 | [5 empty] | 21.45 | 90.7667 | 95.775
14.4 | [5 empty] | 21.02 | 90.58 | 95.65
14.5 | [5 empty] | 20.575 | 90.4143 | 95.6
14.6 | [5 empty] | 20.225 | 90.24 | 95.5125
14.7 | [5 empty] | 19.85 | 90.05 | 95.3875
14.8 | [5 empty] | 19.525 | 89.7833 | 95.2625
14.9 | [5 empty] | 19.2 | 89.5571 | 95.2
15 | [5 empty] | 18.84 | 89.3667 | 95.1
15.1 | [5 empty] | 18.6 | 89.2 | 95.0571
15.2 | [5 empty] | 18.25 | 89.0333 | 95
15.3 | [5 empty] | 18 | 88.8429 | 94.9
15.4 | [5 empty] | 17.66 | 88.5333 | 94.775
15.5 | [5 empty] | 17.42 | 88.25 | 94.6333
15.6 | [5 empty] | 17.15 | 88.0857 | 94.5
15.7 | [5 empty] | 16.825 | 87.92 | 94.3875
15.8 | [5 empty] | 16.675 | 87.6 | 94.2625
15.9 | [5 empty] | 16.325 | 87.4333 | 94.1375
16 | [5 empty] | 16.08 | 87.2333 | 94.0875
16.1 | [5 empty] | 15.9 | 86.9429 | 93.9667
16.2 | [5 empty] | 15.7 | 86.7 | 93.9
16.3 | [5 empty] | 15.45 | 86.44 | 93.7778
16.4 | [5 empty] | 15.2 | 86.2714 | 93.625
16.5 | [5 empty] | 14.9 | 86 | 93.4375
16.6 | [5 empty] | 14.7 | 85.75 | 93.3125
16.7 | [5 empty] | 14.5 | 85.475 | 93.1875
16.8 | [5 empty] | 14.2667 | 85.2667 | 93.1
16.9 | [5 empty] | 14.025 | 84.85 | 93.0429
17 | [5 empty] | 13.85 | 84.65 | 92.9
17.1 | [5 empty] | 13.7 | 84.3667 | 92.8667
17.2 | [5 empty] | 13.3667 | 84 | 92.725
17.3 | [5 empty] | 13.2 | 83.7667 | 92.5714
17.4 | [5 empty] | 12.95 | 83.5333 | 92.4375
17.5 | [5 empty] | 12.7 | 83.2 | 92.3125
17.6 | [5 empty] | 12.6 | 82.8667 | 92.2
17.7 | [5 empty] | 12.46 | 82.5429 | 92.0667
17.8 | [5 empty] | 12.3 | 82.2 | 91.925
17.9 | [5 empty] | 12.08 | 81.8667 | 91.8
18 | [6 empty] | 81.38 | 91.675
18.1 | [6 empty] | 81.0667 | 91.55
18.2 | [6 empty] | 80.7 | 91.3286
18.3 | [6 empty] | 80.2667 | 91.2
18.4 | [6 empty] | 79.7714 | 91.1667
18.5 | [6 empty] | 79.3 | 91.025
18.6 | [6 empty] | 78.74 | 90.9
18.7 | [6 empty] | 78.1 | 90.675
18.8 | [6 empty] | 77.2667 | 90.5429
18.9 | [6 empty] | 76.1667 | 90.4
19 | [6 empty] | 74.7 | 90.2667
19.1 | [6 empty] | 72.62 | 90.15
19.2 | [6 empty] | 69.82 | 89.95
19.3 | [6 empty] | 65.6833 | 89.7
19.4 | [6 empty] | 60.5167 | 89.5875
19.5 | [6 empty] | 54.66 | 89.425
19.6 | [6 empty] | 48.55 | 89.2375
19.7 | [6 empty] | 42.7667 | 89.1125
19.8 | [6 empty] | 37.8857 | 88.95
19.9 | [6 empty] | 33.72 | 88.7556
20 | [6 empty] | 30.8 | 88.6286
20.1 | [6 empty] | 28.9 | 88.5
20.2 | [6 empty] | 27.4 | 88.2667
20.3 | [6 empty] | 26.4667 | 88.1143
20.4 | [6 empty] | 25.7143 | 87.975
20.5 | [6 empty] | 25.1667 | 87.725
20.6 | [6 empty] | 24.7 | 87.5444
20.7 | [6 empty] | 24.25 | 87.3667
20.8 | [6 empty] | 23.8 | 87.1444
20.9 | [6 empty] | 23.4333 | 86.9857
21 | [6 empty] | 23.04 | 86.8556
21.1 | [6 empty] | 22.64 | 86.675
21.2 | [6 empty] | 22.42 | 86.425
21.3 | [6 empty] | 22.0714 | 86.275
21.4 | [6 empty] | 21.7667 | 86.025
21.5 | [6 empty] | 21.4333 | 85.7571
21.6 | [6 empty] | 21.1 | 85.575
21.7 | [6 empty] | 20.95 | 85.3625
21.8 | [6 empty] | 20.66 | 85.2286
21.9 | [6 empty] | 20.3286 | 85
22 | [6 empty] | 20 | 84.75
22.1 | [6 empty] | 19.75 | 84.5
22.2 | [6 empty] | 19.5143 | 84.25
22.3 | [6 empty] | 19.2 | 83.975
22.4 | [6 empty] | 18.9571 | 83.725
22.5 | [6 empty] | 18.7333 | 83.5375
22.6 | [6 empty] | 18.4 | 83.375
22.7 | [6 empty] | 18.2333 | 83.1
22.8 | [6 empty] | 18.0333 | 82.825
22.9 | [6 empty] | 17.8 | 82.475
23 | [6 empty] | 17.5167 | 82.25
23.1 | [6 empty] | 17.2333 | 82.0444
23.2 | [6 empty] | 17.05 | 81.7857
23.3 | [6 empty] | 16.775 | 81.55
23.4 | [6 empty] | 16.5833 | 81.25
23.5 | [6 empty] | 16.4167 | 80.9714
23.6 | [6 empty] | 16.2143 | 80.65
23.7 | [6 empty] | 15.9857 | 80.3
23.8 | [6 empty] | 15.82 | 80.0571
23.9 | [6 empty] | 15.6333 | 79.7429
24 | [6 empty] | 15.35 | 79.4429
24.1 | [6 empty] | 15.25 | 79.1125
24.2 | [6 empty] | 15 | 78.825
24.3 | [6 empty] | 14.8667 | 78.4625
24.4 | [6 empty] | 14.7143 | 78.1143
24.5 | [7 empty] | 77.8286
24.6 | [7 empty] | 77.5
24.7 | [7 empty] | 77.05
24.8 | [7 empty] | 76.675
24.9 | [7 empty] | 76.3
25 | [7 empty] | 75.8286
25.1 | [7 empty] | 75.2667
25.2 | [7 empty] | 74.6625
25.3 | [7 empty] | 73.9167
25.4 | [7 empty] | 72.9375
25.5 | [7 empty] | 71.55
25.6 | [7 empty] | 69.75
25.7 | [7 empty] | 66.95
25.8 | [7 empty] | 63.0125
25.9 | [7 empty] | 58.6
26 | [7 empty] | 53.3429
26.1 | [7 empty] | 47.7
26.2 | [7 empty] | 42.55
26.3 | [7 empty] | 37.8
26.4 | [7 empty] | 34.05
26.5 | [7 empty] | 31.2
26.6 | [7 empty] | 29.325
26.7 | [7 empty] | 28
26.8 | [7 empty] | 27.1571
26.9 | [7 empty] | 26.4429
27 | [7 empty] | 25.8556
27.1 | [7 empty] | 25.3625
27.2 | [7 empty] | 24.9714
27.3 | [7 empty] | 24.5429
27.4 | [7 empty] | 24.1778
27.5 | [7 empty] | 23.925
27.6 | [7 empty] | 23.55
27.7 | [7 empty] | 23.175
27.8 | [7 empty] | 22.9
27.9 | [7 empty] | 22.6143
28 | [7 empty] | 22.35
28.1 | [7 empty] | 22
28.2 | [7 empty] | 21.825
28.3 | [7 empty] | 21.5125
28.4 | [7 empty] | 21.225
28.5 | [7 empty] | 20.975
28.6 | [7 empty] | 20.7143
28.7 | [7 empty] | 20.4286
28.8 | [7 empty] | 20.1
28.9 | [7 empty] | 19.9333
29 | [7 empty] | 19.65
29.1 | [7 empty] | 19.45
29.2 | [7 empty] | 19.2
29.3 | [7 empty] | 18.95
29.4 | [7 empty] | 18.7
29.5 | [7 empty] | 18.4714
29.6 | [7 empty] | 18.275
29.7 | [7 empty] | 18.1125
29.8 | [7 empty] | 17.875
29.9 | [7 empty] | 17.625
30 | [7 empty] | 17.4167
30.1 | [7 empty] | 17.2667
30.2 | [7 empty] | 17.0375
30.3 | [7 empty] | 16.9
30.4 | [7 empty] | 16.6143
30.5 | [7 empty] | 16.425
30.6 | [7 empty] | 16.3
30.7 | [7 empty] | 16.075
30.8 | [7 empty] | 15.9111
30.9 | [7 empty] | 15.675

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-43.7 |  |  |  |  | 7.0743
-43.6 |  |  |  |  | 7.1
-43.5 |  |  |  |  | 7.1343
-43.4 |  |  |  |  | 7.2
-43.3 |  |  |  |  | 7.2032
-43.2 |  |  |  |  | 7.3
-43.1 |  |  |  |  | 7.3
-43 |  |  |  |  | 7.4063
-42.9 |  |  |  |  | 7.5
-42.8 |  |  |  |  | 7.5
-42.7 |  |  |  |  | 7.5739
-42.6 |  |  |  |  | 7.6591
-42.5 |  |  |  |  | 7.7
-42.4 |  |  |  |  | 7.743
-42.3 |  |  |  |  | 7.8242
-42.2 |  |  |  |  | 7.9
-42.1 |  |  |  |  | 7.901
-42 |  |  |  |  | 7.9894
-41.9 |  |  |  |  | 8.0778
-41.8 |  |  |  |  | 8.1662
-41.7 |  |  |  |  | 8.2728
-41.6 |  |  |  |  | 8.3063
-41.5 |  |  |  |  | 8.4
-41.4 |  |  |  |  | 8.4
-41.3 |  |  |  |  | 8.4949
-41.2 |  |  |  |  | 8.6
-41.1 |  |  |  |  | 8.6298
-41 |  |  |  |  | 8.7409
-40.9 |  |  |  | 5.2 | 8.8419
-40.8 |  |  |  | 5.3 | 8.9
-40.7 |  |  |  | 5.3011 | 8.9512
-40.6 |  |  |  | 5.4 | 9.0691
-40.5 |  |  |  | 5.4031 | 9.1
-40.4 |  |  |  | 5.5 | 9.2041
-40.3 |  |  |  | 5.5 | 9.304
-40.2 |  |  |  | 5.5919 | 9.3826
-40.1 |  |  |  | 5.6 | 9.4786
-40 |  |  |  | 5.6939 | 9.5929
-39.9 |  |  |  | 5.7107 | 9.6949
-39.8 |  |  |  | 5.8 | 9.8102
-39.7 |  |  |  | 5.8256 | 9.9112
-39.6 |  |  |  | 5.9 | 10.0122
-39.5 |  |  |  | 5.9322 | 10.1
-39.4 |  |  |  | 6 | 10.1667
-39.3 |  |  |  | 6.0582 | 10.3384
-39.2 |  |  |  | 6.1 | 10.4268
-39.1 |  |  |  | 6.1744 | 10.5
-39 |  |  |  | 6.2755 | 10.5214
-38.9 |  |  |  | 6.3 | 10.6392
-38.8 |  |  |  | 6.3904 | 10.7489
-38.7 |  |  |  | 6.4928 | 10.9
-38.6 |  |  |  | 6.5 | 11.051
-38.5 |  |  |  | 6.5948 | 11.1607
-38.4 |  |  |  | 6.6839 | 11.2673
-38.3 |  |  |  | 6.7826 | 11.3683
-38.2 |  |  |  | 6.8 | 11.5618
-38.1 |  |  | 3.5275 | 6.8846 | 11.674
-38 |  |  | 3.6 | 6.9857 | 11.7999
-37.9 |  |  | 3.6 | 7.1 | 11.9153
-37.8 |  |  | 3.6305 | 7.1019 | 12.019
-37.7 |  |  | 3.7 | 7.2026 | 12.1315
-37.6 |  |  | 3.7 | 7.291 | 12.2651
-37.5 |  |  | 3.7621 | 7.4 | 12.4294
-37.4 |  |  | 3.8 | 7.4203 | 12.5406
-37.3 |  |  | 3.8 | 7.5213 | 12.7213
-37.2 |  |  | 3.857 | 7.6223 | 12.8391
-37.1 |  |  | 3.9 | 7.7233 | 13.0106
-37 |  |  | 3.9529 | 7.8243 | 13.1426
-36.9 |  |  | 4 | 7.9254 | 13.3111
-36.8 |  |  | 4.0642 | 8.0264 | 13.4212
-36.7 |  | 2.7 | 4.1 | 8.124 | 13.5979
-36.6 |  | 2.7832 | 4.1713 | 8.2141 | 13.733
-36.5 |  | 2.8 | 4.2 | 8.3152 | 13.9458
-36.4 |  | 2.8 | 4.2733 | 8.4188 | 14.1034
-36.3 |  | 2.8041 | 4.3 | 8.5367 | 14.28
-36.2 |  | 2.9 | 4.3712 | 8.6545 | 14.4792
-36.1 |  | 2.9 | 4.4 | 8.762 | 14.6115
-36 |  | 2.9 | 4.4236 | 8.8631 | 14.8202
-35.9 |  | 3 | 4.5 | 8.9641 | 14.9495
-35.8 |  | 3 | 4.5426 | 9.0652 | 15.1853
-35.7 | 2.2 | 3 | 4.6 | 9.2323 | 15.3605
-35.6 | 2.3 | 3.0941 | 4.7 | 9.3587 | 15.5567
-35.5 | 2.3 | 3.1 | 4.7396 | 9.4538 | 15.822
-35.4 | 2.3 | 3.1 | 4.8 | 9.5768 | 16.014
-35.3 | 2.3 | 3.2 | 4.8274 | 9.6986 | 16.1718
-35.2 | 2.4 | 3.2 | 4.9331 | 9.814 | 16.4246
-35.1 | 2.4 | 3.2811 | 5 | 9.9151 | 16.6014
-35 | 2.4 | 3.3 | 5.0446 | 10.0321 | 16.7781
-34.9 | 2.4 | 3.3 | 5.1 | 10.2199 | 17.0655
-34.8 | 2.5 | 3.4 | 5.1467 | 10.3324 | 17.309
-34.7 | 2.5 | 3.4 | 5.2477 | 10.4585 | 17.5627
-34.6 | 2.5 | 3.4 | 5.3 | 10.6235 | 17.8279
-34.5 | 2.558 | 3.5 | 5.406 | 10.7413 | 18.0931
-34.4 | 2.6 | 3.5 | 5.4944 | 10.9014 | 18.3518
-34.3 | 2.6 | 3.6 | 5.5 | 11.0724 | 18.5875
-34.2 | 2.6 | 3.6 | 5.6283 | 11.2899 | 18.9346
-34.1 | 2.6942 | 3.7 | 5.7346 | 11.3967 | 19.2563
-34 | 2.7 | 3.7 | 5.8 | 11.5278 | 19.6398
-33.9 | 2.7 | 3.7818 | 5.8273 | 11.7318 | 20.0817
-33.8 | 2.7595 | 3.8 | 5.933 | 11.8993 | 20.5649
-33.7 | 2.8 | 3.9 | 6.061 | 12.0675 | 21.1541
-33.6 | 2.8 | 3.9 | 6.1017 | 12.2048 | 21.812
-33.5 | 2.8839 | 3.9866 | 6.2027 | 12.4548 | 22.5841
-33.4 | 2.9 | 4.0065 | 6.3044 | 12.6755 | 23.5931
-33.3 | 2.9334 | 4.1 | 6.4223 | 12.881 | 24.7932
-33.2 | 3 | 4.1881 | 6.5343 | 12.9915 | 26.1516
-33.1 | 3.0079 | 4.2079 | 6.6413 | 13.2296 | 27.9317
-33 | 3.1 | 4.3 | 6.7 | 13.3887 | 29.8971
-32.9 | 3.1 | 4.3603 | 6.7724 | 13.6121 | 32.236
-32.8 | 3.2 | 4.4 | 6.9099 | 13.8768 | 35.0345
-32.7 | 3.2 | 4.496 | 7.0254 | 14.1182 | 38.1659
-32.6 | 3.2958 | 4.5366 | 7.2 | 14.3809 | 41.3456
-32.5 | 3.3 | 4.6 | 7.2415 | 14.5968 | 44.6276
-32.4 | 3.3 | 4.6995 | 7.3425 | 14.8324 | 48.0011
-32.3 | 3.4 | 4.8 | 7.4871 | 15.0442 | 50.9076
-32.2 | 3.4519 | 4.8779 | 7.6445 | 15.2463 | 53.5944
-32.1 | 3.5 | 4.9198 | 7.7398 | 15.5224 | 55.6262
-32 | 3.5314 | 5 | 7.8644 | 15.8255 | 57.2923
-31.9 | 3.6 | 5.0555 | 8.0332 | 16.0998 | 58.6104
-31.8 | 3.7 | 5.188 | 8.1343 | 16.3885 | 59.5448
-31.7 | 3.7 | 5.2245 | 8.2412 | 16.7554 | 60.3303
-31.6 | 3.8 | 5.3508 | 8.4012 | 17.0945 | 61.0541
-31.5 | 3.8903 | 5.4 | 8.5516 | 17.4637 | 61.7101
-31.4 | 3.9 | 5.4947 | 8.7053 | 17.9348 | 62.3687
-31.3 | 3.9951 | 5.6459 | 8.8626 | 18.3973 | 63.0757
-31.2 | 4.0365 | 5.7 | 9.038 | 18.9828 | 63.783
-31.1 | 4.1 | 5.798 | 9.17 | 19.69 | 64.4035
-31 | 4.2 | 5.9194 | 9.342 | 20.6101 | 65.1716
-30.9 | 4.2808 | 6 | 9.544 | 21.7642 | 65.9039
-30.8 | 4.3 | 6.0852 | 9.7703 | 22.9764 | 66.6111
-30.7 | 4.3636 | 6.2037 | 9.8773 | 24.4597 | 67.2942
-30.6 | 4.5 | 6.3376 | 10.0501 | 26.6109 | 67.9254
-30.5 | 4.6 | 6.4 | 10.2774 | 29.0512 | 68.6322
-30.4 | 4.6479 | 6.5066 | 10.4826 | 32.1307 | 69.2595
-30.3 | 4.7 | 6.6244 | 10.6848 | 35.5422 | 69.8101
-30.2 | 4.8 | 6.7508 | 10.9179 | 38.9491 | 70.3403
-30.1 | 4.8435 | 6.8768 | 11.1535 | 43.2919 | 70.8706
-30 | 4.992 | 6.9947 | 11.367 | 47.3196 | 71.3786
-29.9 | 5.055 | 7.1438 | 11.5501 | 51.6048 | 71.8562
-29.8 | 5.1803 | 7.2965 | 11.7782 | 55.2175 | 72.3867
-29.7 | 5.2316 | 7.4178 | 12.0541 | 58.314 | 72.783
-29.6 | 5.3593 | 7.5593 | 12.2982 | 60.9033 | 73.1483
-29.5 | 5.4 | 7.7007 | 12.6012 | 62.8416 | 73.5307
-29.4 | 5.5221 | 7.8843 | 12.8369 | 64.2481 | 73.8826
-29.3 | 5.6635 | 8 | 13.1218 | 65.4236 | 74.1968
-29.2 | 5.705 | 8.2041 | 13.4248 | 66.3668 | 74.5093
-29.1 | 5.8464 | 8.3528 | 13.766 | 67.261 | 74.7745
-29 | 5.9731 | 8.5565 | 14.1653 | 68.0592 | 75.1259
-28.9 | 6.0115 | 8.7153 | 14.5693 | 68.8882 | 75.5244
-28.8 | 6.1706 | 8.9256 | 15.1379 | 69.7973 | 75.6844
-28.7 | 6.2934 | 9.1041 | 15.6601 | 70.7064 | 75.9802
-28.6 | 6.4 | 9.2779 | 16.3806 | 71.5915 | 76.2559
-28.5 | 6.5583 | 9.4915 | 17.2919 | 72.3996 | 76.4436
-28.4 | 6.6163 | 9.6727 | 18.4632 | 73.2422 | 76.7409
-28.3 | 6.7577 | 9.9296 | 19.8778 | 74.0525 | 77.0943
-28.2 | 6.8992 | 10.1653 | 21.6313 | 74.7518 | 77.2988
-28.1 | 7.0207 | 10.4207 | 23.997 | 75.6319 | 77.4344
-28 | 7.1517 | 10.6551 | 26.9747 | 76.38 | 77.684
-27.9 | 7.3392 | 10.9068 | 30.6656 | 77.0534 | 77.9565
-27.8 | 7.5048 | 11.2121 | 34.8804 | 77.7318 | 78.1927
-27.7 | 7.6463 | 11.5657 | 40.1037 | 78.3316 | 78.4105
-27.6 | 7.7731 | 12.0508 | 45.0167 | 78.8935 | 78.6879
-27.5 | 7.9819 | 12.5383 | 50.8808 | 79.388 | 78.9114
-27.4 | 8.1133 | 13.1527 | 56.6864 | 79.8586 | 79.1176
-27.3 | 8.272 | 13.8597 | 61.7692 | 80.2626 | 79.3533
-27.2 | 8.5112 | 14.9004 | 66.0469 | 80.6334 | 79.5889
-27.1 | 8.658 | 16.3231 | 69.4709 | 81.0245 | 79.7685
-27 | 8.8563 | 18.0523 | 71.9655 | 81.3905 | 80.0491
-26.9 | 9.1354 | 20.2386 | 73.6493 | 81.653 | 80.296
-26.8 | 9.4182 | 23.254 | 74.9513 | 82.0774 | 80.4318
-26.7 | 9.7014 | 27.0472 | 76.1322 | 82.4155 | 80.6289
-26.6 | 10.1096 | 31.4792 | 77.2145 | 82.6457 | 80.7598
-26.5 | 10.7638 | 37.6416 | 78.3122 | 82.9215 | 80.9722
-26.4 | 11.4236 | 44.0258 | 79.4233 | 83.2495 | 81.2296
-26.3 | 12.44 | 50.685 | 80.5344 | 83.5324 | 81.4769
-26.2 | 13.816 | 57.3894 | 81.6456 | 83.7164 | 81.7055
-26.1 | 15.7044 | 64.1198 | 82.8081 | 84.0129 | 81.8151
-26 | 18.2238 | 69.3661 | 83.8313 | 84.2209 | 82.0507
-25.9 | 21.6159 | 74.0164 | 84.889 | 84.5038 | 82.2399
-25.8 | 25.6431 | 77.1989 | 85.9374 | 84.6777 | 82.4037
-25.7 | 30.7993 | 79.5215 | 86.9227 | 84.9294 | 82.6026
-25.6 | 37.6034 | 80.8935 | 87.8525 | 85.2403 | 82.8401
-25.5 | 45.5281 | 82.3605 | 88.6321 | 85.4554 | 82.9973
-25.4 | 53.0488 | 83.5245 | 89.2945 | 85.6127 | 83.1737
-25.3 | 61.6475 | 84.8698 | 90.0019 | 85.8149 | 83.3004
-25.2 | 68.4919 | 86.0835 | 90.6364 | 86.0534 | 83.4363
-25.1 | 74.7193 | 87.4322 | 91.1439 | 86.2903 | 83.6618
-25 | 79.1283 | 88.6467 | 91.6616 | 86.4539 | 83.8638
-24.9 | 82.2226 | 89.9689 | 92.0536 | 86.6435 | 84.0329
-24.8 | 84.5749 | 91.188 | 92.5024 | 86.8535 | 84.1792
-24.7 | 85.9888 | 92.4173 | 92.8089 | 87.0556 | 84.3984
-24.6 | 87.3033 | 93.5028 | 93.0258 | 87.2172 | 84.5879
-24.5 | 88.6141 | 94.5426 | 93.4726 | 87.4227 | 84.7647
-24.4 | 89.8389 | 95.5938 | 93.811 | 87.5665 | 84.9185
-24.3 | 91.1896 | 96.4186 | 93.9495 | 87.8244 | 84.997
-24.2 | 92.5808 | 97.1168 | 94.1773 | 88.0372 | 85.2552
-24.1 | 93.974 | 97.7873 | 94.4535 | 88.2392 | 85.4306
-24 | 95.3296 | 98.3471 | 94.6278 | 88.4412 | 85.6314
-23.9 | 96.8024 | 98.7343 | 94.7576 | 88.5716 | 85.8002
-23.8 | 97.9633 | 99.1542 | 94.9695 | 88.6848 | 85.9363
-23.7 | 98.9822 | 99.5078 | 95.1631 | 88.8539 | 86.136
-23.6 | 100.0673 | 99.8614 | 95.4032 | 89.0045 | 86.2538
-23.5 | 100.916 | 100.08 | 95.5059 | 89.1919 | 86.4075
-23.4 | 101.6372 | 100.2686 | 95.6474 | 89.3874 | 86.5675
-23.3 | 102.3444 | 100.5777 | 95.8481 | 89.5289 | 86.6778
-23.2 | 102.9306 | 100.7085 | 95.9788 | 89.6703 | 86.8576
-23.1 | 103.3033 | 100.8528 | 96.1596 | 89.8116 | 86.9698
-23 | 103.6792 | 101.053 | 96.2707 | 89.953 | 87.1164
-22.9 | 103.9345 | 101.1788 | 96.4575 | 90.1575 | 87.2414
-22.8 | 104.1898 | 101.2966 | 96.5828 | 90.2966 | 87.3399
-22.7 | 104.4174 | 101.4373 | 96.6978 | 90.3981 | 87.497
-22.6 | 104.5588 | 101.5656 | 96.9277 | 90.5982 | 87.7602
-22.5 | 104.7002 | 101.6835 | 97.0402 | 90.6876 | 87.9503
-22.4 | 104.8617 | 101.8 | 97.1816 | 90.8737 | 88.127
-22.3 | 104.903 | 101.8358 | 97.3021 | 90.9879 | 88.2692
-22.2 | 105 | 101.9645 | 97.4032 | 91.1778 | 88.3653
-22.1 | 105.0716 | 102.0059 | 97.5059 | 91.3049 | 88.5073
-22 | 105.2 | 102.1472 | 97.6675 | 91.439 | 88.6481
-21.9 | 105.2239 | 102.2 | 97.8 | 91.6205 | 88.743
-21.8 | 105.3 | 102.2918 | 97.8751 | 91.7215 | 88.8627
-21.7 | 105.3929 | 102.4097 | 97.9083 | 91.8225 | 89.0263
-21.6 | 105.4 | 102.5 | 98.0093 | 91.933 | 89.1379
-21.5 | 105.4 | 102.5787 | 98.118 | 92.062 | 89.2907
-21.4 | 105.4799 | 102.6 | 98.2949 | 92.1959 | 89.4948
-21.3 | 105.5 | 102.7 | 98.3981 | 92.3216 | 89.6358
-21.2 | 105.5786 | 102.7 | 98.4991 | 92.4984 | 89.7429
-21.1 | 105.6 | 102.8001 | 98.6 | 92.6501 | 89.9001
-21 | 105.6 | 102.9 | 98.6019 | 92.7 | 89.9786
-20.9 | 105.7 | 102.9 | 98.7 | 92.8358 | 90.0644
-20.8 | 105.7 | 102.9044 | 98.7745 | 92.946 | 90.1603
-20.7 | 105.7 | 103 | 98.9058 | 93.1097 | 90.2536
-20.6 | 105.7873 | 103.056 | 99 | 93.2 | 90.384
-20.5 | 105.8 | 103.1 | 99.049 | 93.3086 | 90.554
-20.4 | 105.8 | 103.2 | 99.15 | 93.4417 | 90.7056
-20.3 | 105.8 | 103.2 | 99.2714 | 93.5596 | 90.7841
-20.2 | 105.8 | 103.2607 | 99.3 | 93.7 | 90.9254
-20.1 | 105.8 | 103.3 | 99.4 | 93.7543 | 91.1061
-20 | 105.8696 | 103.4 | 99.4631 | 93.9 | 91.2541
-19.9 | 105.9 | 103.4 | 99.5 | 93.9266 | 91.3
-19.8 | 105.9 | 103.4 | 99.6 | 94.0552 | 91.3821
-19.7 | 105.9 | 103.4667 | 99.65 | 94.2 | 91.5
-19.6 | 105.9 | 103.5 | 99.7678 | 94.2259 | 91.5028
-19.5 | 105.9 | 103.5 | 99.8 | 94.3163 | 91.7357
-19.4 | 105.9 | 103.6 | 99.8744 | 94.4174 | 91.8401
-19.3 | 105.9 | 103.6 | 99.9 | 94.5184 | 91.9571
-19.2 | 105.9 | 103.6 | 99.9 | 94.6194 | 92.1226
-19.1 | 106 | 103.6737 | 99.9918 | 94.7204 | 92.2
-19 | 106 | 103.7 | 100.0812 | 94.8187 | 92.2375
-18.9 | 106 | 103.7 | 100.1796 | 94.9094 | 92.4063
-18.8 | 106 | 103.7 | 100.2 | 95.0234 | 92.4849
-18.7 | 106 | 103.7 | 100.2816 | 95.1285 | 92.5
-18.6 | 106 | 103.7797 | 100.3 | 95.2 | 92.7098
-18.5 | 106 | 103.8 | 100.3693 | 95.2407 | 92.7981
-18.4 | 106 | 103.8 | 100.4 | 95.3417 | 92.8769
-18.3 | 106 | 103.8 | 100.4833 | 95.4499 | 92.9624
-18.2 | 106 | 103.8 | 100.5 | 95.5 | 93.0508
-18.1 | 106 | 103.8 | 100.5627 | 95.5733 | 93.2046
-18 | 106 | 103.8642 | 100.6 | 95.6744 | 93.3526
-17.9 | 105.9 | 103.9 | 100.6 | 95.788 | 93.4469
-17.8 | 105.9 | 103.9 | 100.605 | 95.9 | 93.5372
-17.7 | 105.9 | 103.9 | 100.7 | 95.9 | 93.6178
-17.6 | 105.9 | 103.9 | 100.7071 | 95.9898 | 93.7071
-17.5 | 105.9 | 103.9 | 100.8 | 96.1261 | 93.8094
-17.4 | 105.9 | 103.9 | 100.8 | 96.2 | 93.9272
-17.3 | 105.9 | 103.9 | 100.8243 | 96.2784 | 94
-17.2 | 105.9 | 103.9 | 100.9 | 96.3962 | 94.0309
-17.1 | 105.9 | 103.9 | 100.9 | 96.4369 | 94.1106
-17 | 105.9 | 103.9 | 101 | 96.5 | 94.1989
-16.9 | 105.8 | 103.9 | 101.0598 | 96.5999 | 94.3141
-16.8 | 105.8 | 103.9 | 101.1 | 96.7 | 94.4152
-16.7 | 105.8 | 103.9 | 101.1 | 96.7022 | 94.5126
-16.6 | 105.8 | 103.9 | 101.1 | 96.8 | 94.5911
-16.5 | 105.8 | 103.9 | 101.1 | 96.8 | 94.6784
-16.4 | 105.8 | 103.9 | 101.1 | 96.8558 | 94.7
-16.3 | 105.7 | 103.9 | 101.1 | 96.9 | 94.7854
-16.2 | 105.7 | 103.9 | 101.1897 | 96.9784 | 94.9277
-16.1 | 105.7 | 103.9 | 101.2 | 97.0794 | 95.0062
-16 | 105.7 | 103.9 | 101.2 | 97.1804 | 95.0848
-15.9 | 105.6217 | 103.9 | 101.2 | 97.3 | 95.1713
-15.8 | 105.6 | 103.9 | 101.2221 | 97.3 | 95.253
-15.7 | 105.6 | 103.9 | 101.3 | 97.3 | 95.3355
-15.6 | 105.6 | 103.9 | 101.3152 | 97.3986 | 95.4274
-15.5 | 105.6 | 103.9 | 101.4 | 97.4 | 95.5331
-15.4 | 105.6 | 103.9 | 101.4 | 97.5 | 95.6
-15.3 | 105.5645 | 103.9 | 101.4 | 97.5161 | 95.6304
-15.2 | 105.5 | 103.8258 | 101.4 | 97.6171 | 95.7314
-15.1 | 105.5 | 103.8 | 101.4 | 97.7 | 95.8
-15 | 105.4332 | 103.8 | 101.4 | 97.7 | 95.862
-14.9 | 105.4 | 103.8 | 101.4 | 97.7487 | 96.0102
-14.8 | 105.4 | 103.8 | 101.4 | 97.858 | 96.1
-14.7 | 105.3074 | 103.8 | 101.4 | 97.9 | 96.1395
-14.6 | 105.3 | 103.8 | 101.4 | 97.9661 | 96.2203
-14.5 | 105.3 | 103.8 | 101.4616 | 98.0587 | 96.3077
-14.4 | 105.3 | 103.7 | 101.5 | 98.1 | 96.3863
-14.3 | 105.2433 | 103.7 | 101.5 | 98.1548 | 96.4649
-14.2 | 105.2 | 103.7 | 101.5 | 98.2 | 96.5
-14.1 | 105.2 | 103.6 | 101.5 | 98.2497 | 96.5425
-14 | 105.1158 | 103.6 | 101.5 | 98.3 | 96.6
-13.9 | 105.0646 | 103.6 | 101.5 | 98.352 | 96.639
-13.8 | 105 | 103.6 | 101.5 | 98.4 | 96.7244
-13.7 | 105 | 103.6 | 101.5467 | 98.5 | 96.8033
-13.6 | 105 | 103.6 | 101.6 | 98.5 | 96.8917
-13.5 | 104.9265 | 103.6 | 101.6 | 98.5901 | 96.9915
-13.4 | 104.9 | 103.6 | 101.6 | 98.6 | 97.072
-13.3 | 104.9 | 103.5491 | 101.6 | 98.6 | 97.1
-13.2 | 104.9 | 103.5 | 101.6124 | 98.6828 | 97.1402
-13.1 | 104.8062 | 103.5 | 101.7 | 98.7813 | 97.2241
-13 | 104.7706 | 103.5 | 101.7 | 98.8 | 97.3
-12.9 | 104.7 | 103.5 | 101.7 | 98.8 | 97.3208
-12.8 | 104.7 | 103.5 | 101.7 | 98.8 | 97.4975
-12.7 | 104.7 | 103.5 | 101.7 | 98.8996 | 97.5775
-12.6 | 104.7 | 103.5 | 101.7 | 99 | 97.6631
-12.5 | 104.6 | 103.5 | 101.7 | 99.0019 | 97.7
-12.4 | 104.6 | 103.4 | 101.7 | 99.1 | 97.7312
-12.3 | 104.6 | 103.4 | 101.7 | 99.1044 | 97.8
-12.2 | 104.6 | 103.4 | 101.7555 | 99.2 | 97.8037
-12.1 | 104.5767 | 103.4 | 101.8 | 99.2 | 97.8822
-12 | 104.5 | 103.4 | 101.8 | 99.2 | 97.9684
-11.9 | 104.4692 | 103.4 | 101.8 | 99.3 | 98
-11.8 | 104.4 | 103.4 | 101.8 | 99.3 | 98.0374
-11.7 | 104.4 | 103.4 | 101.8 | 99.3 | 98.1384
-11.6 | 104.4 | 103.3 | 101.8 | 99.3536 | 98.2
-11.5 | 104.4 | 103.3 | 101.8 | 99.4 | 98.2304
-11.4 | 104.3017 | 103.3 | 101.8 | 99.4 | 98.3
-11.3 | 104.3 | 103.3 | 101.8 | 99.4 | 98.3567
-11.2 | 104.3 | 103.2 | 101.8 | 99.4863 | 98.4
-11.1 | 104.3 | 103.2 | 101.8 | 99.5 | 98.4514
-11 | 104.3 | 103.2 | 101.8 | 99.5863 | 98.5398
-10.9 | 104.2 | 103.2 | 101.8 | 99.6 | 98.6
-10.8 | 104.2 | 103.2 | 101.8 | 99.6054 | 98.6165
-10.7 | 104.2 | 103.1434 | 101.8 | 99.7 | 98.7
-10.6 | 104.1422 | 103.1 | 101.8 | 99.7353 | 98.7067
-10.5 | 104.1 | 103.1 | 101.8 | 99.8363 | 98.806
-10.4 | 104.1 | 103.1 | 101.8 | 99.9 | 98.8845
-10.3 | 104.0264 | 103.1 | 101.8 | 99.9 | 98.9
-10.2 | 104 | 103.1 | 101.8 | 99.9 | 98.9469
-10.1 | 103.9 | 103 | 101.8 | 99.9 | 99
-10 | 103.9 | 103 | 101.8 | 99.9816 | 99.0237
-9.9 | 103.8607 | 103 | 101.8 | 100 | 99.1107
-9.8 | 103.8 | 103 | 101.8 | 100 | 99.1893
-9.7 | 103.8 | 103 | 101.8 | 100 | 99.2
-9.6 | 103.7565 | 103 | 101.8 | 100.0696 | 99.2
-9.5 | 103.7 | 103 | 101.8 | 100.1 | 99.2875
-9.4 | 103.7 | 102.978 | 101.8 | 100.1 | 99.3
-9.3 | 103.7 | 102.9 | 101.8 | 100.1199 | 99.3
-9.2 | 103.7 | 102.9 | 101.8 | 100.2 | 99.3496
-9.1 | 103.6493 | 102.9 | 101.8 | 100.2 | 99.4
-9 | 103.6 | 102.8 | 101.8 | 100.2086 | 99.4601
-8.9 | 103.6 | 102.8279 | 101.8 | 100.3 | 99.5
-8.8 | 103.6 | 102.8542 | 101.8 | 100.3 | 99.5093
-8.7 | 103.5037 | 102.8 | 101.8 | 100.3 | 99.5977
-8.6 | 103.5 | 102.8 | 101.8 | 100.3 | 99.6
-8.5 | 103.5 | 102.8 | 101.8 | 100.3566 | 99.6
-8.4 | 103.5 | 102.8 | 101.8 | 100.4 | 99.6629
-8.3 | 103.4725 | 102.8 | 101.8 | 100.4 | 99.7
-8.2 | 103.4 | 102.8 | 101.8 | 100.4 | 99.731
-8.1 | 103.3551 | 102.7126 | 101.8 | 100.4749 | 99.8
-8 | 103.3 | 102.7 | 101.8 | 100.5 | 99.8035
-7.9 | 103.3 | 102.7 | 101.8 | 100.5 | 99.8821
-7.8 | 103.3 | 102.7 | 101.8 | 100.5 | 99.9
-7.7 | 103.2079 | 102.7 | 101.7 | 100.5421 | 99.9
-7.6 | 103.2 | 102.668 | 101.7 | 100.6 | 99.9
-7.5 | 103.1266 | 102.6 | 101.7 | 100.6 | 99.963
-7.4 | 103.1 | 102.6 | 101.7 | 100.6 | 100
-7.3 | 103.0638 | 102.5698 | 101.7687 | 100.6 | 100
-7.2 | 103 | 102.5 | 101.8 | 100.6 | 100
-7.1 | 103 | 102.5 | 101.8 | 100.6422 | 100
-7 | 103 | 102.4 | 101.7282 | 100.7 | 100.0289
-6.9 | 102.9226 | 102.4 | 101.7 | 100.7 | 100.1
-6.8 | 102.9 | 102.4 | 101.7 | 100.7 | 100.1
-6.7 | 102.9 | 102.4 | 101.7 | 100.7 | 100.1606
-6.6 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
-6.5 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
-6.4 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
-6.3 | 102.8 | 102.4 | 101.7 | 100.7 | 100.2
-6.2 | 102.8 | 102.3234 | 101.7 | 100.7 | 100.2
-6.1 | 102.7056 | 102.3 | 101.7 | 100.7 | 100.2629
-6 | 102.7 | 102.3 | 101.6 | 100.7 | 100.3
-5.9 | 102.6365 | 102.3 | 101.6 | 100.7 | 100.3
-5.8 | 102.6 | 102.2 | 101.6 | 100.7 | 100.3
-5.7 | 102.6 | 102.2 | 101.6 | 100.7 | 100.3
-5.6 | 102.5995 | 102.1663 | 101.6 | 100.7 | 100.3
-5.5 | 102.5 | 102.1 | 101.5273 | 100.7 | 100.3
-5.4 | 102.5 | 102.1 | 101.5 | 100.7 | 100.3
-5.3 | 102.4 | 102.0922 | 101.4967 | 100.7 | 100.3
-5.2 | 102.4 | 102 | 101.4 | 100.7 | 100.3
-5.1 | 102.3 | 101.9104 | 101.4 | 100.7 | 100.3
-5 | 102.2426 | 101.8759 | 101.4 | 100.665 | 100.3
-4.9 | 102.2 | 101.8 | 101.3069 | 100.6 | 100.3
-4.8 | 102.1735 | 101.7402 | 101.3 | 100.5569 | 100.3
-4.7 | 102.1 | 101.7 | 101.2 | 100.5 | 100.3
-4.6 | 102.0254 | 101.7 | 101.2 | 100.5 | 100.3
-4.5 | 101.9033 | 101.644 | 101.104 | 100.4 | 100.3
-4.4 | 101.8625 | 101.5188 | 101.0825 | 100.4 | 100.3
-4.3 | 101.8 | 101.4009 | 100.9411 | 100.4 | 100.3
-4.2 | 101.6997 | 101.3797 | 100.9 | 100.4169 | 100.3
-4.1 | 101.5652 | 101.3 | 100.8486 | 100.5 | 100.3
-4 | 101.4369 | 101.214 | 100.8 | 100.5 | 100.3
-3.9 | 101.3 | 101.1 | 100.8 | 100.5 | 100.3
-3.8 | 101.3 | 101.034 | 100.8 | 100.5 | 100.3
-3.7 | 101.2326 | 101 | 100.7518 | 100.5 | 100.4
-3.6 | 101.2 | 101 | 100.7 | 100.426 | 100.3794
-3.5 | 101.1498 | 100.9297 | 100.7 | 100.4 | 100.3
-3.4 | 101.1 | 100.9 | 100.7 | 100.4 | 100.3
-3.3 | 101.0669 | 100.9 | 100.7 | 100.4109 | 100.3
-3.2 | 101 | 100.9 | 100.7 | 100.5 | 100.3
-3.1 | 101 | 100.9 | 100.6241 | 100.5 | 100.3
-3 | 101 | 100.8026 | 100.6 | 100.5 | 100.3
-2.9 | 101 | 100.8 | 100.6 | 100.5 | 100.3
-2.8 | 101 | 100.8 | 100.6859 | 100.5 | 100.3
-2.7 | 101 | 100.8 | 100.7 | 100.4384 | 100.4
-2.6 | 100.9 | 100.8 | 100.7 | 100.403 | 100.4
-2.5 | 100.9 | 100.7796 | 100.7 | 100.4555 | 100.4
-2.4 | 100.8141 | 100.7 | 100.6341 | 100.4 | 100.4
-2.3 | 100.8 | 100.7 | 100.6 | 100.4 | 100.3293
-2.2 | 100.8 | 100.7 | 100.6 | 100.4 | 100.3
-2.1 | 100.7082 | 100.7 | 100.5299 | 100.4 | 100.3
-2 | 100.7 | 100.6404 | 100.5 | 100.4 | 100.3
-1.9 | 100.7 | 100.6 | 100.5 | 100.3336 | 100.3
-1.8 | 100.6 | 100.6 | 100.5 | 100.3 | 100.3
-1.7 | 100.6 | 100.6 | 100.4368 | 100.3 | 100.2868
-1.6 | 100.5689 | 100.5523 | 100.4 | 100.3 | 100.2
-1.5 | 100.5 | 100.4344 | 100.3295 | 100.3 | 100.2
-1.4 | 100.5 | 100.4 | 100.3 | 100.2428 | 100.2
-1.3 | 100.4 | 100.4 | 100.3 | 100.2 | 100.1821
-1.2 | 100.4 | 100.4 | 100.2265 | 100.2 | 100.1
-1.1 | 100.313 | 100.2964 | 100.113 | 100.2 | 100.1358
-1 | 100.3 | 100.2 | 100.1 | 100.2 | 100.2
-0.9 | 100.2128 | 100.144 | 100.1 | 100.2 | 100.1955
-0.8 | 100.2 | 100.1 | 100.0714 | 100.1081 | 100.1071
-0.7 | 100.1499 | 100.0699 | 100 | 100.0937 | 100.0187
-0.6 | 100.1 | 100 | 100 | 100.0053 | 100
-0.5 | 100.1 | 100 | 100 | 100 | 100
-0.4 | 100.1 | 100 | 100 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100.0202 | 100 | 100 | 100 | 100
0.4 | 100.1 | 100 | 100 | 100 | 100
0.5 | 100.1 | 100 | 100 | 100 | 100
0.6 | 100.1 | 100 | 100 | 100 | 100
0.7 | 100.175 | 100.0499 | 100 | 100.0642 | 100.0722
0.8 | 100.2 | 100.1 | 100.1 | 100.1761 | 100.1
0.9 | 100.2 | 100.2 | 100.1 | 100.2 | 100.1516
1 | 100.2742 | 100.2452 | 100.1 | 100.2 | 100.2
1.1 | 100.3 | 100.3 | 100.1556 | 100.2 | 100.1888
1.2 | 100.3713 | 100.3971 | 100.2 | 100.2 | 100.1
1.3 | 100.4185 | 100.4 | 100.2703 | 100.2 | 100.1321
1.4 | 100.5 | 100.4 | 100.3 | 100.2 | 100.2
1.5 | 100.5 | 100.5 | 100.3724 | 100.3 | 100.2
1.6 | 100.5 | 100.5227 | 100.4 | 100.3 | 100.2305
1.7 | 100.6 | 100.6 | 100.4 | 100.3 | 100.3
1.8 | 100.6 | 100.6 | 100.488 | 100.3 | 100.3
1.9 | 100.7 | 100.6 | 100.5 | 100.3907 | 100.3
2 | 100.7 | 100.7 | 100.5 | 100.4 | 100.3
2.1 | 100.7249 | 100.7 | 100.5 | 100.4 | 100.3
2.2 | 100.8 | 100.7 | 100.6 | 100.4 | 100.3
2.3 | 100.8 | 100.7 | 100.6 | 100.4 | 100.3
2.4 | 100.8341 | 100.7 | 100.6741 | 100.4 | 100.3713
2.5 | 100.9 | 100.7463 | 100.7 | 100.5 | 100.4
2.6 | 100.9 | 100.8 | 100.7 | 100.443 | 100.4
2.7 | 100.9653 | 100.8 | 100.7 | 100.4 | 100.4
2.8 | 101 | 100.8 | 100.7 | 100.5 | 100.3335
2.9 | 101 | 100.8 | 100.6492 | 100.5 | 100.3
3 | 101 | 100.8 | 100.6 | 100.5 | 100.3
3.1 | 101 | 100.9 | 100.6 | 100.5 | 100.3
3.2 | 101 | 100.9 | 100.6879 | 100.4675 | 100.3
3.3 | 101.1 | 100.9 | 100.7 | 100.4 | 100.3
3.4 | 101.1 | 100.9 | 100.7 | 100.4 | 100.3
3.5 | 101.1872 | 100.9 | 100.7 | 100.4 | 100.3387
3.6 | 101.2 | 101 | 100.7 | 100.4 | 100.4
3.7 | 101.2526 | 101 | 100.7947 | 100.5 | 100.3562
3.8 | 101.3 | 101.095 | 100.8 | 100.5 | 100.3
3.9 | 101.3154 | 101.1 | 100.8 | 100.5 | 100.3
4 | 101.4 | 101.1807 | 100.8 | 100.5 | 100.3
4.1 | 101.4983 | 101.2986 | 100.8 | 100.5 | 100.3
4.2 | 101.7164 | 101.3497 | 100.9 | 100.4503 | 100.3
4.3 | 101.8 | 101.4 | 100.9811 | 100.4 | 100.3
4.4 | 101.8825 | 101.5021 | 101.0188 | 100.4 | 100.3
4.5 | 101.924 | 101.62 | 101.1 | 100.444 | 100.3
4.6 | 102 | 101.7 | 101.1711 | 100.5 | 100.3
4.7 | 102.0723 | 101.7 | 101.2 | 100.5191 | 100.3
4.8 | 102.1902 | 101.8 | 101.2344 | 100.6 | 100.3
4.9 | 102.2 | 101.8 | 101.3 | 100.6 | 100.3
5 | 102.2592 | 101.8592 | 101.3365 | 100.6222 | 100.3
5.1 | 102.3 | 101.9 | 101.4 | 100.7 | 100.3
5.2 | 102.3339 | 101.9949 | 101.4 | 100.7 | 100.3
5.3 | 102.4 | 102.0442 | 101.4395 | 100.7 | 100.3
5.4 | 102.5 | 102.1 | 101.5 | 100.7 | 100.3
5.5 | 102.5 | 102.1 | 101.5 | 100.7 | 100.3
5.6 | 102.6 | 102.1396 | 101.5711 | 100.7 | 100.3
5.7 | 102.6 | 102.2 | 101.6 | 100.7 | 100.3
5.8 | 102.6 | 102.2 | 101.6 | 100.7 | 100.3
5.9 | 102.6639 | 102.2839 | 101.6 | 100.7 | 100.3
6 | 102.7 | 102.3 | 101.6253 | 100.7 | 100.3
6.1 | 102.7389 | 102.3 | 101.7 | 100.7 | 100.3
6.2 | 102.8 | 102.3 | 101.7 | 100.7 | 100.2324
6.3 | 102.8246 | 102.4 | 101.7 | 100.7 | 100.2
6.4 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
6.5 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
6.6 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
6.7 | 102.9 | 102.4 | 101.7 | 100.7 | 100.2
6.8 | 102.9 | 102.4 | 101.7 | 100.7 | 100.1028
6.9 | 102.9581 | 102.4 | 101.7 | 100.7 | 100.0849
7 | 103 | 102.4 | 101.7 | 100.7 | 100
7.1 | 103 | 102.4809 | 101.7721 | 100.6851 | 100
7.2 | 103 | 102.5186 | 101.8 | 100.6 | 100
7.3 | 103.0838 | 102.6 | 101.7259 | 100.6 | 100
7.4 | 103.1 | 102.6 | 101.7 | 100.6 | 99.9968
7.5 | 103.1466 | 102.6 | 101.7 | 100.6 | 99.9084
7.6 | 103.2 | 102.635 | 101.7 | 100.6 | 99.9
7.7 | 103.2 | 102.7 | 101.7294 | 100.5906 | 99.9
7.8 | 103.2709 | 102.7 | 101.8 | 100.5 | 99.9
7.9 | 103.3 | 102.7 | 101.8 | 100.5 | 99.9
8 | 103.3 | 102.7 | 101.8 | 100.5 | 99.8539
8.1 | 103.3939 | 102.7 | 101.8 | 100.5 | 99.7694
8.2 | 103.4 | 102.8 | 101.8 | 100.4167 | 99.7
8.3 | 103.4975 | 102.8 | 101.8 | 100.4 | 99.7
8.4 | 103.5 | 102.8 | 101.8 | 100.4 | 99.7
8.5 | 103.5 | 102.8 | 101.8 | 100.3 | 99.6245
8.6 | 103.5 | 102.8 | 101.8 | 100.3 | 99.6
8.7 | 103.5237 | 102.8031 | 101.8 | 100.3 | 99.5477
8.8 | 103.6 | 102.8749 | 101.8 | 100.3 | 99.5
8.9 | 103.6 | 102.8 | 101.8 | 100.2525 | 99.5
9 | 103.6 | 102.8 | 101.8 | 100.2 | 99.4101
9.1 | 103.6 | 102.8911 | 101.8 | 100.2 | 99.4
9.2 | 103.7 | 102.9 | 101.8 | 100.1558 | 99.394
9.3 | 103.7 | 102.9 | 101.8 | 100.1 | 99.3155
9.4 | 103.7 | 102.9613 | 101.8 | 100.1 | 99.3
9.5 | 103.7 | 103 | 101.8 | 100.1 | 99.2321
9.6 | 103.7956 | 103 | 101.8 | 100.003 | 99.2
9.7 | 103.8 | 103 | 101.8 | 100 | 99.2
9.8 | 103.8 | 103 | 101.8 | 100 | 99.1379
9.9 | 103.8839 | 103 | 101.8 | 100 | 99.0552
10 | 103.9 | 103 | 101.8 | 100 | 99
10.1 | 103.9 | 103 | 101.8 | 99.9 | 98.9852
10.2 | 103.925 | 103.085 | 101.8 | 99.9 | 98.9
10.3 | 104 | 103.1 | 101.8 | 99.9 | 98.9
10.4 | 104.0899 | 103.1 | 101.8 | 99.8935 | 98.8326
10.5 | 104.1 | 103.1 | 101.8 | 99.8 | 98.8
10.6 | 104.1589 | 103.1 | 101.8 | 99.7744 | 98.7495
10.7 | 104.2 | 103.1267 | 101.8 | 99.7 | 98.7
10.8 | 104.2 | 103.2 | 101.8 | 99.6465 | 98.6666
10.9 | 104.2 | 103.2 | 101.8 | 99.6 | 98.5782
11 | 104.2563 | 103.2 | 101.8 | 99.6 | 98.4898
11.1 | 104.3 | 103.2 | 101.8 | 99.5185 | 98.4014
11.2 | 104.3 | 103.2 | 101.8 | 99.5 | 98.4
11.3 | 104.3 | 103.2806 | 101.8 | 99.4246 | 98.3
11.4 | 104.3 | 103.3 | 101.8 | 99.4 | 98.3
11.5 | 104.3543 | 103.3 | 101.8 | 99.4 | 98.2832
11.6 | 104.4 | 103.3 | 101.8 | 99.3844 | 98.1822
11.7 | 104.4 | 103.3886 | 101.8 | 99.3 | 98.0836
11.8 | 104.4 | 103.4 | 101.8 | 99.3 | 98
11.9 | 104.4891 | 103.4 | 101.8 | 99.3 | 98
12 | 104.5 | 103.4 | 101.8 | 99.2095 | 97.9184
12.1 | 104.5933 | 103.4 | 101.8 | 99.2 | 97.8378
12.2 | 104.6 | 103.4 | 101.719 | 99.1618 | 97.8
12.3 | 104.6 | 103.4 | 101.7 | 99.1 | 97.8
12.4 | 104.6 | 103.4 | 101.7 | 99.0523 | 97.7884
12.5 | 104.6 | 103.4981 | 101.7 | 99 | 97.6902
12.6 | 104.6591 | 103.5 | 101.7 | 98.9435 | 97.6116
12.7 | 104.7 | 103.5 | 101.7 | 98.9 | 97.5331
12.8 | 104.7 | 103.5 | 101.7 | 98.8317 | 97.4488
12.9 | 104.7 | 103.5 | 101.7 | 98.8 | 97.3604
13 | 104.7873 | 103.5 | 101.7 | 98.8 | 97.3
13.1 | 104.8262 | 103.5 | 101.7 | 98.7282 | 97.2813
13.2 | 104.9 | 103.5 | 101.6 | 98.7 | 97.1847
13.3 | 104.9 | 103.5242 | 101.6 | 98.6318 | 97.1061
13.4 | 104.9 | 103.6 | 101.6 | 98.6 | 97.0275
13.5 | 104.9 | 103.6 | 101.6 | 98.6 | 97
13.6 | 104.9733 | 103.6 | 101.6 | 98.5066 | 96.9417
13.7 | 105 | 103.6 | 101.5037 | 98.5 | 96.8533
13.8 | 105 | 103.6 | 101.5 | 98.4366 | 96.7688
13.9 | 105.0976 | 103.6 | 101.5 | 98.3 | 96.689
14 | 105.1325 | 103.6 | 101.5 | 98.2865 | 96.6006
14.1 | 105.2 | 103.6604 | 101.5 | 98.2 | 96.5
14.2 | 105.2 | 103.7 | 101.5 | 98.2 | 96.499
14.3 | 105.2632 | 103.7 | 101.5 | 98.1977 | 96.4204
14.4 | 105.3 | 103.7372 | 101.5 | 98.1 | 96.3346
14.5 | 105.3 | 103.8 | 101.5 | 98.0962 | 96.2522
14.6 | 105.3 | 103.8 | 101.4104 | 98.0078 | 96.2
14.7 | 105.3408 | 103.8 | 101.4 | 97.9079 | 96.1944
14.8 | 105.4 | 103.8 | 101.4 | 97.8069 | 96.106
14.9 | 105.4 | 103.8 | 101.4 | 97.7883 | 96.0059
15 | 105.4665 | 103.8 | 101.4 | 97.7 | 95.9167
15.1 | 105.5 | 103.8 | 101.4 | 97.7 | 95.8045
15.2 | 105.5 | 103.8 | 101.4 | 97.66 | 95.7
15.3 | 105.5975 | 103.9 | 101.4 | 97.5521 | 95.6766
15.4 | 105.6 | 103.9 | 101.4 | 97.5 | 95.5865
15.5 | 105.6 | 103.9 | 101.383 | 97.4426 | 95.4855
15.6 | 105.6 | 103.9 | 101.3 | 97.4 | 95.4
15.7 | 105.6 | 103.9 | 101.3 | 97.3307 | 95.3855
15.8 | 105.6 | 103.9 | 101.2824 | 97.3 | 95.2975
15.9 | 105.646 | 103.9 | 101.2 | 97.2675 | 95.2189
16 | 105.7 | 103.9 | 101.2 | 97.2 | 95.1329
16.1 | 105.7 | 103.9 | 101.2 | 97.1222 | 95.0013
16.2 | 105.7 | 103.9 | 101.2 | 97.0212 | 94.8721
16.3 | 105.7117 | 103.9 | 101.1069 | 96.9202 | 94.7927
16.4 | 105.8 | 103.9 | 101.1 | 96.8906 | 94.7043
16.5 | 105.8 | 103.9 | 101.1 | 96.8 | 94.6253
16.6 | 105.8 | 103.9 | 101.1 | 96.8 | 94.5467
16.7 | 105.8 | 103.9 | 101.1 | 96.7447 | 94.4641
16.8 | 105.8 | 103.9 | 101.1 | 96.6437 | 94.3677
16.9 | 105.8168 | 103.9 | 101.0198 | 96.5427 | 94.3
17 | 105.9 | 103.9 | 101 | 96.5 | 94.256
17.1 | 105.9 | 103.9 | 100.9308 | 96.4769 | 94.1606
17.2 | 105.9 | 103.9 | 100.9 | 96.4 | 94.0722
17.3 | 105.9 | 103.9 | 100.8815 | 96.3243 | 93.9815
17.4 | 105.9 | 103.9 | 100.8 | 96.1939 | 93.8772
17.5 | 105.9 | 103.9 | 100.8 | 96.1 | 93.8
17.6 | 105.9 | 103.9 | 100.7641 | 96.0298 | 93.7123
17.7 | 105.9 | 103.9 | 100.7 | 95.9 | 93.5713
17.8 | 105.9 | 103.9 | 100.6765 | 95.9 | 93.5
17.9 | 105.9 | 103.9 | 100.6 | 95.8213 | 93.5
18 | 105.9465 | 103.8868 | 100.6 | 95.7173 | 93.3059
18.1 | 106 | 103.8 | 100.5228 | 95.6163 | 93.2019
18.2 | 106 | 103.8 | 100.5 | 95.5153 | 93.1009
18.3 | 106 | 103.8 | 100.4499 | 95.3857 | 92.9249
18.4 | 106 | 103.8 | 100.4 | 95.3 | 92.8325
18.5 | 106 | 103.8 | 100.3356 | 95.2836 | 92.7481
18.6 | 106 | 103.8 | 100.3 | 95.2 | 92.654
18.7 | 106 | 103.7 | 100.2387 | 95.1673 | 92.5452
18.8 | 106 | 103.7 | 100.2 | 95.0607 | 92.4377
18.9 | 106 | 103.7 | 100.1367 | 94.902 | 92.3507
19 | 106 | 103.7 | 100.0437 | 94.8 | 92.2722
19.1 | 106 | 103.6904 | 100 | 94.7553 | 92.1904
19.2 | 105.9471 | 103.6 | 99.9479 | 94.6622 | 92.0765
19.3 | 105.9 | 103.6 | 99.9 | 94.5612 | 91.9785
19.4 | 105.9 | 103.6 | 99.9 | 94.4602 | 91.8776
19.5 | 105.9 | 103.5 | 99.8306 | 94.3592 | 91.6857
19.6 | 105.9 | 103.5 | 99.7296 | 94.2634 | 91.5678
19.7 | 105.9 | 103.5 | 99.7 | 94.1714 | 91.45
19.8 | 105.9 | 103.4 | 99.6 | 94.0704 | 91.3322
19.9 | 105.9 | 103.4 | 99.5476 | 93.9694 | 91.3
20 | 105.9 | 103.3298 | 99.4298 | 93.8447 | 91.2112
20.1 | 105.8 | 103.3 | 99.3943 | 93.7943 | 91.1102
20.2 | 105.8 | 103.2941 | 99.3 | 93.6529 | 90.9412
20.3 | 105.8 | 103.2 | 99.2262 | 93.4939 | 90.8397
20.4 | 105.8 | 103.13 | 99.2 | 93.4 | 90.7125
20.5 | 105.8 | 103.1 | 99.1062 | 93.3487 | 90.543
20.6 | 105.756 | 103.0727 | 99.0052 | 93.2226 | 90.392
20.7 | 105.7 | 103 | 98.9 | 93.0898 | 90.3036
20.8 | 105.7 | 102.9243 | 98.837 | 92.9869 | 90.1063
20.9 | 105.7 | 102.9 | 98.7191 | 92.8629 | 90.0143
21 | 105.6 | 102.9 | 98.6815 | 92.7215 | 89.9342
21.1 | 105.6 | 102.8167 | 98.5572 | 92.5801 | 89.85
21.2 | 105.5483 | 102.7 | 98.4617 | 92.4489 | 89.7659
21.3 | 105.5 | 102.6477 | 98.3695 | 92.3409 | 89.6716
21.4 | 105.5 | 102.6 | 98.2685 | 92.2298 | 89.4941
21.5 | 105.4287 | 102.5953 | 98.1675 | 92.112 | 89.2953
21.6 | 105.4 | 102.5 | 98.0665 | 91.873 | 89.1614
21.7 | 105.3716 | 102.443 | 97.9654 | 91.8 | 88.9797
21.8 | 105.3 | 102.3102 | 97.9 | 91.7563 | 88.8834
21.9 | 105.2 | 102.2 | 97.8073 | 91.6267 | 88.8049
22 | 105.1728 | 102.1728 | 97.6788 | 91.4121 | 88.6841
22.1 | 105.1 | 102.0323 | 97.4716 | 91.2382 | 88.5074
22.2 | 105.0245 | 101.9844 | 97.3603 | 91.1317 | 88.4247
22.3 | 104.9788 | 101.8525 | 97.2525 | 91.0307 | 88.3192
22.4 | 104.8347 | 101.7347 | 97.1347 | 90.8594 | 88.1309
22.5 | 104.6802 | 101.7 | 97.0001 | 90.6574 | 87.9669
22.6 | 104.5388 | 101.5646 | 96.8823 | 90.5277 | 87.784
22.7 | 104.3955 | 101.4 | 96.7289 | 90.4359 | 87.5623
22.8 | 104.1699 | 101.3359 | 96.5542 | 90.33 | 87.3724
22.9 | 103.9291 | 101.1954 | 96.4454 | 90.1243 | 87.2717
23 | 103.6327 | 101.0776 | 96.3379 | 89.9931 | 87.1146
23.1 | 103.2792 | 100.9033 | 96.1737 | 89.8516 | 86.9396
23.2 | 102.8109 | 100.7252 | 95.9717 | 89.7252 | 86.8216
23.3 | 102.3034 | 100.5146 | 95.8348 | 89.5378 | 86.7305
23.4 | 101.6368 | 100.3186 | 95.7074 | 89.3274 | 86.5457
23.5 | 100.8958 | 100.1217 | 95.5432 | 89.186 | 86.405
23.6 | 100.1231 | 99.8113 | 95.3208 | 89.0446 | 86.3032
23.7 | 99.0416 | 99.5067 | 95.129 | 88.8386 | 86.086
23.8 | 98.0127 | 99.1453 | 94.9739 | 88.6522 | 85.9681
23.9 | 96.669 | 98.7721 | 94.7719 | 88.517 | 85.8502
24 | 95.389 | 98.3556 | 94.5849 | 88.4134 | 85.6376
24.1 | 94.0745 | 97.7725 | 94.4518 | 88.3124 | 85.4248
24.2 | 92.5721 | 97.1802 | 94.1701 | 88.0342 | 85.2622
24.3 | 91.1311 | 96.422 | 93.991 | 87.791 | 85.0577
24.4 | 89.8585 | 95.5104 | 93.7463 | 87.5732 | 84.8707
24.5 | 88.5175 | 94.612 | 93.4051 | 87.4025 | 84.7647
24.6 | 87.4038 | 93.5774 | 93.0258 | 87.2419 | 84.588
24.7 | 85.9894 | 92.4405 | 92.8378 | 87.0483 | 84.4111
24.8 | 84.4825 | 91.3059 | 92.4823 | 86.8392 | 84.1896
24.9 | 82.3781 | 89.9284 | 92.1152 | 86.6372 | 84.0576
25 | 79.1409 | 88.6461 | 91.6542 | 86.4782 | 83.8782
25.1 | 74.4123 | 87.4458 | 91.2108 | 86.276 | 83.672
25.2 | 68.4072 | 86.1181 | 90.6648 | 86.0609 | 83.4682
25.3 | 61.1844 | 84.7706 | 90.0021 | 85.7754 | 83.3504
25.4 | 53.3294 | 83.558 | 89.379 | 85.5993 | 83.1842
25.5 | 45.8389 | 82.2976 | 88.688 | 85.496 | 82.9969
25.6 | 37.7191 | 81.0144 | 87.8745 | 85.2403 | 82.7935
25.7 | 31.3458 | 79.3912 | 86.9192 | 84.9095 | 82.5719
25.8 | 25.673 | 77.1737 | 85.9282 | 84.6554 | 82.4148
25.9 | 21.5851 | 74.0535 | 84.989 | 84.4839 | 82.2398
26 | 18.092 | 69.914 | 83.9285 | 84.2587 | 82.0782
26.1 | 15.6766 | 64.3492 | 82.8422 | 84.0479 | 81.9113
26.2 | 13.8989 | 57.5875 | 81.6969 | 83.7448 | 81.7127
26.3 | 12.4077 | 51.0954 | 80.4964 | 83.5078 | 81.477
26.4 | 11.4887 | 43.5693 | 79.3522 | 83.2296 | 81.2413
26.5 | 10.7167 | 37.4489 | 78.3094 | 82.9085 | 80.9861
26.6 | 10.1082 | 31.4489 | 77.299 | 82.6042 | 80.8185
26.7 | 9.7414 | 27.216 | 76.0879 | 82.3385 | 80.5718
26.8 | 9.3781 | 23.3139 | 74.9909 | 82.0477 | 80.4318
26.9 | 9.097 | 20.4572 | 73.6446 | 81.7295 | 80.2481
27 | 8.8724 | 17.9765 | 71.9097 | 81.4406 | 80.0604
27.1 | 8.6247 | 16.2521 | 69.4426 | 81.0446 | 79.8028
27.2 | 8.4867 | 14.9675 | 66.1019 | 80.681 | 79.5557
27.3 | 8.2399 | 13.9866 | 61.6172 | 80.2799 | 79.3399
27.4 | 8.0921 | 13.1529 | 56.4556 | 79.9265 | 79.151
27.5 | 7.9692 | 12.5165 | 51.0237 | 79.3611 | 78.9153
27.6 | 7.7795 | 12.0255 | 45.6873 | 78.8561 | 78.652
27.7 | 7.6219 | 11.5588 | 40.0739 | 78.3412 | 78.4104
27.8 | 7.4849 | 11.2081 | 35.0074 | 77.7352 | 78.1811
27.9 | 7.3529 | 10.9086 | 30.5044 | 77.1172 | 78.0042
28 | 7.2221 | 10.604 | 26.8515 | 76.4098 | 77.7042
28.1 | 7.0009 | 10.3812 | 23.9129 | 75.6464 | 77.4343
28.2 | 6.8792 | 10.132 | 21.6022 | 74.7665 | 77.2566
28.3 | 6.8 | 9.8977 | 19.8164 | 73.9584 | 77.0443
28.4 | 6.6963 | 9.7126 | 18.3747 | 73.1452 | 76.7605
28.5 | 6.5624 | 9.4297 | 17.4021 | 72.3283 | 76.4748
28.6 | 6.4335 | 9.2946 | 16.3677 | 71.5676 | 76.2479
28.7 | 6.2601 | 9.1443 | 15.6945 | 70.7576 | 75.9301
28.8 | 6.1307 | 8.8844 | 15.0945 | 69.818 | 75.6922
28.9 | 6 | 8.7244 | 14.5475 | 68.9736 | 75.523
29 | 5.9399 | 8.5463 | 14.1647 | 68.1018 | 75.1695
29.1 | 5.822 | 8.3553 | 13.8212 | 67.2555 | 74.8421
29.2 | 5.7 | 8.2375 | 13.4124 | 66.333 | 74.5124
29.3 | 5.6294 | 8.0071 | 13.1191 | 65.3902 | 74.206
29.4 | 5.5018 | 7.8621 | 12.8702 | 64.3656 | 73.9505
29.5 | 5.4607 | 7.7207 | 12.5021 | 63.0065 | 73.5574
29.6 | 5.3327 | 7.5792 | 12.2988 | 61.0869 | 73.1549
29.7 | 5.2 | 7.4379 | 12.0524 | 58.4469 | 72.7606
29.8 | 5.1564 | 7.3137 | 11.7663 | 54.9492 | 72.3387
29.9 | 5.015 | 7.1687 | 11.5687 | 51.4406 | 71.8832
30 | 4.967 | 7.028 | 11.3864 | 47.493 | 71.4009
30.1 | 4.9 | 6.8935 | 11.0922 | 43.2771 | 70.8708
30.2 | 4.7923 | 6.7756 | 10.9179 | 39.1491 | 70.3403
30.3 | 4.7 | 6.6578 | 10.7117 | 35.1567 | 69.8098
30.4 | 4.6279 | 6.5279 | 10.5257 | 31.7567 | 69.2096
30.5 | 4.5774 | 6.4054 | 10.3108 | 28.8289 | 68.5909
30.6 | 4.485 | 6.3542 | 10.0929 | 26.5763 | 67.9669
30.7 | 4.4 | 6.2364 | 9.8909 | 24.6005 | 67.3182
30.8 | 4.3023 | 6.1185 | 9.6704 | 23.0222 | 66.6109
30.9 | 4.2608 | 5.9808 | 9.5292 | 21.7782 | 65.9039
31 | 4.1995 | 5.9 | 9.3562 | 20.6522 | 65.1212
31.1 | 4.1 | 5.815 | 9.1541 | 19.7155 | 64.3898
31.2 | 4.0165 | 5.6972 | 9.0261 | 18.9854 | 63.7348
31.3 | 3.9751 | 5.6 | 8.8252 | 18.3823 | 63.1115
31.4 | 3.9 | 5.5337 | 8.7098 | 17.8841 | 62.452
31.5 | 3.8523 | 5.3935 | 8.5176 | 17.5307 | 61.6698
31.6 | 3.8 | 5.3 | 8.4078 | 17.0389 | 61.1199
31.7 | 3.7 | 5.2494 | 8.2848 | 16.7203 | 60.3857
31.8 | 3.7 | 5.2 | 8.0914 | 16.4173 | 59.55
31.9 | 3.6 | 5.0888 | 7.9916 | 16.1142 | 58.5546
32 | 3.5065 | 5 | 7.9033 | 15.8112 | 57.3143
32.1 | 3.5 | 4.9364 | 7.7053 | 15.5446 | 55.546
32.2 | 3.4223 | 4.8023 | 7.6016 | 15.3319 | 53.5107
32.3 | 3.4 | 4.8 | 7.5006 | 15.0448 | 51.0277
32.4 | 3.3595 | 4.7163 | 7.3996 | 14.7991 | 48.062
32.5 | 3.3 | 4.6 | 7.3113 | 14.5687 | 44.7434
32.6 | 3.2708 | 4.5767 | 7.1966 | 14.35 | 41.5194
32.7 | 3.2 | 4.5 | 7.068 | 14.0646 | 38.1625
32.8 | 3.1938 | 4.4115 | 6.9615 | 13.8626 | 34.9859
32.9 | 3.1 | 4.377 | 6.8324 | 13.6606 | 32.3936
33 | 3.0925 | 4.3 | 6.7221 | 13.4586 | 29.8912
33.1 | 3 | 4.2296 | 6.6079 | 13.2392 | 27.894
33.2 | 3 | 4.2 | 6.4901 | 13.0401 | 26.2522
33.3 | 2.9066 | 4.1 | 6.4 | 12.8667 | 24.7836
33.4 | 2.9 | 4.0377 | 6.3609 | 12.6422 | 23.5458
33.5 | 2.8532 | 4 | 6.2599 | 12.4678 | 22.6679
33.6 | 2.8 | 3.902 | 6.152 | 12.2424 | 21.8311
33.7 | 2.8 | 3.9 | 6.021 | 12.101 | 21.105
33.8 | 2.8 | 3.833 | 5.9996 | 11.883 | 20.5273
33.9 | 2.7151 | 3.7985 | 5.8864 | 11.7402 | 20.0455
34 | 2.7 | 3.7 | 5.7977 | 11.5567 | 19.6898
34.1 | 2.6553 | 3.7 | 5.6961 | 11.379 | 19.2477
34.2 | 2.6 | 3.6 | 5.6 | 11.2385 | 18.9197
34.3 | 2.6 | 3.6 | 5.5604 | 11.0249 | 18.6541
34.4 | 2.6 | 3.5 | 5.4569 | 10.8926 | 18.3776
34.5 | 2.5297 | 3.5 | 5.4 | 10.7747 | 18.0348
34.6 | 2.5 | 3.4082 | 5.3482 | 10.6261 | 17.747
34.7 | 2.5 | 3.4 | 5.3 | 10.4668 | 17.5113
34.8 | 2.5 | 3.4 | 5.2038 | 10.3423 | 17.2756
34.9 | 2.4 | 3.3 | 5.1028 | 10.1533 | 17.0799
35 | 2.4 | 3.3 | 5.0018 | 10.0516 | 16.8893
35.1 | 2.4 | 3.2011 | 5 | 9.958 | 16.6521
35.2 | 2.3997 | 3.2 | 4.9 | 9.8569 | 16.4247
35.3 | 2.3 | 3.2 | 4.8864 | 9.6766 | 16.2404
35.4 | 2.3 | 3.1168 | 4.7978 | 9.6 | 16.0384
35.5 | 2.3 | 3.1 | 4.7 | 9.4935 | 15.8257
35.6 | 2.3 | 3.1 | 4.7 | 9.2963 | 15.5679
35.7 | 2.2 | 3 | 4.6325 | 9.2079 | 15.321
35.8 |  | 3 | 4.5944 | 9.108 | 15.1852
35.9 |  | 2.9415 | 4.5061 | 8.9139 | 14.9496
36 |  | 2.9 | 4.4 | 8.806 | 14.7631
36.1 |  | 2.9 | 4.4 | 8.7049 | 14.5938
36.2 |  | 2.9 | 4.3455 | 8.5896 | 14.4409
36.3 |  | 2.83 | 4.3 | 8.5 | 14.2915
36.4 |  | 2.8 | 4.2304 | 8.459 | 14.1141
36.5 |  | 2.8 | 4.2 | 8.358 | 13.9446
36.6 |  | 2.7998 | 4.1284 | 8.257 | 13.733
36.7 |  |  | 4.1 | 8.156 | 13.556
36.8 |  |  | 4.0141 | 8.0606 | 13.4606
36.9 |  |  | 4 | 7.9364 | 13.3364
37 |  |  | 4 | 7.8 | 13.1343
37.1 |  |  | 3.9091 | 7.7662 | 12.9662
37.2 |  |  | 3.9 | 7.6652 | 12.839
37.3 |  |  | 3.8213 | 7.5641 | 12.6749
37.4 |  |  | 3.8 | 7.457 | 12.5713
37.5 |  |  | 3.7193 | 7.4 | 12.4814
37.6 |  |  | 3.7 | 7.3325 | 12.2897
37.7 |  |  | 3.6887 | 7.2315 | 12.1868
37.8 |  |  | 3.6 | 7.1305 | 12.0689
37.9 |  |  | 3.6 | 7.0295 | 11.9162
38 |  |  | 3.5874 | 7 | 11.7499
38.1 |  |  |  | 6.9275 | 11.632
38.2 |  |  |  | 6.8265 | 11.5265
38.3 |  |  |  | 6.7255 | 11.3259
38.4 |  |  |  | 6.7 | 11.2101
38.5 |  |  |  | 6.6234 | 11.1091
38.6 |  |  |  | 6.5321 | 11.0081
38.7 |  |  |  | 6.4357 | 10.9071
38.8 |  |  |  | 6.4 | 10.7918
38.9 |  |  |  | 6.3226 | 10.6893
39 |  |  |  | 6.2184 | 10.5755
39.1 |  |  |  | 6.2 | 10.4702
39.2 |  |  |  | 6.1024 | 10.3683
39.3 |  |  |  | 6.1 | 10.2868
39.4 |  |  |  | 6 | 10.1833
39.5 |  |  |  | 5.9704 | 10.0704
39.6 |  |  |  | 5.9 | 9.9694
39.7 |  |  |  | 5.8684 | 9.8631
39.8 |  |  |  | 5.7619 | 9.8
39.9 |  |  |  | 5.7 | 9.7441
40 |  |  |  | 5.6368 | 9.6368
40.1 |  |  |  | 5.6 | 9.525
40.2 |  |  |  | 5.5347 | 9.4381
40.3 |  |  |  | 5.5 | 9.3545
40.4 |  |  |  | 5.5 | 9.2548
40.5 |  |  |  | 5.4402 | 9.137
40.6 |  |  |  | 5.4 | 9.0191
40.7 |  |  |  | 5.3439 | 9
40.8 |  |  |  | 5.3 | 8.8858
40.9 |  |  |  | 5.2155 | 8.7867
41 |  |  |  |  | 8.7
41.1 |  |  |  |  | 8.6799
41.2 |  |  |  |  | 8.562
41.3 |  |  |  |  | 8.5
41.4 |  |  |  |  | 8.4447
41.5 |  |  |  |  | 8.3563
41.6 |  |  |  |  | 8.3
41.7 |  |  |  |  | 8.2227
41.8 |  |  |  |  | 8.2
41.9 |  |  |  |  | 8.1175
42 |  |  |  |  | 8.035
42.1 |  |  |  |  | 7.951
42.2 |  |  |  |  | 7.8626
42.3 |  |  |  |  | 7.8
42.4 |  |  |  |  | 7.7859
42.5 |  |  |  |  | 7.6974
42.6 |  |  |  |  | 7.6091
42.7 |  |  |  |  | 7.6
42.8 |  |  |  |  | 7.5448
42.9 |  |  |  |  | 7.4502
43 |  |  |  |  | 7.4
43.1 |  |  |  |  | 7.3421
43.2 |  |  |  |  | 7.2471
43.3 |  |  |  |  | 7.2
43.4 |  |  |  |  | 7.1684
43.5 |  |  |  |  | 7.1
43.6 |  |  |  |  | 7.1
43.7 |  |  |  |  | 7.0216

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.88 | 0.891 | 0.899 | 0.908 | 0.915 | 0.92 | 0.923 | 0.926 | 0.926
| 4 | 0.892 | 0.906 | 0.915 | 0.927 | 0.936 | 0.943 | 0.946 | 0.95 | 0.951
| 5 | 0.901 | 0.917 | 0.928 | 0.942 | 0.952 | 0.961 | 0.966 | 0.97 | 0.971
| 7 | 0.913 | 0.932 | 0.946 | 0.963 | 0.977 | 0.988 | 0.995 | 1.001 | 1.003
| 10 | 0.924 | 0.946 | 0.962 | 0.983 | 1 | 1.015 | 1.024 | 1.032 | 1.035
| 15 | 0.932 | 0.956 | 0.974 | 0.999 | 1.02 | 1.039 | 1.052 | 1.063 | 1.067
| 20 | 0.936 | 0.961 | 0.98 | 1.007 | 1.031 | 1.053 | 1.067 | 1.082 | 1.086
| 30 | 0.941 | 0.967 | 0.987 | 1.016 | 1.043 | 1.069 | 1.086 | 1.104 | 1.11
| 40 | 0.943 | 0.97 | 0.99 | 1.02 | 1.048 | 1.076 | 1.095 | 1.113 | 1.119

## Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 43.4 | 48.8 | 57.2 | 64.4 | 67
0.1 | 47.5 | 52.975 | 61.2 | 67.9667 | 70.55
0.2 | 54.3333 | 59.2 | 66.64 | 72.825 | 75.56
0.3 | 64.14 | 68.05 | 74.36 | 79.8333 | 81.4
0.4 | 72.84 | 76.325 | 81.3 | 85.68 | 87.0667
0.5 | 80.04 | 82.9667 | 86.9667 | 90.075 | 91.34
0.6 | 85.4 | 87.82 | 90.8 | 93.48 | 94.28
0.7 | 89.8 | 91.6 | 93.9 | 95.82 | 96.38
0.8 | 92.7 | 94 | 95.95 | 97.42 | 97.8
0.9 | 95.26 | 96.02 | 97.52 | 98.5 | 98.8
1 | 96.8 | 97.6 | 98.7 | 99.2333 | 99.5
1.1 | 98.1 | 98.55 | 99.4 | 99.7 | 99.78
1.2 | 99.16 | 99.34 | 99.76 | 99.9 | 99.925
1.3 | 99.6 | 99.7 | 99.98 | 100 | 100
1.4 | 99.88 | 99.95 | 100 | 100 | 100
1.5 | 100 | 100 | 100 | 99.9 | 99.8333
1.6 | 99.975 | 99.98 | 99.9 | 99.64 | 99.625
1.7 | 99.84 | 99.9 | 99.78 | 99.5 | 99.4
1.8 | 99.64 | 99.68 | 99.5 | 99.32 | 99.125
1.9 | 99.4 | 99.3333 | 99.2 | 98.9 | 98.8571
2 | 99.04 | 99.1333 | 98.9 | 98.62 | 98.6
2.1 | 98.6 | 98.775 | 98.5333 | 98.22 | 98.1333
2.2 | 98.04 | 98.32 | 98.16 | 97.88 | 97.8
2.3 | 97.64 | 97.98 | 97.825 | 97.55 | 97.5333
2.4 | 97.16 | 97.5667 | 97.5 | 97.2 | 97.15
2.5 | 96.6333 | 97.2 | 97.1667 | 96.875 | 96.84
2.6 | 96.22 | 96.825 | 96.75 | 96.52 | 96.42
2.7 | 95.65 | 96.3 | 96.35 | 96.125 | 96.1333
2.8 | 95.12 | 95.95 | 95.96 | 95.8 | 95.74
2.9 | 94.725 | 95.4 | 95.65 | 95.4 | 95.425
3 | 94.2333 | 95.05 | 95.2333 | 95.075 | 95.1
3.1 | 93.62 | 94.65 | 94.84 | 94.76 | 94.7
3.2 | 93.15 | 94.18 | 94.48 | 94.375 | 94.4
3.3 | 92.66 | 93.82 | 94.15 | 94.0667 | 94
3.4 | 92.1333 | 93.5 | 93.76 | 93.8 | 93.55
3.5 | 91.7 | 93 | 93.45 | 93.275 | 93.3
3.6 | 91.2 | 92.55 | 92.96 | 93 | 92.98
3.7 | 90.7 | 92.16 | 92.58 | 92.6 | 92.6
3.8 | 90.1857 | 91.7 | 92.15 | 92.3 | 92.2
3.9 | 89.78 | 91.38 | 91.9 | 91.9 | 91.8
4 | 89.24 | 90.9 | 91.48 | 91.46 | 91.475
4.1 | 88.72 | 90.5667 | 91.1 | 91.1333 | 91.18
4.2 | 88.18 | 90.05 | 90.75 | 90.8 | 90.825
4.3 | 87.65 | 89.62 | 90.3 | 90.42 | 90.4
4.4 | 87.225 | 89.3 | 89.9333 | 90.02 | 90.0667
4.5 | 86.78 | 88.86 | 89.5 | 89.65 | 89.7
4.6 | 86.2 | 88.4 | 89.2 | 89.25 | 89.3667
4.7 | 85.75 | 88 | 88.8 | 88.9 | 89.075
4.8 | 85.35 | 87.6 | 88.5 | 88.65 | 88.68
4.9 | 84.76 | 87.1 | 88.2 | 88.3 | 88.32
5 | 84.32 | 86.74 | 87.75 | 87.95 | 87.86
5.1 | 83.8 | 86.3 | 87.34 | 87.55 | 87.6
5.2 | 83.45 | 85.88 | 87 | 87.15 | 87.1667
5.3 | 82.95 | 85.55 | 86.52 | 86.82 | 86.7667
5.4 | 82.4 | 85.1 | 86.12 | 86.48 | 86.46
5.5 | 81.94 | 84.66 | 85.72 | 86.1 | 86.075
5.6 | 81.46 | 84.3 | 85.425 | 85.78 | 85.72
5.7 | 80.975 | 83.9 | 85.05 | 85.44 | 85.38
5.8 | 80.52 | 83.4 | 84.7 | 84.9429 | 85.12
5.9 | 79.96 | 82.95 | 84.4 | 84.72 | 84.72
6 | 79.6 | 82.56 | 83.9 | 84.375 | 84.4
6.1 | 79 | 82.25 | 83.5 | 84.02 | 84.1333
6.2 | 78.6 | 81.82 | 83.18 | 83.625 | 83.7333
6.3 | 78.1 | 81.38 | 82.88 | 83.34 | 83.4
6.4 | 77.8 | 81.04 | 82.6 | 82.96 | 83.04
6.5 | 77.38 | 80.775 | 82.2 | 82.56 | 82.7714
6.6 | 77 | 80.1 | 81.7 | 82.2333 | 82.44
6.7 | 76.35 | 79.75 | 81.45 | 81.875 | 82
6.8 | 75.9 | 79.3429 | 81.1 | 81.55 | 81.6
6.9 | 75.52 | 79.1 | 80.74 | 81.2 | 81.3667
7 | 75.05 | 78.7333 | 80.3333 | 80.88 | 81.04
7.1 | 74.5667 | 78.34 | 79.94 | 80.6 | 80.7
7.2 | 74.2333 | 77.94 | 79.54 | 80.14 | 80.38
7.3 | 73.7 | 77.55 | 79.32 | 79.8333 | 80
7.4 | 73.36 | 77.1 | 78.925 | 79.5 | 79.65
7.5 | 72.9 | 76.7 | 78.5 | 79.175 | 79.325
7.6 | 72.5 | 76.28 | 78.2 | 78.94 | 79.1
7.7 | 71.95 | 75.9667 | 77.8833 | 78.5 | 78.6333
7.8 | 71.525 | 75.54 | 77.56 | 78.15 | 78.25
7.9 | 71.2 | 75.1667 | 77.1333 | 77.82 | 78
8 | 70.75 | 74.78 | 76.825 | 77.5 | 77.72
8.1 | 70.25 | 74.38 | 76.425 | 77.12 | 77.44
8.2 | 69.8 | 74.08 | 76.175 | 76.85 | 77.0333
8.3 | 69.48 | 73.62 | 75.7667 | 76.54 | 76.6667
8.4 | 69.06 | 73.35 | 75.3667 | 76.1 | 76.325
8.5 | 68.6 | 72.85 | 75.1 | 75.775 | 76.0333
8.6 | 68.2 | 72.5 | 74.8 | 75.56 | 75.75
8.7 | 67.85 | 72.15 | 74.4 | 75.2 | 75.3333
8.8 | 67.4333 | 71.82 | 74.125 | 74.85 | 75
8.9 | 67.2 | 71.4833 | 73.7 | 74.5 | 74.76
9 | 66.6429 | 71.1 | 73.375 | 74.2 | 74.4667
9.1 | 66.2 | 70.7 | 73.025 | 73.8667 | 74.1
9.2 | 65.7667 | 70.36 | 72.8 | 73.6143 | 73.7667
9.3 | 65.36 | 69.95 | 72.5 | 73.3 | 73.45
9.4 | 65 | 69.55 | 72 | 72.975 | 73.1
9.5 | 64.6 | 69.2 | 71.7 | 72.6 | 72.8
9.6 | 64.25 | 68.9 | 71.3 | 72.225 | 72.4
9.7 | 63.78 | 68.55 | 70.9667 | 71.88 | 72.15
9.8 | 63.46 | 68.15 | 70.6333 | 71.5833 | 71.76
9.9 | 63.0667 | 67.6667 | 70.3 | 71.275 | 71.48
10 | 62.66 | 67.4167 | 69.95 | 70.9 | 71.22
10.1 | 62.3 | 67.04 | 69.6 | 70.5333 | 70.8667
10.2 | 61.86 | 66.75 | 69.35 | 70.2333 | 70.55
10.3 | 61.5429 | 66.3 | 69.025 | 69.975 | 70.22
10.4 | 61.25 | 66.02 | 68.8 | 69.7 | 69.9
10.5 | 60.86 | 65.6333 | 68.375 | 69.3 | 69.66
10.6 | 60.38 | 65.1833 | 68.04 | 68.975 | 69.375
10.7 | 60.05 | 64.9333 | 67.6667 | 68.6833 | 69.05
10.8 | 59.75 | 64.7 | 67.4 | 68.475 | 68.72
10.9 | 59.4 | 64.25 | 67.15 | 68.15 | 68.4
11 | 59.1 | 63.8667 | 66.8 | 67.9167 | 68.1
11.1 | 58.7 | 63.625 | 66.56 | 67.58 | 67.9
11.2 | 58.3 | 63.24 | 66.16 | 67.3 | 67.625
11.3 | 58.02 | 62.9 | 65.85 | 67.0333 | 67.3
11.4 | 57.54 | 62.7 | 65.6 | 66.7 | 66.88
11.5 | 57.24 | 62.32 | 65.18 | 66.38 | 66.6333
11.6 | 56.94 | 61.96 | 64.96 | 66.075 | 66.35
11.7 | 56.68 | 61.5667 | 64.65 | 65.825 | 66.0667
11.8 | 56.3 | 61.2333 | 64.4 | 65.55 | 65.78
11.9 | 55.95 | 60.85 | 64.06 | 65.1667 | 65.46
12 | 55.66 | 60.64 | 63.725 | 64.875 | 65.2
12.1 | 55.28 | 60.28 | 63.46 | 64.64 | 64.88
12.2 | 54.9 | 59.98 | 63.175 | 64.25 | 64.575
12.3 | 54.7 | 59.7 | 62.84 | 63.9667 | 64.2
12.4 | 54.34 | 59.325 | 62.5667 | 63.7167 | 63.925
12.5 | 54.06 | 59.04 | 62.25 | 63.42 | 63.74
12.6 | 53.65 | 58.6333 | 61.9 | 63.12 | 63.4667
12.7 | 53.2333 | 58.3333 | 61.6 | 62.9 | 63.225
12.8 | 53 | 58.12 | 61.42 | 62.65 | 62.9
12.9 | 52.65 | 57.76 | 60.95 | 62.2667 | 62.6
13 | 52.46 | 57.4 | 60.725 | 61.95 | 62.32
13.1 | 52.05 | 57.2 | 60.4 | 61.7 | 62.08
13.2 | 51.775 | 56.875 | 60.16 | 61.4333 | 61.825
13.3 | 51.525 | 56.525 | 59.925 | 61.175 | 61.5
13.4 | 51.14 | 56.175 | 59.675 | 60.9 | 61.3
13.5 | 50.82 | 55.925 | 59.4 | 60.5 | 60.775
13.6 | 50.5 | 55.7 | 58.9333 | 60.26 | 60.6
13.7 | 50.3 | 55.275 | 58.75 | 60.0667 | 60.3667
13.8 | 49.975 | 54.95 | 58.475 | 59.775 | 60.1
13.9 | 49.6 | 54.7167 | 58.18 | 59.44 | 59.84
14 | 49.35 | 54.5 | 57.9 | 59.14 | 59.62
14.1 | 49.025 | 54.08 | 57.58 | 58.95 | 59.38
14.2 | 48.72 | 53.75 | 57.375 | 58.725 | 59.0667
14.3 | 48.4333 | 53.45 | 57.06 | 58.44 | 58.8
14.4 | 48.14 | 53.25 | 56.75 | 58.125 | 58.4667
14.5 | 47.85 | 53.0143 | 56.4667 | 57.8 | 58.25
14.6 | 47.5667 | 52.68 | 56.15 | 57.6167 | 57.925
14.7 | 47.3 | 52.5 | 55.86 | 57.4 | 57.6833
14.8 | 47.0667 | 52.14 | 55.6333 | 57.1 | 57.4333
14.9 | 46.7333 | 51.8 | 55.4 | 56.8 | 57.14
15 | 46.4 | 51.575 | 55.14 | 56.5 | 57
15.1 | 46.12 | 51.35 | 54.8333 | 56.3167 | 56.64
15.2 | 45.875 | 51 | 54.4833 | 56 | 56.4
15.3 | 45.55 | 50.675 | 54.2333 | 55.7667 | 56.1
15.4 | 45.25 | 50.48 | 53.98 | 55.54 | 55.8667
15.5 | 45.025 | 50.18 | 53.8 | 55.2 | 55.6
15.6 | 44.8 | 49.86 | 53.575 | 54.9667 | 55.3
15.7 | 44.5667 | 49.55 | 53.3 | 54.7667 | 55.1
15.8 | 44.2667 | 49.325 | 52.9667 | 54.5 | 54.825
15.9 | 44 | 49.05 | 52.72 | 54.26 | 54.6667
16 | 43.725 | 48.82 | 52.5 | 53.98 | 54.375
16.1 | 43.5167 | 48.55 | 52.22 | 53.76 | 54.14
16.2 | 43.25 | 48.36 | 51.975 | 53.55 | 53.9333
16.3 | 43.04 | 48.1 | 51.74 | 53.24 | 53.65
16.4 | 42.7 | 47.8 | 51.4571 | 53 | 53.46
16.5 | 42.475 | 47.5333 | 51.26 | 52.8 | 53.15
16.6 | 42.16 | 47.32 | 51 | 52.5333 | 52.94
16.7 | 41.9 | 47.04 | 50.7 | 52.26 | 52.62
16.8 | 41.6833 | 46.8 | 50.525 | 52.05 | 52.42
16.9 | 41.5 | 46.6 | 50.26 | 51.8 | 52.175
17 | 41.2167 | 46.2 | 49.9714 | 51.5667 | 51.94
17.1 | 40.9667 | 46.0167 | 49.7333 | 51.3286 | 51.7
17.2 | 40.7 | 45.75 | 49.525 | 51.04 | 51.5
17.3 | 40.45 | 45.4833 | 49.26 | 50.875 | 51.1833
17.4 | 40.26 | 45.2333 | 48.9833 | 50.6 | 51.0167
17.5 | 40.08 | 44.94 | 48.78 | 50.3167 | 50.8
17.6 | 39.825 | 44.7 | 48.6 | 50.1167 | 50.625
17.7 | 39.5667 | 44.5667 | 48.32 | 49.94 | 50.3
17.8 | 39.3167 | 44.3 | 48.08 | 49.68 | 50.1
17.9 | 39.08 | 44.04 | 47.8 | 49.42 | 49.8333
18 | 38.92 | 43.78 | 47.6 | 49.1 | 49.62
18.1 | 38.58 | 43.55 | 47.4 | 48.975 | 49.38
18.2 | 38.34 | 43.42 | 47.14 | 48.84 | 49.15
18.3 | 38.15 | 43.1333 | 46.94 | 48.6 | 49.02
18.4 | 37.85 | 42.9 | 46.6667 | 48.3167 | 48.75
18.5 | 37.72 | 42.7 | 46.4333 | 48.05 | 48.44
18.6 | 37.5 | 42.46 | 46.2 | 47.8 | 48.25
18.7 | 37.325 | 42.125 | 46.02 | 47.675 | 48.1
18.8 | 37.0667 | 41.9333 | 45.7333 | 47.425 | 47.8
18.9 | 36.8 | 41.7 | 45.5667 | 47.175 | 47.62
19 | 36.6 | 41.45 | 45.3167 | 47 | 47.34
19.1 | 36.4167 | 41.2 | 45.0667 | 46.82 | 47.14
19.2 | 36.1833 | 41.075 | 44.84 | 46.5333 | 46.94
19.3 | 35.9333 | 40.8571 | 44.64 | 46.3 | 46.68
19.4 | 35.8 | 40.6 | 44.4 | 46.08 | 46.45
19.5 | 35.55 | 40.3 | 44.18 | 45.9 | 46.36
19.6 | 35.34 | 40.2 | 43.9667 | 45.68 | 46.0667
19.7 | 35.1667 | 39.95 | 43.7286 | 45.38 | 45.8667
19.8 | 35 | 39.75 | 43.5143 | 45.2 | 45.65
19.9 | 34.7833 | 39.425 | 43.2833 | 45.02 | 45.5
20 | 34.6167 | 39.2571 | 43.0333 | 44.825 | 45.2
20.1 | 34.35 | 38.975 | 42.88 | 44.5667 | 44.975
20.2 | 34.14 | 38.85 | 42.64 | 44.4 | 44.75
20.3 | 33.925 | 38.7 | 42.45 | 44.2 | 44.575
20.4 | 33.7 | 38.4333 | 42.25 | 43.9 | 44.35
20.5 | 33.52 | 38.22 | 42.1 | 43.7429 | 44.125
20.6 | 33.375 | 38 | 41.84 | 43.5667 | 44
20.7 | 33.1 | 37.7667 | 41.6 | 43.3167 | 43.825
20.8 | 32.9 | 37.6 | 41.3667 | 43.18 | 43.6
20.9 | 32.8 | 37.4 | 41.125 | 43 | 43.325
21 | 32.575 | 37.2 | 40.95 | 42.7 | 43.2
21.1 | 32.44 | 37.0667 | 40.7857 | 42.56 | 43.02
21.2 | 32.14 | 36.825 | 40.5833 | 42.275 | 42.7667
21.3 | 31.98 | 36.6 | 40.375 | 42.125 | 42.5
21.4 | 31.8 | 36.48 | 40.2 | 42 | 42.3833
21.5 | 31.6667 | 36.2 | 40.08 | 41.775 | 42.2
21.6 | 31.46 | 36 | 39.82 | 41.55 | 42
21.7 | 31.36 | 35.8 | 39.55 | 41.34 | 41.875
21.8 | 31.1 | 35.58 | 39.4 | 41.15 | 41.6
21.9 | 30.9 | 35.4 | 39.2 | 41.025 | 41.45
22 | 30.7 | 35.225 | 39 | 40.8 | 41.22
22.1 | 30.525 | 35.15 | 38.8 | 40.6167 | 41.1
22.2 | 30.4333 | 34.95 | 38.6167 | 40.4 | 40.86
22.3 | 30.2333 | 34.7333 | 38.4 | 40.18 | 40.6667
22.4 | 30.02 | 34.46 | 38.3 | 40.05 | 40.4167
22.5 | 29.92 | 34.2667 | 38.0333 | 39.8571 | 40.225
22.6 | 29.775 | 34.15 | 37.85 | 39.68 | 40.1
22.7 | 29.54 | 33.9 | 37.6 | 39.475 | 39.9
22.8 | 29.35 | 33.7 | 37.4667 | 39.24 | 39.7
22.9 | 29.1667 | 33.5833 | 37.3 | 39.05 | 39.52
23 | 28.9833 | 33.38 | 37.15 | 38.86 | 39.34
23.1 | 28.775 | 33.1667 | 36.95 | 38.7 | 39.12
23.2 | 28.64 | 33.1 | 36.74 | 38.5 | 38.9
23.3 | 28.5 | 32.875 | 36.5667 | 38.3167 | 38.7
23.4 | 28.3 | 32.65 | 36.38 | 38.2 | 38.6286
23.5 | 28.2333 | 32.5 | 36.2 | 37.95 | 38.45
23.6 | 28.04 | 32.32 | 36 | 37.8333 | 38.24
23.7 | 27.8333 | 32.12 | 35.8 | 37.65 | 38.125
23.8 | 27.7 | 32 | 35.675 | 37.45 | 37.96
23.9 | 27.54 | 31.8 | 35.4833 | 37.275 | 37.725
24 | 27.4 | 31.6429 | 35.3 | 37.06 | 37.6
24.1 | 27.2833 | 31.5 | 35.1857 | 36.9 | 37.36
24.2 | 27.1 | 31.325 | 35.02 | 36.775 | 37.15
24.3 | 26.98 | 31.1 | 34.8 | 36.6 | 36.96
24.4 | 26.8 | 30.96 | 34.7 | 36.38 | 36.825
24.5 | 26.6 | 30.8 | 34.4667 | 36.2 | 36.68
24.6 | 26.5 | 30.6 | 34.3 | 36.1 | 36.475
24.7 | 26.34 | 30.4667 | 34.175 | 35.85 | 36.325
24.8 | 26.2 | 30.3 | 33.95 | 35.7167 | 36.175
24.9 | 25.96 | 30.1 | 33.75 | 35.54 | 36
25 | 25.85 | 29.95 | 33.64 | 35.35 | 35.8
25.1 | 25.7 | 29.725 | 33.5 | 35.2 | 35.68
25.2 | 25.5 | 29.6 | 33.3 | 35 | 35.46
25.3 | 25.4 | 29.5 | 33.1 | 34.88 | 35.28
25.4 | 25.3 | 29.3 | 32.98 | 34.66 | 35.1
25.5 | 25.1 | 29.2 | 32.76 | 34.525 | 35
25.6 | 25 | 29 | 32.625 | 34.4 | 34.8
25.7 | 24.85 | 28.82 | 32.475 | 34.24 | 34.65
25.8 | 24.75 | 28.7 | 32.35 | 34 | 34.52
25.9 | 24.525 | 28.5833 | 32.16 | 33.9 | 34.35
26 | 24.4 | 28.4 | 31.9833 | 33.7 | 34.2
26.1 | 24.25 | 28.28 | 31.8 | 33.58 | 34.02
26.2 | 24.15 | 28.1 | 31.7 | 33.4 | 33.8167
26.3 | 24.0429 | 27.9333 | 31.52 | 33.2 | 33.7
26.4 | 23.9 | 27.8 | 31.3 | 33.0833 | 33.4857
26.5 | 23.7667 | 27.65 | 31.2 | 32.9667 | 33.32
26.6 | 23.6 | 27.5 | 31.0167 | 32.7 | 33.2
26.7 | 23.5 | 27.36 | 30.9 | 32.6 | 33.0333
26.8 | 23.35 | 27.2 | 30.7167 | 32.5 | 32.9
26.9 | 23.2 | 27.0667 | 30.6 | 32.3 | 32.76
27 | 23.0667 | 26.9 | 30.46 | 32.1333 | 32.525
27.1 | 22.9 | 26.7 | 30.24 | 31.9667 | 32.5
27.2 | 22.78 | 26.6 | 30.1 | 31.85 | 32.34
27.3 | 22.7 | 26.42 | 29.9333 | 31.725 | 32.22
27.4 | 22.54 | 26.4 | 29.85 | 31.52 | 32
27.5 | 22.34 | 26.1833 | 29.7833 | 31.3833 | 31.8167
27.6 | 22.24 | 26.05 | 29.6167 | 31.2 | 31.7
27.7 | 22.1667 | 25.825 | 29.35 | 31.12 | 31.6
27.8 | 22.1 | 25.8 | 29.2 | 30.9333 | 31.4
27.9 | 21.9 | 25.6 | 29.1 | 30.8 | 31.28
28 | 21.8 | 25.5 | 29 | 30.7 | 31.1
28.1 | 21.6833 | 25.38 | 28.8 | 30.5667 | 31
28.2 | 21.5167 | 25.2 | 28.6833 | 30.4 | 30.88
28.3 | 21.4 | 25.1 | 28.5167 | 30.275 | 30.68
28.4 | 21.3 | 24.975 | 28.4 | 30.0667 | 30.5
28.5 | 21.2 | 24.825 | 28.24 | 29.8833 | 30.35
28.6 | 21.0667 | 24.725 | 28.12 | 29.8 | 30.25
28.7 | 20.9667 | 24.6 | 28 | 29.68 | 30.1
28.8 | 20.875 | 24.4 | 27.88 | 29.58 | 29.98
28.9 | 20.75 | 24.3333 | 27.68 | 29.38 | 29.875
29 | 20.6 | 24.125 | 27.575 | 29.2667 | 29.66
29.1 | 20.5 | 24.0429 | 27.4 | 29.1 | 29.5333
29.2 | 20.4 | 23.9 | 27.3 | 28.925 | 29.4833
29.3 | 20.275 | 23.8 | 27.18 | 28.8167 | 29.3167
29.4 | 20.125 | 23.6 | 27.08 | 28.64 | 29.2
29.5 | 20.0333 | 23.5 | 26.875 | 28.55 | 29
29.6 | 19.9 | 23.4 | 26.76 | 28.4 | 28.86
29.7 | 19.8 | 23.3 | 26.6 | 28.3 | 28.76
29.8 | 19.7 | 23.125 | 26.5 | 28.175 | 28.6
29.9 | 19.52 | 23 | 26.4 | 28.04 | 28.45
30 | 19.4167 | 22.9 | 26.3 | 27.95 | 28.3

## SSD80 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 45.3 | 51.9 | 63.8 | 73.3 | 76.7
0.1 | 50.35 | 57 | 68.2 | 77.65 | 80.5
0.2 | 58.8 | 64.55 | 75.05 | 82.6714 | 85.3
0.3 | 69.4 | 73.975 | 82.8 | 88.3778 | 90.65
0.4 | 77.85 | 81.3667 | 88.1 | 92.45 | 94.15
0.5 | 84.125 | 87.125 | 92.025 | 95.45 | 96.5
0.6 | 88.45 | 90.9333 | 94.8333 | 97.4 | 98.025
0.7 | 92 | 93.85 | 96.825 | 98.7 | 99
0.8 | 94.75 | 96.16 | 98.2 | 99.475 | 99.6429
0.9 | 96.7 | 97.75 | 99.125 | 99.875 | 99.9
1 | 98.15 | 98.8 | 99.75 | 100 | 100
1.1 | 99.1 | 99.55 | 99.975 | 100 | 100
1.2 | 99.7 | 99.9 | 100 | 100 | 99.9333
1.3 | 99.9 | 100 | 100 | 99.8625 | 99.7
1.4 | 100 | 100 | 99.94 | 99.575 | 99.45
1.5 | 100 | 99.94 | 99.7 | 99.28 | 99.0625
1.6 | 99.9 | 99.7 | 99.34 | 98.925 | 98.7571
1.7 | 99.6 | 99.5 | 99.1 | 98.65 | 98.34
1.8 | 99.275 | 99.2 | 98.6 | 98.15 | 98.0125
1.9 | 98.85 | 98.7 | 98.2 | 97.825 | 97.6375
2 | 98.375 | 98.4 | 97.85 | 97.275 | 97.25
2.1 | 97.95 | 97.8857 | 97.5 | 96.9667 | 96.7875
2.2 | 97.325 | 97.5 | 97.02 | 96.55 | 96.4
2.3 | 96.75 | 97.05 | 96.5 | 96.175 | 95.95
2.4 | 96.3 | 96.475 | 96.1333 | 95.8 | 95.6
2.5 | 95.7667 | 95.95 | 95.65 | 95.3 | 95.1
2.6 | 95.2 | 95.575 | 95.175 | 94.8 | 94.7333
2.7 | 94.7 | 95.08 | 94.75 | 94.45 | 94.3
2.8 | 94.1 | 94.65 | 94.35 | 94.05 | 93.92
2.9 | 93.55 | 94.1 | 93.95 | 93.7 | 93.55
3 | 93 | 93.7 | 93.5 | 93.15 | 93
3.1 | 92.4 | 93.1 | 93.05 | 92.75 | 92.65
3.2 | 91.8 | 92.7 | 92.6 | 92.3 | 92.2
3.3 | 91.35 | 92.2 | 92.15 | 91.95 | 91.82
3.4 | 90.9 | 91.7 | 91.65 | 91.5 | 91.4
3.5 | 90.2333 | 91.25 | 91.25 | 91.05 | 91.05
3.6 | 89.66 | 90.85 | 90.8667 | 90.7 | 90.55
3.7 | 89.15 | 90.3333 | 90.4 | 90.275 | 90.1625
3.8 | 88.575 | 89.875 | 89.95 | 89.85 | 89.75
3.9 | 88.0333 | 89.3333 | 89.5 | 89.35 | 89.3125
4 | 87.45 | 88.8 | 89.1333 | 88.9875 | 88.85
4.1 | 86.85 | 88.4667 | 88.65 | 88.55 | 88.46
4.2 | 86.375 | 87.925 | 88.3 | 88.18 | 88.1
4.3 | 85.92 | 87.4667 | 87.8 | 87.7625 | 87.7
4.4 | 85.25 | 86.95 | 87.35 | 87.35 | 87.3
4.5 | 84.72 | 86.65 | 87 | 86.9286 | 86.8
4.6 | 84.32 | 86.1 | 86.58 | 86.5 | 86.4667
4.7 | 83.7286 | 85.6 | 86.1 | 86.15 | 86.0667
4.8 | 83.1 | 85.0571 | 85.6 | 85.7714 | 85.7
4.9 | 82.65 | 84.7 | 85.2 | 85.275 | 85.2
5 | 82.1 | 84.2 | 84.76 | 84.95 | 84.85
5.1 | 81.55 | 83.75 | 84.35 | 84.5 | 84.5
5.2 | 81.06 | 83.2333 | 84 | 84.1 | 84.15
5.3 | 80.525 | 82.8 | 83.55 | 83.7 | 83.7
5.4 | 80.125 | 82.275 | 83.1333 | 83.2 | 83.2667
5.5 | 79.45 | 81.95 | 82.7667 | 82.8 | 82.9
5.6 | 78.9 | 81.45 | 82.35 | 82.4 | 82.5
5.7 | 78.5 | 80.94 | 81.85 | 82.05 | 82.1
5.8 | 77.95 | 80.6 | 81.45 | 81.675 | 81.7
5.9 | 77.5 | 80.1 | 81.1 | 81.2875 | 81.325
6 | 77 | 79.6 | 80.65 | 80.85 | 81
6.1 | 76.4 | 79.05 | 80.2 | 80.525 | 80.55
6.2 | 75.9 | 78.725 | 79.8 | 80.1556 | 80.18
6.3 | 75.5 | 78.45 | 79.4 | 79.7333 | 79.9
6.4 | 75.05 | 77.92 | 78.95 | 79.4 | 79.36
6.5 | 74.55 | 77.55 | 78.625 | 78.9 | 78.95
6.6 | 74.05 | 77.05 | 78.175 | 78.62 | 78.6
6.7 | 73.55 | 76.55 | 77.85 | 78.225 | 78.3
6.8 | 72.975 | 76.15 | 77.45 | 77.85 | 77.8
6.9 | 72.6 | 75.8 | 77.04 | 77.35 | 77.45
7 | 72.1 | 75.3 | 76.7 | 77.1 | 77.125
7.1 | 71.7 | 75 | 76.3 | 76.65 | 76.85
7.2 | 71.2286 | 74.4 | 75.85 | 76.3 | 76.425
7.3 | 70.725 | 74 | 75.5 | 75.95 | 76.05
7.4 | 70.4 | 73.575 | 75.1111 | 75.6 | 75.68
7.5 | 69.925 | 73.15 | 74.6875 | 75.25 | 75.3
7.6 | 69.4333 | 72.7875 | 74.3125 | 74.825 | 74.95
7.7 | 69.04 | 72.4125 | 73.975 | 74.45 | 74.6
7.8 | 68.52 | 71.95 | 73.625 | 74.0429 | 74.15
7.9 | 68.15 | 71.6 | 73.175 | 73.7333 | 73.85
8 | 67.725 | 71.225 | 72.8333 | 73.3429 | 73.475
8.1 | 67.275 | 70.8 | 72.4 | 73 | 73.05
8.2 | 66.85 | 70.3429 | 72.05 | 72.75 | 72.7875
8.3 | 66.325 | 69.85 | 71.7 | 72.2875 | 72.4
8.4 | 65.94 | 69.5 | 71.3 | 71.9125 | 72
8.5 | 65.5 | 69.2 | 71 | 71.625 | 71.75
8.6 | 65.1 | 68.8 | 70.5667 | 71.2625 | 71.4286
8.7 | 64.7 | 68.4 | 70.25 | 70.8875 | 71.05
8.8 | 64.3 | 67.9 | 69.9 | 70.5571 | 70.6
8.9 | 63.925 | 67.5 | 69.5 | 70.1 | 70.325
9 | 63.375 | 67.15 | 69.1 | 69.85 | 69.95
9.1 | 63.05 | 66.7 | 68.8 | 69.5 | 69.66
9.2 | 62.625 | 66.3333 | 68.42 | 69.175 | 69.32
9.3 | 62.125 | 65.925 | 68.05 | 68.75 | 68.95
9.4 | 61.775 | 65.65 | 67.725 | 68.36 | 68.6
9.5 | 61.38 | 65.225 | 67.35 | 68.025 | 68.25
9.6 | 61 | 64.8 | 66.9 | 67.775 | 67.9
9.7 | 60.5333 | 64.45 | 66.55 | 67.325 | 67.55
9.8 | 60.125 | 64.1 | 66.2 | 67 | 67.3
9.9 | 59.75 | 63.7 | 65.9 | 66.75 | 66.9
10 | 59.375 | 63.3 | 65.5 | 66.4 | 66.55
10.1 | 59 | 62.96 | 65.14 | 66.05 | 66.275
10.2 | 58.65 | 62.6 | 64.8667 | 65.7 | 65.925
10.3 | 58.2 | 62.22 | 64.475 | 65.375 | 65.575
10.4 | 57.85 | 61.75 | 64.125 | 65.05 | 65.225
10.5 | 57.4375 | 61.45 | 63.775 | 64.65 | 64.9
10.6 | 57.15 | 61.025 | 63.525 | 64.3333 | 64.525
10.7 | 56.65 | 60.675 | 63.1 | 64 | 64.175
10.8 | 56.3 | 60.35 | 62.8 | 63.75 | 63.85
10.9 | 55.9714 | 60 | 62.5 | 63.5 | 63.575
11 | 55.6 | 59.6667 | 62.15 | 63.15 | 63.275
11.1 | 55.25 | 59.35 | 61.7 | 62.8 | 62.975
11.2 | 54.9 | 58.8667 | 61.45 | 62.46 | 62.675
11.3 | 54.5 | 58.575 | 61.1 | 62.2 | 62.36
11.4 | 54.15 | 58.2333 | 60.875 | 61.9 | 62.1
11.5 | 53.775 | 57.85 | 60.525 | 61.5375 | 61.7857
11.6 | 53.4667 | 57.575 | 60.175 | 61.1833 | 61.35
11.7 | 53.1125 | 57.3 | 59.925 | 60.9 | 61.1
11.8 | 52.775 | 57 | 59.6 | 60.55 | 60.825
11.9 | 52.45 | 56.675 | 59.25 | 60.225 | 60.5
12 | 52.175 | 56.2 | 58.925 | 59.95 | 60.25
12.1 | 51.85 | 55.85 | 58.575 | 59.65 | 59.875
12.2 | 51.48 | 55.55 | 58.225 | 59.35 | 59.625
12.3 | 51.2 | 55.2286 | 57.9667 | 59.075 | 59.275
12.4 | 50.8333 | 54.95 | 57.675 | 58.8 | 59.025
12.5 | 50.5 | 54.625 | 57.3375 | 58.52 | 58.7
12.6 | 50.2 | 54.375 | 57.075 | 58.175 | 58.48
12.7 | 49.84 | 54.025 | 56.75 | 57.825 | 58.2
12.8 | 49.5 | 53.65 | 56.4 | 57.58 | 57.8
12.9 | 49.2333 | 53.2667 | 56.15 | 57.34 | 57.55
13 | 48.86 | 53 | 55.9 | 57 | 57.24
13.1 | 48.55 | 52.6667 | 55.45 | 56.75 | 57.02
13.2 | 48.3 | 52.4 | 55.2 | 56.4 | 56.725
13.3 | 47.95 | 52.075 | 54.975 | 56.0857 | 56.375
13.4 | 47.65 | 51.725 | 54.65 | 55.825 | 56.025
13.5 | 47.375 | 51.475 | 54.3 | 55.55 | 55.875
13.6 | 47.125 | 51.225 | 54.0667 | 55.325 | 55.5667
13.7 | 46.775 | 50.875 | 53.8 | 54.975 | 55.3
13.8 | 46.44 | 50.6 | 53.5 | 54.8 | 54.9667
13.9 | 46.2 | 50.28 | 53.2 | 54.4 | 54.6667
14 | 45.85 | 49.9667 | 52.9 | 54.15 | 54.4
14.1 | 45.5 | 49.6333 | 52.65 | 53.9 | 54.25
14.2 | 45.3 | 49.4 | 52.4 | 53.55 | 53.925
14.3 | 45 | 49.14 | 52.075 | 53.3111 | 53.675
14.4 | 44.65 | 48.8 | 51.825 | 53.075 | 53.25
14.5 | 44.4 | 48.5 | 51.475 | 52.825 | 53.075
14.6 | 44.16 | 48.3 | 51.225 | 52.475 | 52.825
14.7 | 43.925 | 47.9667 | 50.95 | 52.24 | 52.5625
14.8 | 43.5333 | 47.7 | 50.625 | 52 | 52.3
14.9 | 43.3 | 47.375 | 50.3 | 51.7 | 52.05
15 | 43.025 | 47.1 | 50.1 | 51.4 | 51.7333
15.1 | 42.65 | 46.8333 | 49.8 | 51.15 | 51.5
15.2 | 42.5 | 46.55 | 49.56 | 50.8 | 51.175
15.3 | 42.2 | 46.325 | 49.325 | 50.575 | 50.9
15.4 | 41.85 | 45.975 | 49.0333 | 50.42 | 50.675
15.5 | 41.65 | 45.725 | 48.7333 | 50.075 | 50.4667
15.6 | 41.325 | 45.375 | 48.44 | 49.825 | 50.175
15.7 | 41.075 | 45.1333 | 48.2222 | 49.575 | 49.875
15.8 | 40.9 | 44.8 | 47.96 | 49.325 | 49.64
15.9 | 40.6 | 44.55 | 47.6286 | 49.08 | 49.4
16 | 40.35 | 44.34 | 47.4714 | 48.84 | 49.125
16.1 | 40.2 | 44.1 | 47.2 | 48.55 | 48.88
16.2 | 39.85 | 43.9 | 47 | 48.3571 | 48.7
16.3 | 39.6 | 43.6 | 46.7 | 48.08 | 48.4333
16.4 | 39.35 | 43.3 | 46.4 | 47.8333 | 48.1667
16.5 | 39.1 | 43.075 | 46.15 | 47.55 | 47.925
16.6 | 38.875 | 42.8 | 45.9 | 47.3 | 47.75
16.7 | 38.6 | 42.5 | 45.65 | 47.05 | 47.425
16.8 | 38.375 | 42.3 | 45.38 | 46.85 | 47.2
16.9 | 38.125 | 42 | 45.15 | 46.6 | 46.9667
17 | 37.9 | 41.7667 | 44.92 | 46.35 | 46.75
17.1 | 37.675 | 41.5 | 44.68 | 46.1667 | 46.4
17.2 | 37.375 | 41.3429 | 44.45 | 45.8 | 46.25
17.3 | 37.125 | 41.08 | 44.2 | 45.675 | 46.025
17.4 | 36.9 | 40.875 | 43.975 | 45.38 | 45.7857
17.5 | 36.725 | 40.625 | 43.72 | 45.2111 | 45.525
17.6 | 36.5 | 40.4167 | 43.55 | 44.925 | 45.3429
17.7 | 36.3 | 40.1 | 43.225 | 44.7 | 45.075
17.8 | 36.1 | 39.875 | 43.075 | 44.525 | 44.8625
17.9 | 35.85 | 39.7125 | 42.825 | 44.2 | 44.675
18 | 35.6 | 39.44 | 42.575 | 44 | 44.425
18.1 | 35.375 | 39.28 | 42.325 | 43.8 | 44.275
18.2 | 35.2 | 39.05 | 42.1 | 43.625 | 43.925
18.3 | 34.975 | 38.7 | 41.85 | 43.375 | 43.78
18.4 | 34.7 | 38.45 | 41.675 | 43.18 | 43.54
18.5 | 34.4 | 38.325 | 41.4 | 42.95 | 43.3
18.6 | 34.225 | 38.075 | 41.225 | 42.7333 | 43.05
18.7 | 34 | 37.825 | 41 | 42.5 | 42.8
18.8 | 33.85 | 37.575 | 40.725 | 42.2 | 42.575
18.9 | 33.6 | 37.34 | 40.5 | 42 | 42.425
19 | 33.375 | 37.16 | 40.34 | 41.8 | 42.275
19.1 | 33.2 | 36.95 | 40.1 | 41.5667 | 42.025
19.2 | 32.975 | 36.7 | 39.85 | 41.4 | 41.78
19.3 | 32.725 | 36.55 | 39.7 | 41.15 | 41.54
19.4 | 32.58 | 36.3 | 39.45 | 40.9 | 41.3
19.5 | 32.35 | 36.05 | 39.2 | 40.7 | 41.1667
19.6 | 32.1 | 35.9 | 39.1 | 40.5 | 40.95
19.7 | 32 | 35.675 | 38.8429 | 40.35 | 40.75
19.8 | 31.7 | 35.5 | 38.6 | 40.16 | 40.5
19.9 | 31.575 | 35.275 | 38.34 | 39.9143 | 40.3
20 | 31.3571 | 35.025 | 38.125 | 39.65 | 40.16
20.1 | 31.175 | 34.8167 | 37.9 | 39.45 | 39.875
20.2 | 30.9625 | 34.6167 | 37.74 | 39.2667 | 39.725
20.3 | 30.8375 | 34.4375 | 37.58 | 39.075 | 39.45
20.4 | 30.625 | 34.175 | 37.35 | 38.85 | 39.3
20.5 | 30.425 | 34 | 37.1 | 38.7125 | 39.075
20.6 | 30.2 | 33.8 | 36.875 | 38.475 | 38.9
20.7 | 30.025 | 33.55 | 36.725 | 38.25 | 38.6667
20.8 | 29.875 | 33.3 | 36.475 | 38.075 | 38.475
20.9 | 29.65 | 33.2 | 36.325 | 37.85 | 38.3
21 | 29.5 | 33 | 36.075 | 37.6 | 38.1
21.1 | 29.25 | 32.8 | 35.95 | 37.5 | 37.9
21.2 | 29.1 | 32.7 | 35.72 | 37.3 | 37.68
21.3 | 28.95 | 32.45 | 35.55 | 37.1 | 37.55
21.4 | 28.8 | 32.22 | 35.4 | 36.9333 | 37.3
21.5 | 28.6 | 32.125 | 35.2 | 36.7 | 37.05
21.6 | 28.425 | 31.875 | 35 | 36.525 | 36.9
21.7 | 28.2 | 31.725 | 34.82 | 36.3 | 36.7333
21.8 | 28.1 | 31.5 | 34.625 | 36.1333 | 36.55
21.9 | 27.95 | 31.3 | 34.45 | 36 | 36.4
22 | 27.75 | 31.1 | 34.225 | 35.7667 | 36.2
22.1 | 27.6 | 31.025 | 34.075 | 35.6 | 35.95
22.2 | 27.35 | 30.8 | 33.9 | 35.4 | 35.86
22.3 | 27.2 | 30.625 | 33.75 | 35.2667 | 35.6
22.4 | 27.05 | 30.4 | 33.5 | 35.05 | 35.4286
22.5 | 26.9 | 30.225 | 33.3667 | 34.9 | 35.3111
22.6 | 26.7 | 30.1 | 33.1667 | 34.6667 | 35.075
22.7 | 26.525 | 29.9 | 33 | 34.5 | 34.8625
22.8 | 26.375 | 29.8 | 32.825 | 34.3 | 34.7375
22.9 | 26.2 | 29.6 | 32.6 | 34.125 | 34.525
23 | 26.075 | 29.4 | 32.475 | 33.975 | 34.3875
23.1 | 25.9 | 29.3 | 32.3 | 33.825 | 34.225
23.2 | 25.8 | 29.05 | 32.1 | 33.65 | 34.0286
23.3 | 25.6 | 28.9 | 31.9 | 33.4 | 33.8
23.4 | 25.5 | 28.7 | 31.725 | 33.2333 | 33.75
23.5 | 25.2667 | 28.6 | 31.5571 | 33.1 | 33.5
23.6 | 25.15 | 28.35 | 31.375 | 32.925 | 33.35
23.7 | 25.025 | 28.2 | 31.225 | 32.775 | 33.1
23.8 | 24.86 | 28.125 | 31.075 | 32.625 | 33.0667
23.9 | 24.7 | 27.9 | 30.9 | 32.48 | 32.825
24 | 24.5 | 27.8 | 30.775 | 32.3 | 32.6
24.1 | 24.425 | 27.675 | 30.525 | 32.1 | 32.525
24.2 | 24.275 | 27.425 | 30.4 | 31.95 | 32.4
24.3 | 24.125 | 27.325 | 30.24 | 31.8 | 32.15
24.4 | 23.975 | 27.1 | 30.0833 | 31.6 | 32
24.5 | 23.8 | 26.94 | 29.9167 | 31.4 | 31.85
24.6 | 23.68 | 26.8 | 29.7833 | 31.275 | 31.7
24.7 | 23.5 | 26.6667 | 29.6167 | 31.1 | 31.5
24.8 | 23.4 | 26.48 | 29.44 | 30.9667 | 31.325
24.9 | 23.2 | 26.3 | 29.24 | 30.8 | 31.175
25 | 23.1 | 26.2 | 29.1 | 30.6375 | 31.025
25.1 | 22.9833 | 26 | 28.94 | 30.5 | 30.9
25.2 | 22.8167 | 25.85 | 28.8 | 30.3375 | 30.7
25.3 | 22.7 | 25.7 | 28.65 | 30.125 | 30.6
25.4 | 22.6 | 25.6 | 28.5 | 30 | 30.4
25.5 | 22.3714 | 25.45 | 28.3 | 29.85 | 30.3
25.6 | 22.3 | 25.3 | 28.2 | 29.6429 | 30.06
25.7 | 22.1875 | 25.16 | 28.075 | 29.5 | 29.925
25.8 | 22.025 | 25 | 27.825 | 29.3667 | 29.8
25.9 | 21.9 | 24.9 | 27.7 | 29.275 | 29.7
26 | 21.775 | 24.725 | 27.525 | 29.1 | 29.5
26.1 | 21.65 | 24.6 | 27.4 | 28.95 | 29.35
26.2 | 21.5 | 24.45 | 27.3 | 28.8 | 29.225
26.3 | 21.35 | 24.3 | 27.2 | 28.675 | 29.1
26.4 | 21.2 | 24.15 | 27 | 28.4571 | 28.94
26.5 | 21.06 | 24 | 26.9 | 28.35 | 28.8
26.6 | 21 | 23.9 | 26.6667 | 28.2 | 28.65
26.7 | 20.8 | 23.7333 | 26.525 | 28.05 | 28.5
26.8 | 20.725 | 23.6 | 26.4 | 27.9 | 28.3
26.9 | 20.5 | 23.4833 | 26.2833 | 27.7667 | 28.2
27 | 20.425 | 23.375 | 26.16 | 27.6 | 28.05
27.1 | 20.375 | 23.2 | 26.025 | 27.475 | 27.9
27.2 | 20.25 | 23.075 | 25.875 | 27.3 | 27.8
27.3 | 20.1 | 22.925 | 25.75 | 27.2 | 27.6
27.4 | 20 | 22.8 | 25.6 | 27.025 | 27.5
27.5 | 19.875 | 22.66 | 25.475 | 26.975 | 27.35
27.6 | 19.75 | 22.525 | 25.3 | 26.8 | 27.2
27.7 | 19.6 | 22.4 | 25.2 | 26.6333 | 27.1
27.8 | 19.5 | 22.3 | 25.025 | 26.5 | 26.9
27.9 | 19.4 | 22.175 | 24.9 | 26.4 | 26.8
28 | 19.3 | 22 | 24.8 | 26.2 | 26.7
28.1 | 19.1 | 21.9 | 24.675 | 26.1 | 26.52
28.2 | 19.05 | 21.8 | 24.5 | 26 | 26.34
28.3 | 18.9 | 21.7 | 24.4 | 25.85 | 26.225
28.4 | 18.8 | 21.5 | 24.3 | 25.7 | 26.175
28.5 | 18.725 | 21.4 | 24.1 | 25.55 | 25.9625
28.6 | 18.6 | 21.3 | 24 | 25.4 | 25.875
28.7 | 18.5 | 21.2 | 23.9 | 25.3 | 25.7
28.8 | 18.4 | 21.1 | 23.75 | 25.2 | 25.5875
28.9 | 18.275 | 21 | 23.6 | 25 | 25.44
29 | 18.125 | 20.8333 | 23.5 | 24.9 | 25.2875
29.1 | 18.075 | 20.7 | 23.4 | 24.775 | 25.1625
29.2 | 17.925 | 20.6 | 23.2 | 24.7 | 25.0444
29.3 | 17.8 | 20.5 | 23.125 | 24.5 | 24.925
29.4 | 17.7 | 20.4 | 23 | 24.425 | 24.8
29.5 | 17.6333 | 20.275 | 22.9 | 24.3 | 24.6667
29.6 | 17.6 | 20.125 | 22.7 | 24.125 | 24.5667
29.7 | 17.4333 | 20.0167 | 22.6 | 24 | 24.425
29.8 | 17.3 | 19.9 | 22.5 | 23.9333 | 24.3111
29.9 | 17.2 | 19.8 | 22.4 | 23.8 | 24.2
30 | 17.1 | 19.6625 | 22.3 | 23.7 | 24.075

## SSD120 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 45 | 50 | 57.5 | 63.8 | 66.2
0.1 | 50.4 | 55.1 | 62.3 | 68.48 | 70.7
0.2 | 59 | 62.8 | 69.65 | 74.8 | 76.8
0.3 | 69.5 | 72.3 | 78.2 | 81.6 | 83.55
0.4 | 77.4 | 79.6667 | 84.15 | 87.15 | 88.4
0.5 | 83.45 | 85.08 | 88.7 | 90.9571 | 92.1
0.6 | 88 | 89.45 | 92.05 | 93.98 | 94.75
0.7 | 91.5 | 92.425 | 94.6 | 95.85 | 96.55
0.8 | 93.95 | 94.725 | 96.3 | 97.45 | 97.9
0.9 | 95.9333 | 96.575 | 97.7 | 98.58 | 98.84
1 | 97.4333 | 97.8429 | 98.725 | 99.3333 | 99.375
1.1 | 98.5 | 98.8625 | 99.3667 | 99.8 | 99.825
1.2 | 99.275 | 99.42 | 99.7667 | 99.975 | 100
1.3 | 99.775 | 99.8 | 99.925 | 100 | 100
1.4 | 99.925 | 99.98 | 100 | 100 | 100
1.5 | 100 | 100 | 100 | 100 | 99.9
1.6 | 100 | 100 | 100 | 99.8 | 99.8
1.7 | 99.8667 | 99.7 | 99.775 | 99.6 | 99.6333
1.8 | 99.6 | 99.575 | 99.525 | 99.4 | 99.3
1.9 | 99.4 | 99.4 | 99.3 | 99.2 | 99.05
2 | 98.9667 | 99.1667 | 99.125 | 98.85 | 98.8
2.1 | 98.725 | 98.7 | 98.7 | 98.65 | 98.5667
2.2 | 98.3 | 98.4625 | 98.35 | 98.325 | 98.2
2.3 | 97.8 | 98.05 | 98 | 97.975 | 97.85
2.4 | 97.35 | 97.65 | 97.8 | 97.68 | 97.6
2.5 | 96.9 | 97.425 | 97.4 | 97.4 | 97.35
2.6 | 96.6 | 97 | 97.15 | 97 | 97.05
2.7 | 95.9 | 96.6 | 96.7667 | 96.725 | 96.675
2.8 | 95.6 | 96.25 | 96.35 | 96.35 | 96.325
2.9 | 95.14 | 96 | 96.1333 | 96.1143 | 96.025
3 | 94.675 | 95.5667 | 95.7 | 95.7 | 95.7667
3.1 | 94.325 | 95.2 | 95.4 | 95.4571 | 95.4
3.2 | 93.85 | 94.82 | 95.1 | 95.1111 | 95.125
3.3 | 93.325 | 94.425 | 94.75 | 94.8125 | 94.68
3.4 | 92.75 | 93.92 | 94.5 | 94.425 | 94.4333
3.5 | 92.3 | 93.625 | 94.125 | 94.1 | 94.2
3.6 | 91.98 | 93.225 | 93.65 | 93.8 | 93.84
3.7 | 91.4 | 92.7 | 93.325 | 93.55 | 93.45
3.8 | 91.1 | 92.4 | 92.975 | 93.12 | 93.125
3.9 | 90.6 | 92.1 | 92.65 | 92.825 | 92.85
4 | 90.2 | 91.8 | 92.35 | 92.475 | 92.525
4.1 | 89.65 | 91.38 | 92.025 | 92.15 | 92.1
4.2 | 89.3 | 90.9 | 91.56 | 91.84 | 91.8
4.3 | 88.8 | 90.5 | 91.14 | 91.45 | 91.5
4.4 | 88.3 | 90.15 | 90.8 | 91.1 | 91.15
4.5 | 87.95 | 89.84 | 90.6333 | 90.86 | 90.9
4.6 | 87.3333 | 89.4333 | 90.25 | 90.55 | 90.56
4.7 | 86.9333 | 89.06 | 89.8 | 90.125 | 90.25
4.8 | 86.6143 | 88.75 | 89.45 | 89.775 | 89.925
4.9 | 86.2 | 88.375 | 89.1 | 89.5333 | 89.7
5 | 85.75 | 88 | 88.8 | 89.2 | 89.3
5.1 | 85.15 | 87.55 | 88.4 | 88.7333 | 89
5.2 | 84.6 | 87.125 | 88.1333 | 88.48 | 88.6
5.3 | 84.2 | 86.9167 | 87.85 | 88.2333 | 88.4
5.4 | 83.7 | 86.425 | 87.45 | 88 | 87.9
5.5 | 83.3 | 86.05 | 87.1 | 87.6 | 87.6
5.6 | 82.9 | 85.55 | 86.7 | 87.3667 | 87.3333
5.7 | 82.55 | 85.2286 | 86.375 | 86.9 | 87.02
5.8 | 81.95 | 84.925 | 86.025 | 86.6 | 86.7
5.9 | 81.575 | 84.3667 | 85.775 | 86.3 | 86.3333
6 | 81.225 | 84.1 | 85.525 | 85.875 | 85.975
6.1 | 80.875 | 83.7125 | 85.18 | 85.725 | 85.74
6.2 | 80.35 | 83.3375 | 84.85 | 85.35 | 85.375
6.3 | 80 | 83 | 84.54 | 85.05 | 85.05
6.4 | 79.4857 | 82.5667 | 84.1 | 84.76 | 84.7
6.5 | 79 | 82.15 | 83.75 | 84.4 | 84.4
6.6 | 78.55 | 81.92 | 83.425 | 84 | 84.1
6.7 | 78.16 | 81.45 | 83.0625 | 83.75 | 83.75
6.8 | 77.65 | 81.125 | 82.725 | 83.5 | 83.325
6.9 | 77.35 | 80.675 | 82.4125 | 83.1333 | 83.1
7 | 76.85 | 80.425 | 82.0375 | 82.8333 | 82.85
7.1 | 76.425 | 80.1 | 81.775 | 82.4 | 82.54
7.2 | 76.125 | 79.55 | 81.5 | 82.2 | 82.25
7.3 | 75.65 | 79.3 | 81.2 | 81.825 | 81.925
7.4 | 75.2 | 78.85 | 80.95 | 81.525 | 81.7
7.5 | 74.8 | 78.6 | 80.55 | 81.275 | 81.25
7.6 | 74.4 | 78.25 | 80.125 | 80.925 | 81
7.7 | 74.05 | 77.9 | 79.875 | 80.56 | 80.65
7.8 | 73.6 | 77.6667 | 79.54 | 80.2 | 80.4
7.9 | 73.15 | 77.2 | 79.26 | 79.9 | 80.16
8 | 72.85 | 76.9333 | 78.925 | 79.6 | 79.825
8.1 | 72.5 | 76.45 | 78.575 | 79.2 | 79.55
8.2 | 71.95 | 76.05 | 78.34 | 79.0333 | 79.0667
8.3 | 71.625 | 75.68 | 78 | 78.62 | 78.85
8.4 | 71.2 | 75.38 | 77.6 | 78.26 | 78.4
8.5 | 70.75 | 75.12 | 77.1 | 78.125 | 78.1667
8.6 | 70.3571 | 74.7 | 76.86 | 77.68 | 77.98
8.7 | 69.9 | 74.1667 | 76.625 | 77.44 | 77.6
8.8 | 69.55 | 73.9 | 76.35 | 77.2 | 77.25
8.9 | 69.225 | 73.5857 | 75.95 | 76.8333 | 77
9 | 68.975 | 73.18 | 75.675 | 76.475 | 76.65
9.1 | 68.55 | 72.8 | 75.4 | 76.2 | 76.24
9.2 | 68.075 | 72.525 | 75.1 | 75.875 | 76
9.3 | 67.65 | 72.2111 | 74.7 | 75.58 | 75.75
9.4 | 67.3 | 71.85 | 74.3 | 75.3 | 75.475
9.5 | 66.9 | 71.525 | 74.05 | 75 | 75.15
9.6 | 66.5 | 71.225 | 73.8 | 74.6 | 74.9
9.7 | 66.2429 | 70.8 | 73.4 | 74.3 | 74.6
9.8 | 65.8 | 70.55 | 73.04 | 74.05 | 74.2
9.9 | 65.3 | 70.2 | 72.75 | 73.8 | 73.85
10 | 65.1 | 69.8 | 72.425 | 73.5 | 73.6
10.1 | 64.65 | 69.5 | 72.12 | 73.225 | 73.36
10.2 | 64.325 | 69.15 | 71.825 | 72.875 | 73
10.3 | 63.975 | 68.75 | 71.575 | 72.55 | 72.8
10.4 | 63.625 | 68.475 | 71.15 | 72.3 | 72.5
10.5 | 63.3 | 68.125 | 70.9 | 72.0667 | 72.1
10.6 | 62.9333 | 67.875 | 70.64 | 71.7 | 71.8333
10.7 | 62.5 | 67.45 | 70.3 | 71.4 | 71.6
10.8 | 62 | 67.1667 | 70.05 | 71.1667 | 71.325
10.9 | 61.75 | 66.9 | 69.8 | 70.875 | 70.975
11 | 61.5 | 66.45 | 69.375 | 70.625 | 70.725
11.1 | 61.2 | 66.2 | 69.0875 | 70.25 | 70.475
11.2 | 60.7667 | 65.95 | 68.775 | 69.94 | 70.225
11.3 | 60.35 | 65.6 | 68.525 | 69.6 | 69.875
11.4 | 60.075 | 65.22 | 68.1857 | 69.3333 | 69.55
11.5 | 59.725 | 64.925 | 67.9 | 69.1 | 69.35
11.6 | 59.475 | 64.575 | 67.7 | 68.8 | 68.975
11.7 | 59.125 | 64.3 | 67.4 | 68.5 | 68.825
11.8 | 58.85 | 64.075 | 67.1 | 68.2333 | 68.475
11.9 | 58.4 | 63.675 | 66.925 | 67.95 | 68.2625
12 | 58 | 63.375 | 66.48 | 67.7 | 67.925
12.1 | 57.75 | 63.05 | 66.25 | 67.4 | 67.7167
12.2 | 57.4 | 62.7 | 65.9 | 67.075 | 67.425
12.3 | 57.16 | 62.4 | 65.65 | 66.9 | 67.175
12.4 | 56.7 | 62.1 | 65.3 | 66.6 | 66.825
12.5 | 56.4286 | 61.85 | 65.05 | 66.4 | 66.5125
12.6 | 56.125 | 61.5 | 64.7 | 66.05 | 66.3
12.7 | 55.85 | 61.175 | 64.475 | 65.8 | 66
12.8 | 55.45 | 60.825 | 64.2714 | 65.4333 | 65.76
12.9 | 55.225 | 60.6167 | 63.9 | 65.2 | 65.4
13 | 54.875 | 60.275 | 63.625 | 64.9333 | 65.2
13.1 | 54.64 | 59.9875 | 63.25 | 64.6667 | 64.9667
13.2 | 54.35 | 59.675 | 63.02 | 64.4 | 64.6
13.3 | 54.025 | 59.3375 | 62.7125 | 64.15 | 64.4
13.4 | 53.74 | 59.1 | 62.5125 | 63.9 | 64.05
13.5 | 53.4333 | 58.76 | 62.2 | 63.6 | 63.8667
13.6 | 53.1375 | 58.525 | 61.9667 | 63.3 | 63.6
13.7 | 52.775 | 58.35 | 61.625 | 63.05 | 63.25
13.8 | 52.525 | 57.925 | 61.35 | 62.75 | 63
13.9 | 52.2125 | 57.65 | 61.15 | 62.5 | 62.775
14 | 51.8375 | 57.425 | 60.825 | 62.275 | 62.525
14.1 | 51.55 | 57.1 | 60.6 | 62.02 | 62.275
14.2 | 51.2714 | 56.85 | 60.35 | 61.8 | 62.025
14.3 | 50.95 | 56.5 | 60.1 | 61.4 | 61.775
14.4 | 50.7 | 56.1 | 59.75 | 61.125 | 61.525
14.5 | 50.36 | 55.85 | 59.525 | 60.9333 | 61.225
14.6 | 50.2 | 55.55 | 59.275 | 60.575 | 60.9333
14.7 | 49.875 | 55.25 | 58.9143 | 60.4 | 60.65
14.8 | 49.64 | 55.0125 | 58.6286 | 60.175 | 60.4
14.9 | 49.3 | 54.775 | 58.3556 | 59.825 | 60.175
15 | 48.95 | 54.48 | 58.2375 | 59.6 | 59.925
15.1 | 48.7 | 54.2 | 57.9375 | 59.25 | 59.6125
15.2 | 48.45 | 53.9 | 57.675 | 58.9375 | 59.3
15.3 | 48.1 | 53.6 | 57.3875 | 58.7444 | 59.175
15.4 | 47.875 | 53.35 | 57.075 | 58.4875 | 58.925
15.5 | 47.6 | 53.2 | 56.85 | 58.35 | 58.75
15.6 | 47.35 | 52.8714 | 56.525 | 58.08 | 58.44
15.7 | 47.025 | 52.7167 | 56.275 | 57.85 | 58.125
15.8 | 46.78 | 52.35 | 55.95 | 57.65 | 58
15.9 | 46.6143 | 52.12 | 55.8 | 57.425 | 57.7
16 | 46.3714 | 51.85 | 55.58 | 57.15 | 57.4
16.1 | 46.1 | 51.525 | 55.3667 | 56.925 | 57.2
16.2 | 45.8333 | 51.35 | 55 | 56.725 | 56.975
16.3 | 45.65 | 51.05 | 54.85 | 56.46 | 56.825
16.4 | 45.4 | 50.8 | 54.6 | 56.14 | 56.65
16.5 | 45.05 | 50.55 | 54.35 | 55.9 | 56.325
16.6 | 44.8 | 50.2 | 54.2 | 55.75 | 56
16.7 | 44.6 | 50.05 | 53.85 | 55.5 | 55.75
16.8 | 44.3 | 49.825 | 53.6 | 55.25 | 55.5
16.9 | 44.05 | 49.475 | 53.35 | 55 | 55.35
17 | 43.825 | 49.225 | 53.1 | 54.775 | 55.1
17.1 | 43.65 | 49.05 | 52.95 | 54.625 | 54.85
17.2 | 43.4 | 48.75 | 52.7 | 54.375 | 54.6
17.3 | 43.25 | 48.575 | 52.4 | 54.05 | 54.35
17.4 | 42.9 | 48.4 | 52.2 | 53.875 | 54.12
17.5 | 42.65 | 48 | 51.96 | 53.625 | 53.875
17.6 | 42.3 | 47.8 | 51.7 | 53.325 | 53.625
17.7 | 42.05 | 47.475 | 51.475 | 53.1 | 53.5
17.8 | 41.925 | 47.2625 | 51.2333 | 53 | 53.32
17.9 | 41.675 | 47.075 | 51 | 52.6 | 53.025
18 | 41.425 | 46.85 | 50.8 | 52.3667 | 52.75
18.1 | 41.175 | 46.62 | 50.6 | 52.1333 | 52.525
18.2 | 40.94 | 46.325 | 50.3 | 51.9 | 52.275
18.3 | 40.7 | 46.075 | 50.06 | 51.75 | 52.1
18.4 | 40.5 | 45.9 | 49.9 | 51.6 | 51.9
18.5 | 40.3 | 45.7 | 49.7167 | 51.35 | 51.64
18.6 | 40.05 | 45.5 | 49.375 | 51.1 | 51.4
18.7 | 39.8 | 45.15 | 49.125 | 50.875 | 51.25
18.8 | 39.65 | 45 | 48.9444 | 50.625 | 51
18.9 | 39.4 | 44.75 | 48.75 | 50.375 | 50.8
19 | 39.1 | 44.5 | 48.5 | 50.225 | 50.6
19.1 | 38.9 | 44.3 | 48.175 | 49.975 | 50.3
19.2 | 38.7 | 44.1 | 48.02 | 49.72 | 50.15
19.3 | 38.525 | 43.94 | 47.775 | 49.575 | 49.875
19.4 | 38.2 | 43.55 | 47.575 | 49.4 | 49.7
19.5 | 38.0667 | 43.3 | 47.325 | 49.15 | 49.4
19.6 | 37.9 | 43.15 | 47.15 | 48.9 | 49.25
19.7 | 37.6857 | 42.9 | 46.9 | 48.675 | 48.975
19.8 | 37.5 | 42.75 | 46.68 | 48.525 | 48.7
19.9 | 37.25 | 42.4 | 46.5 | 48.35 | 48.6
20 | 37 | 42.175 | 46.375 | 48.1 | 48.38
20.1 | 36.7833 | 41.925 | 46.1 | 47.75 | 48.2
20.2 | 36.6 | 41.775 | 45.85 | 47.6 | 48
20.3 | 36.34 | 41.525 | 45.575 | 47.35 | 47.7667
20.4 | 36.2 | 41.425 | 45.425 | 47.22 | 47.525
20.5 | 36 | 41.275 | 45.22 | 46.975 | 47.3
20.6 | 35.9 | 41.04 | 45.05 | 46.725 | 47.1
20.7 | 35.6333 | 40.8 | 44.9 | 46.575 | 46.975
20.8 | 35.4 | 40.65 | 44.6 | 46.425 | 46.725
20.9 | 35.1333 | 40.3 | 44.4 | 46.275 | 46.575
21 | 34.92 | 40.15 | 44.25 | 45.95 | 46.425
21.1 | 34.875 | 39.925 | 44 | 45.8 | 46.2
21.2 | 34.625 | 39.775 | 43.75 | 45.65 | 45.9
21.3 | 34.4 | 39.525 | 43.55 | 45.5 | 45.8
21.4 | 34.2333 | 39.375 | 43.4143 | 45.25 | 45.6
21.5 | 34.08 | 39.15 | 43.2 | 45 | 45.35
21.6 | 33.95 | 38.9625 | 43.025 | 44.8 | 45.2
21.7 | 33.775 | 38.775 | 42.7833 | 44.625 | 45.075
21.8 | 33.525 | 38.55 | 42.575 | 44.4 | 44.825
21.9 | 33.3 | 38.4 | 42.4 | 44.225 | 44.575
22 | 33.125 | 38.175 | 42.2 | 43.975 | 44.3
22.1 | 32.875 | 37.9571 | 42.05 | 43.725 | 44.175
22.2 | 32.625 | 37.75 | 41.8333 | 43.6 | 44.025
22.3 | 32.5167 | 37.6111 | 41.6 | 43.4 | 43.775
22.4 | 32.3875 | 37.4429 | 41.4 | 43.3 | 43.55
22.5 | 32.2625 | 37.25 | 41.22 | 43.15 | 43.4
22.6 | 32.075 | 37 | 41.075 | 42.9 | 43.25
22.7 | 31.9125 | 36.925 | 40.925 | 42.65 | 43
22.8 | 31.8 | 36.7 | 40.675 | 42.425 | 42.76
22.9 | 31.5714 | 36.55 | 40.425 | 42.275 | 42.625
23 | 31.4 | 36.3 | 40.3 | 42.1125 | 42.5
23.1 | 31.25 | 36.1333 | 40.15 | 41.975 | 42.35
23.2 | 31.1 | 35.9 | 39.9 | 41.725 | 42.0833
23.3 | 30.85 | 35.75 | 39.775 | 41.5375 | 41.875
23.4 | 30.675 | 35.625 | 39.6 | 41.3444 | 41.725
23.5 | 30.525 | 35.375 | 39.425 | 41.1 | 41.575
23.6 | 30.3 | 35.225 | 39.2 | 41.0143 | 41.34
23.7 | 30.2 | 35 | 39.075 | 40.88 | 41.2
23.8 | 30 | 34.84 | 38.8625 | 40.65 | 41
23.9 | 29.85 | 34.58 | 38.7375 | 40.4 | 40.9
24 | 29.7 | 34.4 | 38.525 | 40.3 | 40.7
24.1 | 29.5 | 34.3 | 38.3 | 40.15 | 40.45
24.2 | 29.3 | 34.175 | 38.2 | 39.9 | 40.32
24.3 | 29.2 | 33.925 | 37.95 | 39.7 | 40.125
24.4 | 29 | 33.8 | 37.7 | 39.5 | 39.9
24.5 | 28.9 | 33.625 | 37.55 | 39.35 | 39.725
24.6 | 28.7333 | 33.5 | 37.42 | 39.2 | 39.6
24.7 | 28.5333 | 33.35 | 37.2 | 39 | 39.5
24.8 | 28.35 | 33.1 | 37.1 | 38.85 | 39.3
24.9 | 28.3 | 32.96 | 36.875 | 38.7 | 39.1
25 | 28.15 | 32.725 | 36.7 | 38.575 | 38.9
25.1 | 28 | 32.575 | 36.58 | 38.35 | 38.7
25.2 | 27.8 | 32.425 | 36.34 | 38.175 | 38.55
25.3 | 27.625 | 32.275 | 36.2 | 38 | 38.4
25.4 | 27.475 | 32.125 | 36.05 | 37.8 | 38.2
25.5 | 27.3 | 31.95 | 35.8 | 37.7 | 38.1
25.6 | 27.175 | 31.8 | 35.7 | 37.475 | 37.9
25.7 | 26.95 | 31.55 | 35.5 | 37.325 | 37.725
25.8 | 26.85 | 31.4 | 35.36 | 37.2 | 37.5667
25.9 | 26.7 | 31.25 | 35.1571 | 37 | 37.38
26 | 26.55 | 31.1 | 34.9857 | 36.8 | 37.2
26.1 | 26.4 | 30.9 | 34.82 | 36.6 | 37.05
26.2 | 26.3 | 30.825 | 34.7 | 36.5 | 36.9
26.3 | 26.2 | 30.675 | 34.6 | 36.3 | 36.8
26.4 | 26.075 | 30.5125 | 34.4 | 36.1667 | 36.6
26.5 | 25.925 | 30.375 | 34.24 | 36 | 36.4
26.6 | 25.675 | 30.1625 | 34.075 | 35.9 | 36.2667
26.7 | 25.6 | 30.0286 | 33.925 | 35.7 | 36.1
26.8 | 25.475 | 29.9 | 33.7 | 35.525 | 35.9
26.9 | 25.325 | 29.75 | 33.5667 | 35.375 | 35.775
27 | 25.2 | 29.6 | 33.4 | 35.225 | 35.625
27.1 | 25.05 | 29.5 | 33.1778 | 35.05 | 35.5
27.2 | 24.9 | 29.35 | 33.0625 | 34.8667 | 35.35
27.3 | 24.8 | 29.2 | 32.9375 | 34.7125 | 35.2
27.4 | 24.6 | 29.05 | 32.775 | 34.5875 | 35.05
27.5 | 24.55 | 28.8 | 32.64 | 34.4625 | 34.8
27.6 | 24.4 | 28.7 | 32.4 | 34.3375 | 34.66
27.7 | 24.25 | 28.5 | 32.3 | 34.125 | 34.525
27.8 | 24.1 | 28.35 | 32.2 | 33.8875 | 34.35
27.9 | 23.95 | 28.3 | 32.05 | 33.8 | 34.125
28 | 23.8 | 28.1 | 31.9 | 33.7 | 34.075
28.1 | 23.75 | 28 | 31.75 | 33.5 | 33.92
28.2 | 23.6 | 27.9 | 31.6 | 33.3 | 33.75
28.3 | 23.4 | 27.75 | 31.45 | 33.2 | 33.7
28.4 | 23.3 | 27.6 | 31.275 | 33.1 | 33.475
28.5 | 23.2 | 27.4 | 31.125 | 32.925 | 33.3
28.6 | 23.1 | 27.3 | 31 | 32.775 | 33.175
28.7 | 23 | 27.175 | 30.8 | 32.6 | 33.1
28.8 | 22.875 | 27.0667 | 30.7 | 32.48 | 32.9
28.9 | 22.725 | 26.9 | 30.55 | 32.35 | 32.8
29 | 22.6 | 26.7 | 30.42 | 32.2 | 32.7
29.1 | 22.5 | 26.65 | 30.3 | 32.075 | 32.45
29.2 | 22.38 | 26.5 | 30.1667 | 31.825 | 32.275
29.3 | 22.25 | 26.3 | 30.0125 | 31.7 | 32.15
29.4 | 22.1 | 26.15 | 29.8875 | 31.625 | 31.9625
29.5 | 22 | 26 | 29.725 | 31.4875 | 31.8286
29.6 | 21.8 | 25.9667 | 29.575 | 31.3625 | 31.7
29.7 | 21.7 | 25.84 | 29.425 | 31.2375 | 31.6
29.8 | 21.6 | 25.7 | 29.3 | 31.025 | 31.425
29.9 | 21.5 | 25.55 | 29.1667 | 30.8833 | 31.375
30 | 21.4 | 25.4 | 29 | 30.775 | 31.225

## Wedge 15 Lateral Profiles
Lateral Profiles
Wedge 15
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.5 cm FieldSize:4x4 cm2 | Depth:1.5 cm FieldSize:10x10 cm2 | Depth:1.5 cm FieldSize:20x20 cm2 | Depth:1.5 cm FieldSize:30x30 cm2 | Depth:1.5 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [8 empty] | 6.4
-22.8 | [8 empty] | 6.5 | 7.5 |  |  |  | 8.6667 | 9.8 |  |  |  | 14.6667 | 15.8 |  |  |  | 21.0429 | 22.4667
-22.7 | [8 empty] | 6.5286 | 7.5875 |  |  |  | 8.75 | 9.9 |  |  |  | 14.86 | 16.1 |  |  |  | 21.44 | 22.8
-22.6 | [8 empty] | 6.6 | 7.65 |  |  |  | 8.9167 | 10.0333 |  |  |  | 14.95 | 16.3 |  |  |  | 21.6667 | 23.0167
-22.5 | [8 empty] | 6.7286 | 7.7 |  |  |  | 9.1 | 10.1667 |  |  |  | 15.1667 | 16.52 |  |  |  | 22 | 23.2667
-22.4 | [8 empty] | 6.8625 | 7.85 |  |  |  | 9.2 | 10.375 |  |  |  | 15.4 | 16.7 |  |  |  | 22.3 | 23.7
-22.3 |  |  |  | 5.9 | 6.92 |  |  |  | 6.9 | 7.9 |  |  |  | 9.3167 | 10.5 |  |  |  | 15.6 | 16.9 |  |  |  | 22.5333 | 23.9333
-22.2 |  |  |  | 6 | 7 |  |  |  | 7 | 8 |  |  |  | 9.5 | 10.6 |  |  |  | 15.85 | 17.15 |  |  |  | 22.8 | 24.1833
-22.1 |  |  |  | 6.025 | 7.025 |  |  |  | 7.1 | 8.1 |  |  |  | 9.6833 | 10.7625 |  |  |  | 16.0667 | 17.4 |  |  |  | 23.0667 | 24.56
-22 |  |  |  | 6.1 | 7.1 |  |  |  | 7.18 | 8.2 |  |  |  | 9.7429 | 10.9125 |  |  |  | 16.3 | 17.6 |  |  | 8.4 | 23.44 | 24.8
-21.9 |  |  |  | 6.2 | 7.18 |  |  |  | 7.2833 | 8.2833 |  |  |  | 9.95 | 11.0857 |  |  |  | 16.5 | 17.8 |  |  | 8.46 | 23.725 | 25.12
-21.8 |  |  |  | 6.2 | 7.2 |  |  |  | 7.38 | 8.4 |  |  |  | 10.1 | 11.2167 |  |  |  | 16.74 | 18.125 |  |  | 8.64 | 24.075 | 25.4
-21.7 |  |  |  | 6.325 | 7.3333 |  |  |  | 7.46 | 8.4857 |  |  |  | 10.2667 | 11.3833 |  |  |  | 17.0333 | 18.35 |  |  | 8.7333 | 24.35 | 25.8667
-21.6 |  |  |  | 6.4167 | 7.4 |  |  |  | 7.6 | 8.6125 |  |  |  | 10.4 | 11.5333 |  |  |  | 17.2833 | 18.5333 |  |  | 8.8 | 24.65 | 26.16
-21.5 |  |  |  | 6.5 | 7.5 |  |  |  | 7.7125 | 8.75 |  |  |  | 10.6 | 11.7333 |  |  |  | 17.48 | 18.7833 |  |  | 8.9 | 25.05 | 26.4667
-21.4 |  |  |  | 6.6 | 7.6 |  |  |  | 7.8375 | 8.8333 |  |  |  | 10.7 | 11.9167 |  |  |  | 17.78 | 19.05 |  |  | 8.9833 | 25.44 | 26.8
-21.3 |  |  |  | 6.6 | 7.7 |  |  |  | 7.9 | 8.9857 |  |  |  | 10.86 | 12 |  |  |  | 18.06 | 19.34 |  |  | 9.05 | 25.74 | 27.1
-21.2 |  |  |  | 6.7 | 7.7167 |  |  |  | 8.0375 | 9.1 |  |  |  | 11 | 12.1833 |  |  |  | 18.2 | 19.6667 |  |  | 9.1 | 26.075 | 27.5667
-21.1 |  |  |  | 6.85 | 7.8 |  |  |  | 8.1714 | 9.22 |  |  |  | 11.16 | 12.3571 |  |  |  | 18.4667 | 20 |  |  | 9.2 | 26.525 | 27.9
-21 |  |  |  | 6.9 | 7.9 |  |  |  | 8.3 | 9.4 |  |  |  | 11.375 | 12.52 |  |  |  | 18.84 | 20.2 |  |  | 9.3 | 26.88 | 28.22
-20.9 |  |  |  | 7 | 8.0833 |  |  |  | 8.4 | 9.5 |  |  |  | 11.5 | 12.7167 |  |  |  | 19.1667 | 20.5 |  |  | 9.4 | 27.34 | 28.64
-20.8 |  |  |  | 7.1 | 8.1833 |  |  |  | 8.5571 | 9.5833 |  |  |  | 11.7667 | 12.8714 |  |  |  | 19.44 | 20.7333 |  |  | 9.5 | 27.78 | 29
-20.7 |  |  |  | 7.2 | 8.3 |  |  |  | 8.7167 | 9.75 |  |  |  | 11.88 | 13.1286 |  |  |  | 19.68 | 21.08 |  |  | 9.6 | 28.1667 | 29.5
-20.6 |  |  |  | 7.3 | 8.4 |  |  |  | 8.8714 | 9.9 |  |  |  | 12.02 | 13.2833 |  |  |  | 20.0333 | 21.38 |  |  | 9.8 | 28.5571 | 30.05
-20.5 |  |  |  | 7.4667 | 8.5 |  |  |  | 8.94 | 10 |  |  |  | 12.2667 | 13.4667 |  |  |  | 20.42 | 21.68 |  |  | 10 | 29.1 | 30.6
-20.4 |  |  |  | 7.5667 | 8.6333 |  |  |  | 9.0143 | 10.15 |  |  |  | 12.5 | 13.72 |  |  |  | 20.65 | 22.025 |  |  | 10 | 29.7 | 31.24
-20.3 |  |  |  | 7.7 | 8.75 |  |  |  | 9.1667 | 10.25 |  |  |  | 12.7 | 13.9 |  |  | 7.1 | 21.1 | 22.32 |  |  | 10.1333 | 30.6167 | 31.98
-20.2 |  |  |  | 7.8 | 8.85 |  |  |  | 9.3 | 10.4 |  |  |  | 12.88 | 14.08 |  |  | 7.1 | 21.3 | 22.6333 |  |  | 10.2 | 31.7 | 32.9
-20.1 |  |  |  | 7.9 | 8.95 |  |  |  | 9.425 | 10.5833 |  |  |  | 13.1 | 14.2667 |  |  | 7.14 | 21.72 | 23 |  |  | 10.3571 | 33 | 34.2
-20 |  |  |  | 8 | 9.1 |  |  |  | 9.5667 | 10.6667 |  |  |  | 13.325 | 14.4667 |  |  | 7.2 | 22 | 23.3 |  |  | 10.56 | 34.9667 | 35.95
-19.9 |  |  |  | 8.15 | 9.2 |  |  |  | 9.72 | 10.8167 |  |  |  | 13.525 | 14.7167 |  |  | 7.34 | 22.375 | 23.64 |  |  | 10.7 | 37.7167 | 38.9667
-19.8 |  |  |  | 8.28 | 9.35 |  |  |  | 9.9 | 10.9714 |  |  |  | 13.74 | 14.9 |  |  | 7.475 | 22.8 | 23.96 |  |  | 10.86 | 41.82 | 42.6143
-19.7 |  |  |  | 8.3667 | 9.475 |  |  |  | 10.08 | 11.1571 |  |  |  | 14.02 | 15.18 |  |  | 7.6 | 23.15 | 24.45 |  |  | 10.98 | 47.3333 | 48.22
-19.6 |  |  |  | 8.46 | 9.575 |  |  |  | 10.2 | 11.2667 |  |  |  | 14.2 | 15.38 |  |  | 7.6667 | 23.475 | 24.76 |  |  | 11.1 | 54 | 54.4
-19.5 |  |  |  | 8.6333 | 9.7 |  |  |  | 10.35 | 11.5 |  |  |  | 14.44 | 15.66 |  |  | 7.72 | 23.9 | 25.24 |  |  | 11.2 | 61.88 | 61.925
-19.4 |  |  |  | 8.72 | 9.78 |  |  |  | 10.5167 | 11.6429 |  |  |  | 14.64 | 15.95 |  |  | 7.8 | 24.38 | 25.62 |  |  | 11.35 | 69 | 68.95
-19.3 |  |  |  | 8.825 | 9.92 |  |  |  | 10.6833 | 11.7857 |  |  |  | 14.95 | 16.1667 |  |  | 7.9 | 24.8 | 26.16 |  |  | 11.5 | 76.4167 | 76.38
-19.2 |  |  |  | 8.94 | 10.1 |  |  |  | 10.8667 | 11.9429 |  |  |  | 15.275 | 16.4 |  |  | 8.05 | 25.26 | 26.56 |  |  | 11.6 | 82.3333 | 82.1
-19.1 |  |  |  | 9.1 | 10.2 |  |  |  | 11.0667 | 12.1167 |  |  |  | 15.54 | 16.7286 |  |  | 8.1167 | 25.9 | 27.1 |  |  | 11.8 | 87.15 | 86.4
-19 |  |  |  | 9.2167 | 10.3667 |  |  |  | 11.3 | 12.4 |  |  |  | 15.82 | 16.9667 |  |  | 8.25 | 26.45 | 27.68 |  |  | 11.9667 | 90.5167 | 89.6667
-18.9 |  |  |  | 9.3833 | 10.48 |  |  |  | 11.4167 | 12.5333 |  |  |  | 16.12 | 17.3 |  |  | 8.3 | 27.125 | 28.375 |  |  | 12.06 | 92.5 | 91.7667
-18.8 |  |  |  | 9.46 | 10.6 |  |  |  | 11.6 | 12.76 |  |  |  | 16.38 | 17.575 |  |  | 8.4167 | 27.925 | 29.275 |  |  | 12.2 | 93.8 | 93.2
-18.7 |  |  |  | 9.66 | 10.74 |  |  |  | 11.8333 | 12.9333 |  |  |  | 16.6833 | 17.9 |  |  | 8.5 | 29.08 | 30.2 |  |  | 12.3667 | 94.8 | 94.0333
-18.6 |  |  |  | 9.78 | 10.9 |  |  |  | 12.0167 | 13.1 |  |  | 5.4 | 17 | 18.2667 |  |  | 8.6 | 30.48 | 31.84 |  |  | 12.5 | 95.5 | 94.85
-18.5 |  |  |  | 9.9167 | 11.0833 |  |  |  | 12.2667 | 13.3571 |  |  | 5.4 | 17.3667 | 18.5 |  |  | 8.8 | 32.86 | 33.8 |  |  | 12.7 | 96.1 | 95.5
-18.4 |  |  |  | 10.1 | 11.18 |  |  |  | 12.45 | 13.54 |  |  | 5.4143 | 17.7 | 18.8286 |  |  | 8.9 | 36.15 | 37.3 |  |  | 12.85 | 96.6 | 95.75
-18.3 |  |  |  | 10.275 | 11.36 |  |  |  | 12.68 | 13.8 |  |  | 5.5 | 17.98 | 19.15 |  |  | 9 | 41.54 | 42.42 |  |  | 13 | 96.9667 | 96.1667
-18.2 |  |  |  | 10.4 | 11.5 |  |  |  | 12.9333 | 14.0333 |  |  | 5.6 | 18.38 | 19.5429 |  |  | 9.1 | 48.44 | 49.55 |  |  | 13.2 | 97.325 | 96.48
-18.1 |  |  |  | 10.62 | 11.725 |  |  |  | 13.225 | 14.3167 |  |  | 5.68 | 18.75 | 19.9 |  |  | 9.2 | 57 | 57.7 |  |  | 13.375 | 97.62 | 96.75
-18 |  |  |  | 10.74 | 11.92 |  |  |  | 13.4 | 14.6 |  |  | 5.7 | 19.16 | 20.35 |  |  | 9.34 | 66.5 | 66.4333 |  |  | 13.5 | 97.78 | 97.0167
-17.9 |  |  |  | 10.95 | 12.025 |  |  |  | 13.6571 | 14.8 |  |  | 5.8 | 19.56 | 20.7 |  |  | 9.45 | 74.9 | 75.16 |  |  | 13.6333 | 98.0333 | 97.2333
-17.8 |  |  |  | 11.08 | 12.25 |  |  | 4.5 | 13.9667 | 15.0667 |  |  | 5.8 | 20 | 21.2 |  |  | 9.55 | 83.5333 | 83.25 |  |  | 13.9 | 98.2833 | 97.6
-17.7 |  |  |  | 11.28 | 12.3667 |  |  | 4.5 | 14.2 | 15.3 |  |  | 5.94 | 20.55 | 21.5714 |  |  | 9.65 | 90.2 | 89.9 |  |  | 14.16 | 98.45 | 97.8
-17.6 |  |  |  | 11.425 | 12.58 |  |  | 4.6 | 14.52 | 15.6333 |  |  | 6 | 21 | 22.15 |  |  | 9.8429 | 95.34 | 94.8667 |  |  | 14.3 | 98.6167 | 98
-17.5 |  |  |  | 11.64 | 12.75 |  |  | 4.6 | 14.74 | 15.8667 |  |  | 6.1 | 21.6 | 22.7571 |  |  | 10 | 98.46 | 98.0333 |  |  | 14.4333 | 98.9 | 98.15
-17.4 |  |  |  | 11.85 | 13 |  |  | 4.7 | 15.1 | 16.2 |  |  | 6.2 | 22.275 | 23.475 |  |  | 10.1 | 100.5333 | 99.72 |  |  | 14.6 | 99 | 98.3
-17.3 |  |  |  | 12.05 | 13.2 |  |  | 4.8 | 15.3857 | 16.5 |  |  | 6.3 | 23.1333 | 24.3 |  |  | 10.35 | 101.58 | 101.1 |  |  | 14.85 | 99.1 | 98.4
-17.2 |  |  | 4.2 | 12.25 | 13.42 |  |  | 4.8 | 15.7 | 16.86 |  |  | 6.3333 | 24.2 | 25.4333 |  |  | 10.5 | 102.5 | 101.775 |  |  | 15.05 | 99.3 | 98.6
-17.1 |  |  | 4.3 | 12.44 | 13.6 |  |  | 4.86 | 16.1125 | 17.22 |  |  | 6.4167 | 25.95 | 27.1 |  |  | 10.6333 | 103.05 | 102.325 |  |  | 15.3333 | 99.4 | 98.7667
-17 |  |  | 4.4 | 12.675 | 13.8 |  |  | 4.9 | 16.5 | 17.6 |  |  | 6.5 | 28.5 | 29.5333 |  |  | 10.8 | 103.55 | 102.78 |  |  | 15.4857 | 99.5667 | 98.875
-16.9 |  |  | 4.4 | 12.8667 | 14.025 |  |  | 4.9 | 16.9 | 18.05 |  |  | 6.6 | 32.28 | 33.3 |  |  | 10.9667 | 104 | 103.3333 |  |  | 15.7333 | 99.64 | 99
-16.8 |  |  | 4.4333 | 13.06 | 14.3 |  |  | 5 | 17.52 | 18.55 |  |  | 6.7 | 38.8 | 40.05 |  |  | 11.1333 | 104.32 | 103.575 |  |  | 16.0167 | 99.7333 | 99.1833
-16.7 |  |  | 4.5 | 13.38 | 14.5 |  |  | 5.025 | 18.14 | 19.1857 |  |  | 6.8 | 47.95 | 49.3429 |  |  | 11.2 | 104.5 | 103.78 |  |  | 16.2 | 99.9 | 99.36
-16.6 |  |  | 4.5 | 13.5833 | 14.74 |  |  | 5.1 | 18.9 | 19.925 |  |  | 6.8833 | 59.14 | 59.3857 |  |  | 11.5 | 104.6333 | 103.98 |  |  | 16.425 | 100.0667 | 99.45
-16.5 |  |  | 4.6 | 13.8667 | 15.0333 |  |  | 5.2 | 19.94 | 21.06 |  |  | 6.96 | 70.35 | 70.5167 |  |  | 11.7167 | 104.875 | 104.18 |  |  | 16.775 | 100.15 | 99.5167
-16.4 |  |  | 4.6 | 14.2 | 15.35 |  |  | 5.25 | 21.4 | 22.45 |  |  | 7.0667 | 81.82 | 82.05 |  |  | 11.8 | 104.98 | 104.4 |  |  | 16.95 | 100.275 | 99.6
-16.3 |  |  | 4.6 | 14.54 | 15.7 |  |  | 5.3167 | 23.8167 | 24.5 |  |  | 7.18 | 93 | 92.7 |  |  | 12 | 105.1333 | 104.5 |  |  | 17.2167 | 100.4 | 99.68
-16.2 |  |  | 4.7 | 14.975 | 16.24 |  |  | 5.4 | 27.7571 | 28.3 |  |  | 7.2714 | 100.6 | 100.4714 |  |  | 12.2 | 105.28 | 104.6 |  |  | 17.4667 | 100.4 | 99.7667
-16.1 |  |  | 4.8 | 15.55 | 16.7667 |  |  | 5.4714 | 34.0333 | 34.7 |  |  | 7.34 | 106.7 | 106.1571 |  |  | 12.34 | 105.4 | 104.7 |  |  | 17.6714 | 100.5 | 99.8333
-16 |  |  | 4.8 | 16.3333 | 17.5333 |  |  | 5.5 | 43.5 | 43.75 |  |  | 7.5 | 110.1 | 109.35 |  |  | 12.5333 | 105.5 | 104.825 |  |  | 17.9667 | 100.6 | 99.9
-15.9 |  |  | 4.9 | 17.55 | 18.7 |  |  | 5.6 | 55.34 | 55.9286 |  |  | 7.6 | 112.0429 | 111.34 |  |  | 12.8 | 105.6 | 104.9 |  |  | 18.24 | 100.6 | 99.98
-15.8 |  |  | 4.9 | 19.4 | 20.4 |  |  | 5.625 | 68.7286 | 68.75 |  |  | 7.7 | 113.1 | 112.4143 |  |  | 13.02 | 105.7 | 105 |  |  | 18.5 | 100.76 | 100.08
-15.7 |  |  | 5 | 22.5 | 23.7 |  |  | 5.76 | 82.3667 | 81.4 |  |  | 7.8333 | 113.85 | 113.15 |  |  | 13.14 | 105.72 | 105.1 |  |  | 18.7333 | 100.8 | 100.1833
-15.6 |  |  | 5 | 28.22 | 29.325 |  |  | 5.8 | 95.1 | 94.25 |  |  | 8 | 114.35 | 113.5429 |  |  | 13.3333 | 105.8 | 105.14 |  |  | 19.05 | 100.9 | 100.25
-15.5 |  |  | 5.1 | 38 | 39.5 |  |  | 5.9 | 105.2 | 104.66 |  |  | 8.1667 | 114.66 | 113.9286 |  |  | 13.52 | 105.8 | 105.275 |  | 4.9 | 19.35 | 101 | 100.3167
-15.4 |  |  | 5.2 | 50.6 | 52 |  |  | 6 | 113.3 | 112.1333 |  |  | 8.275 | 114.875 | 114.1571 |  |  | 13.82 | 105.8 | 105.3 |  | 4.9 | 19.65 | 101 | 100.4833
-15.3 |  |  | 5.2 | 66.075 | 66.82 |  |  | 6.1 | 117.2571 | 116.5 |  |  | 8.38 | 115 | 114.3 |  |  | 14.0333 | 105.9 | 105.325 |  | 4.9667 | 19.88 | 101.1167 | 100.56
-15.2 |  |  | 5.3 | 81.1 | 80.925 |  |  | 6.2 | 119.6 | 118.6714 |  |  | 8.5167 | 115.1 | 114.3667 |  |  | 14.35 | 105.9 | 105.4 |  | 5 | 20.2 | 101.2 | 100.7
-15.1 |  |  | 5.4 | 95.525 | 95.4 |  |  | 6.3 | 120.75 | 119.9 |  |  | 8.6833 | 115.1 | 114.4 |  |  | 14.5 | 106 | 105.4 |  | 5.1 | 20.625 | 101.3667 | 100.775
-15 |  |  | 5.45 | 107.65 | 107.4 |  |  | 6.3667 | 121.3625 | 120.64 |  |  | 8.8333 | 115.1 | 114.4 |  |  | 14.7 | 106 | 105.4 |  | 5.2 | 20.94 | 101.5 | 100.9
-14.9 |  |  | 5.5 | 117.525 | 116.9 |  |  | 6.44 | 121.76 | 120.9571 |  |  | 9.06 | 115.1 | 114.4 |  |  | 14.975 | 106 | 105.4 |  | 5.2 | 21.24 | 101.5857 | 101
-14.8 |  |  | 5.6 | 122.82 | 122.2 |  |  | 6.5 | 122.0167 | 121.1167 |  |  | 9.1667 | 115.1 | 114.4 |  |  | 15.3 | 106.1 | 105.5 |  | 5.3 | 21.45 | 101.62 | 101.1429
-14.7 |  |  | 5.6 | 125.24 | 124.42 |  |  | 6.64 | 122.1 | 121.2 |  |  | 9.3 | 115.075 | 114.3 |  |  | 15.5333 | 106.1 | 105.5 |  | 5.3333 | 21.85 | 101.8 | 101.2
-14.6 |  |  | 5.7 | 126.4 | 125.65 |  |  | 6.8 | 122.1 | 121.2 |  |  | 9.4667 | 115 | 114.2 |  |  | 15.8 | 106.1 | 105.5 |  | 5.4 | 22.25 | 101.9333 | 101.3
-14.5 |  |  | 5.8333 | 126.875 | 126.1 |  |  | 6.8833 | 122.0167 | 121.2 |  |  | 9.6 | 114.88 | 114.2 |  |  | 16.12 | 106.1 | 105.5 |  | 5.55 | 22.6 | 102 | 101.4
-14.4 |  |  | 5.9 | 127 | 126.2 |  |  | 6.96 | 121.8667 | 121 |  |  | 9.8 | 114.75 | 114.1625 |  |  | 16.44 | 106.1 | 105.5 |  | 5.6 | 23 | 102.1667 | 101.525
-14.3 |  |  | 6 | 126.975 | 126.2 |  |  | 7.1 | 121.7 | 120.8333 |  |  | 9.98 | 114.62 | 114.0167 |  | 3.8 | 16.675 | 106.1 | 105.5 |  | 5.6 | 23.4 | 102.3333 | 101.66
-14.2 |  |  | 6.1 | 126.9 | 126.1 |  |  | 7.2 | 121.5667 | 120.75 |  |  | 10.18 | 114.5 | 113.9 |  | 3.8 | 16.98 | 106.1 | 105.5 |  | 5.7 | 23.8 | 102.5 | 101.8
-14.1 |  |  | 6.2 | 126.625 | 125.875 |  |  | 7.34 | 121.3714 | 120.4833 |  |  | 10.425 | 114.4 | 113.8 |  | 3.9 | 17.35 | 106.1 | 105.5167 |  | 5.7 | 24.4333 | 102.56 | 102
-14 |  |  | 6.3 | 126.4 | 125.5667 |  |  | 7.5 | 121.2167 | 120.2571 |  |  | 10.575 | 114.3 | 113.62 |  | 3.975 | 17.675 | 106.1 | 105.6 |  | 5.8 | 24.9 | 102.6 | 102.1
-13.9 |  |  | 6.4 | 126.05 | 125.22 |  |  | 7.6 | 120.9143 | 120.0857 |  |  | 10.7667 | 114.2 | 113.5333 |  | 4 | 18 | 106.2 | 105.62 |  | 5.875 | 25.6 | 102.7 | 102.22
-13.8 |  |  | 6.475 | 125.6333 | 124.92 |  |  | 7.7167 | 120.725 | 119.92 |  |  | 11.075 | 114.075 | 113.5 |  | 4 | 18.28 | 106.2 | 105.74 |  | 5.925 | 26.4667 | 102.76 | 102.4
-13.7 |  |  | 6.6 | 125.34 | 124.5 |  |  | 7.94 | 120.42 | 119.58 |  |  | 11.3 | 113.925 | 113.3429 |  | 4.1 | 18.675 | 106.2833 | 105.8 |  | 6.1 | 27.4571 | 102.9 | 102.4667
-13.6 |  |  | 6.7 | 124.86 | 124.14 |  |  | 8.05 | 120.1333 | 119.36 |  |  | 11.5 | 113.8 | 113.2 |  | 4.1 | 19.1 | 106.4 | 105.9 |  | 6.175 | 28.7 | 103 | 102.5286
-13.5 |  |  | 6.82 | 124.5 | 123.7 |  |  | 8.2 | 119.7667 | 119.05 |  |  | 11.7667 | 113.7 | 113.0333 |  | 4.2 | 19.35 | 106.425 | 105.96 |  | 6.3 | 30.5 | 103.1 | 102.7
-13.4 |  |  | 6.92 | 124.1 | 123.35 |  |  | 8.32 | 119.5375 | 118.8 |  |  | 12 | 113.475 | 112.8833 |  | 4.2 | 19.8 | 106.5 | 106.05 |  | 6.3 | 33.4333 | 103.2 | 102.7833
-13.3 |  |  | 7.04 | 123.7 | 122.94 |  |  | 8.5 | 119.275 | 118.66 |  |  | 12.28 | 113.3 | 112.7375 |  | 4.3 | 20.1667 | 106.56 | 106.1 |  | 6.4 | 37.35 | 103.25 | 102.84
-13.2 |  |  | 7.1167 | 123.3 | 122.52 |  |  | 8.65 | 119.05 | 118.3286 |  |  | 12.52 | 113.175 | 112.5667 |  | 4.4 | 20.575 | 106.6 | 106.2 |  | 6.5 | 43.48 | 103.3167 | 103
-13.1 |  |  | 7.3 | 122.75 | 122.075 |  |  | 8.9 | 118.7 | 117.9286 |  | 2.6 | 12.8 | 113.02 | 112.4875 |  | 4.4 | 21.16 | 106.8 | 106.2667 |  | 6.6 | 49.9333 | 103.4833 | 103
-13 |  |  | 7.3833 | 122.5 | 121.8 |  |  | 9.05 | 118.3286 | 117.66 |  | 2.6667 | 13.0667 | 113 | 112.3625 |  | 4.5 | 21.72 | 106.8667 | 106.3 |  | 6.7 | 57.58 | 103.58 | 103.1
-12.9 |  |  | 7.5 | 122 | 121.4 |  |  | 9.3 | 118.0667 | 117.36 |  | 2.7 | 13.36 | 112.9 | 112.3 |  | 4.6 | 22.45 | 106.9 | 106.4 |  | 6.775 | 64.8 | 103.7 | 103.2
-12.8 |  |  | 7.7 | 121.65 | 120.95 |  |  | 9.525 | 117.8167 | 117.16 |  | 2.7 | 13.725 | 112.84 | 112.2 |  | 4.6 | 23.2 | 107.0167 | 106.5 |  | 6.9 | 72.8 | 103.8 | 103.3
-12.7 |  |  | 7.8 | 121.18 | 120.6 |  |  | 9.7 | 117.48 | 116.925 |  | 2.7 | 14.04 | 112.8 | 112.2 |  | 4.7 | 24.2333 | 107.1 | 106.5 |  | 7 | 78.7 | 103.8 | 103.3333
-12.6 |  |  | 7.98 | 120.9 | 120.1667 |  |  | 9.88 | 117.3 | 116.5667 |  | 2.8 | 14.32 | 112.7 | 112.1 |  | 4.7 | 25.6 | 107.1 | 106.6 |  | 7.0333 | 83.28 | 103.8 | 103.4
-12.5 |  |  | 8.1333 | 120.36 | 119.74 |  | 2.1 | 10.125 | 116.9429 | 116.26 |  | 2.8 | 14.65 | 112.62 | 112.1 |  | 4.8 | 27.76 | 107.15 | 106.6 |  | 7.1 | 86.8571 | 103.9 | 103.4
-12.4 |  |  | 8.26 | 120 | 119.5 |  | 2.1 | 10.36 | 116.7167 | 116.1 |  | 2.825 | 15 | 112.6 | 112.1 |  | 4.9 | 30.76 | 107.2 | 106.7 |  | 7.3 | 89.1 | 104 | 103.58
-12.3 |  |  | 8.475 | 119.65 | 119.1 |  | 2.2 | 10.6 | 116.45 | 115.92 |  | 2.9 | 15.45 | 112.6 | 112.0714 |  | 4.95 | 35.45 | 107.2 | 106.7 |  | 7.4 | 90.58 | 104.0167 | 103.6
-12.2 |  |  | 8.68 | 119.2 | 118.6 |  | 2.2 | 10.8 | 116.28 | 115.775 |  | 2.9667 | 15.9 | 112.56 | 112 |  | 5 | 42.46 | 107.2 | 106.76 |  | 7.45 | 91.6 | 104.1 | 103.7
-12.1 |  | 2 | 8.875 | 118.82 | 118.28 |  | 2.2 | 11.12 | 116.2 | 115.65 |  | 3 | 16.3 | 112.5 | 112 |  | 5.1 | 50.15 | 107.2167 | 106.8 |  | 7.5 | 92.4333 | 104.2429 | 103.75
-12 |  | 2 | 8.98 | 118.5 | 118.1 |  | 2.2 | 11.425 | 116.05 | 115.5 |  | 3 | 16.9333 | 112.5 | 112 |  | 5.2 | 59.275 | 107.3 | 106.85 |  | 7.6 | 92.95 | 104.3 | 103.8
-11.9 |  | 2 | 9.175 | 118.3 | 117.75 |  | 2.3 | 11.75 | 116 | 115.4 |  | 3.1 | 17.58 | 112.4 | 112 |  | 5.28 | 67.5667 | 107.32 | 106.9 |  | 7.8 | 93.7 | 104.3667 | 103.8667
-11.8 |  | 2.1 | 9.3429 | 118.02 | 117.5 |  | 2.3 | 12.125 | 115.9 | 115.3 |  | 3.1 | 18.4 | 112.4 | 112 |  | 5.3 | 75.95 | 107.4 | 106.9 |  | 7.9 | 94.1 | 104.4333 | 103.9
-11.7 |  | 2.1 | 9.575 | 117.8 | 117.3 |  | 2.4 | 12.4 | 115.8 | 115.2375 |  | 3.1667 | 19.5 | 112.4 | 112 |  | 5.4 | 82.6667 | 107.4 | 107 |  | 8.0333 | 94.46 | 104.5 | 104
-11.6 |  | 2.1 | 9.76 | 117.6 | 117.2 |  | 2.4 | 12.9 | 115.7 | 115.1 |  | 3.2 | 20.925 | 112.4 | 112 |  | 5.5 | 87.9 | 107.4 | 107.06 | 1.875 | 8.2 | 94.8 | 104.5571 | 104.05
-11.5 |  | 2.1 | 10.06 | 117.5 | 117.06 |  | 2.4 | 13.3667 | 115.6 | 115 |  | 3.3 | 23.2 | 112.3 | 112 |  | 5.6 | 91.6 | 107.4 | 107.1 | 1.875 | 8.3 | 95.14 | 104.6 | 104.1
-11.4 |  | 2.1 | 10.2667 | 117.3 | 116.9 |  | 2.4 | 14 | 115.5143 | 114.9 |  | 3.3 | 26.85 | 112.3 | 111.9167 |  | 5.7 | 93.7 | 107.5 | 107.1 | 1.9 | 8.425 | 95.54 | 104.6 | 104.2
-11.3 |  | 2.2 | 10.55 | 117.2 | 116.72 |  | 2.4 | 14.7 | 115.35 | 114.9 |  | 3.4 | 32.9 | 112.3 | 111.9 |  | 5.8 | 95.08 | 107.5 | 107.1 | 2 | 8.575 | 95.78 | 104.7 | 104.3
-11.2 |  | 2.2 | 10.88 | 117.1 | 116.62 |  | 2.5 | 15.9 | 115.2667 | 114.8 |  | 3.4 | 41.6 | 112.3 | 111.88 |  | 5.9 | 96.06 | 107.5 | 107.1 | 2 | 8.7 | 96.1 | 104.7 | 104.3
-11.1 |  | 2.2 | 11.35 | 116.9167 | 116.5167 |  | 2.5 | 17.3 | 115.1 | 114.7375 |  | 3.5 | 51 | 112.2 | 111.7833 |  | 6 | 96.76 | 107.6 | 107.1167 | 2 | 8.9 | 96.5 | 104.7 | 104.4
-11 |  | 2.2 | 11.925 | 116.84 | 116.32 |  | 2.5 | 19.82 | 115 | 114.66 |  | 3.5 | 61.38 | 112.2 | 111.7 |  | 6.1 | 97.3667 | 107.7 | 107.2 | 2 | 9 | 96.7667 | 104.8 | 104.475
-10.9 |  | 2.3 | 12.7 | 116.7 | 116.2 |  | 2.6 | 23.8 | 114.94 | 114.5833 |  | 3.6 | 72.05 | 112.1 | 111.6 |  | 6.2 | 97.95 | 107.7 | 107.2 | 2 | 9.15 | 97.06 | 104.8 | 104.5
-10.8 |  | 2.3 | 13.86 | 116.525 | 116.1 |  | 2.6 | 30.96 | 114.88 | 114.4 |  | 3.6667 | 81.875 | 112 | 111.6 |  | 6.4 | 98.4 | 107.7 | 107.3 | 2.1 | 9.3 | 97.34 | 104.9 | 104.5
-10.7 |  | 2.3 | 15.7 | 116.48 | 116 |  | 2.6 | 40.25 | 114.8 | 114.3 |  | 3.7 | 90.175 | 111.9 | 111.6 | 1.5 | 6.5 | 98.775 | 107.7 | 107.3 | 2.1 | 9.5 | 97.6333 | 104.9 | 104.5
-10.6 |  | 2.3333 | 18.8333 | 116.4 | 115.9 |  | 2.7 | 50.95 | 114.6333 | 114.175 |  | 3.8 | 95.8 | 111.9 | 111.5 | 1.5 | 6.6 | 99.1 | 107.7 | 107.3 | 2.1 | 9.65 | 97.9 | 105 | 104.5
-10.5 |  | 2.4 | 24.66 | 116.2 | 115.7 |  | 2.7 | 63.6 | 114.5667 | 114.0333 |  | 3.825 | 98.7 | 111.9 | 111.4625 | 1.5 | 6.7 | 99.4429 | 107.7 | 107.3 | 2.2 | 9.8 | 97.94 | 105 | 104.6
-10.4 |  | 2.4 | 34.35 | 116.1 | 115.6 |  | 2.8 | 75.2 | 114.38 | 113.9714 |  | 3.9 | 100.85 | 111.8 | 111.4 | 1.5 | 6.9 | 99.8 | 107.625 | 107.3 | 2.2 | 10 | 98.15 | 105 | 104.64
-10.3 |  | 2.4 | 45.7167 | 116 | 115.45 |  | 2.8 | 85.8 | 114.3 | 113.8167 |  | 4 | 102 | 111.8 | 111.3 | 1.6 | 7 | 100.04 | 107.7 | 107.3 | 2.2 | 10.2 | 98.5167 | 105.1 | 104.7
-10.2 |  | 2.45 | 58.7 | 115.775 | 115.32 |  | 2.8 | 95 | 114.2 | 113.7625 |  | 4.0667 | 103.04 | 111.7 | 111.2143 | 1.6 | 7.0333 | 100.3 | 107.6 | 107.3 | 2.2 | 10.3 | 98.7 | 105.1 | 104.7
-10.1 |  | 2.5 | 72.5 | 115.6 | 115.25 |  | 2.9 | 100.85 | 114.1 | 113.6 |  | 4.125 | 103.68 | 111.6 | 111.1833 | 1.6 | 7.1667 | 100.5667 | 107.6 | 107.2 | 2.2 | 10.525 | 99 | 105.1 | 104.7
-10 |  | 2.5 | 83.95 | 115.5 | 115.08 |  | 2.9 | 103.9 | 114 | 113.52 |  | 4.2 | 104.15 | 111.6 | 111.1 | 1.7 | 7.3 | 100.84 | 107.6 | 107.3 | 2.2 | 10.7 | 99.15 | 105.2 | 104.7
-9.9 |  | 2.5667 | 95.525 | 115.3 | 114.9 |  | 3 | 105.925 | 113.9167 | 113.475 |  | 4.3 | 104.6 | 111.5 | 111.0833 | 1.7 | 7.425 | 101.1333 | 107.6 | 107.3 | 2.3 | 10.825 | 99.28 | 105.2 | 104.7
-9.8 |  | 2.6 | 102.8 | 115.125 | 114.775 |  | 3 | 107.2 | 113.8 | 113.3429 | 0.925 | 4.4 | 104.95 | 111.4 | 111 | 1.7 | 7.675 | 101.3 | 107.6 | 107.3 | 2.325 | 11.1 | 99.525 | 105.2 | 104.7
-9.7 |  | 2.6 | 107.28 | 115 | 114.65 |  | 3.1 | 107.85 | 113.7 | 113.2667 | 1 | 4.5 | 105.28 | 111.4 | 110.94 | 1.7 | 7.8 | 101.5 | 107.6 | 107.3 | 2.4 | 11.2 | 99.725 | 105.2 | 104.8
-9.6 |  | 2.7 | 109.52 | 114.975 | 114.5 |  | 3.1 | 108.5 | 113.6143 | 113.2 | 1 | 4.6 | 105.65 | 111.3 | 110.85 | 1.7 | 8 | 101.6833 | 107.6 | 107.3 | 2.5 | 11.4 | 99.92 | 105.2 | 104.8
-9.5 |  | 2.7 | 110.76 | 114.84 | 114.45 |  | 3.2 | 108.9 | 113.5 | 113.0571 | 1 | 4.6667 | 105.8 | 111.2 | 110.78 | 1.7 | 8.1 | 101.875 | 107.6 | 107.3 | 2.5 | 11.6 | 100.1 | 105.2 | 104.8
-9.4 |  | 2.8 | 111.5 | 114.66 | 114.3 | 0.73333 | 3.275 | 109.22 | 113.3333 | 112.9 | 1 | 4.8 | 106 | 111.1 | 110.7 | 1.7333 | 8.3 | 102.0333 | 107.6 | 107.3 | 2.6 | 12 | 100.24 | 105.2 | 104.8
-9.3 |  | 2.8 | 111.86 | 114.54 | 114.175 | 0.8 | 3.3 | 109.475 | 113.2571 | 112.8 | 1 | 4.9 | 106.14 | 111.1 | 110.7 | 1.725 | 8.5 | 102.2833 | 107.6 | 107.3 | 2.6 | 12.1667 | 100.35 | 105.2 | 104.82
-9.2 |  | 2.9 | 112.14 | 114.4 | 114 | 0.8 | 3.4 | 109.6 | 113.125 | 112.7167 | 1.0333 | 5 | 106.4 | 111 | 110.6 | 1.8 | 8.7 | 102.42 | 107.5 | 107.26 | 2.6 | 12.3333 | 100.5 | 105.2 | 104.85
-9.1 | 0.7 | 2.9 | 112.36 | 114.34 | 113.9 | 0.8 | 3.5 | 109.8 | 113.0857 | 112.66 | 1.1 | 5.1 | 106.5 | 110.9167 | 110.5 | 1.8 | 8.9 | 102.5 | 107.5 | 107.2 | 2.7 | 12.6 | 100.625 | 105.275 | 104.8
-9 | 0.7 | 3 | 112.4 | 114.2 | 113.8 | 0.8 | 3.5 | 109.9 | 112.9333 | 112.56 | 1.1 | 5.3 | 106.66 | 110.8 | 110.5 | 1.8 | 9 | 102.6 | 107.5 | 107.2 | 2.7 | 12.925 | 100.78 | 105.2 | 104.9
-8.9 | 0.7 | 3 | 112.5 | 114.1 | 113.65 | 0.8 | 3.6 | 110 | 112.8875 | 112.3667 | 1.1 | 5.4 | 106.775 | 110.78 | 110.4 | 1.8 | 9.25 | 102.88 | 107.5 | 107.2 | 2.7 | 13.175 | 100.9 | 105.3 | 104.8333
-8.8 | 0.7 | 3.1 | 112.6 | 113.9 | 113.5 | 0.8 | 3.7 | 110.1 | 112.725 | 112.3 | 1.1 | 5.45 | 106.9 | 110.7 | 110.4 | 1.9 | 9.4333 | 103 | 107.4 | 107.2 | 2.7 | 13.3667 | 101 | 105.3 | 104.9
-8.7 | 0.7 | 3.1 | 112.6 | 113.75 | 113.3 | 0.8 | 3.8 | 110.2 | 112.5167 | 112.1667 | 1.1 | 5.625 | 107 | 110.6 | 110.4 | 1.9 | 9.6667 | 103.1 | 107.4 | 107.1 | 2.8 | 13.52 | 101.15 | 105.3 | 104.9
-8.6 | 0.8 | 3.2 | 112.6 | 113.625 | 113.225 | 0.8 | 3.9 | 110.2 | 112.38 | 112.0143 | 1.1 | 5.8 | 107.1 | 110.56 | 110.3 | 1.9 | 9.8667 | 103.2 | 107.4 | 107.1 | 2.8 | 13.8 | 101.28 | 105.3 | 104.9
-8.5 | 0.8 | 3.3 | 112.6 | 113.5 | 113 | 0.9 | 4 | 110.3 | 112.3 | 111.9 | 1.1 | 5.9 | 107.1 | 110.44 | 110.2857 | 1.925 | 10.2 | 103.3 | 107.3 | 107.1 | 2.8 | 14.2 | 101.4 | 105.3 | 104.9
-8.4 | 0.8 | 3.3 | 112.5 | 113.4 | 112.9167 | 0.9 | 4.0333 | 110.3 | 112.2 | 111.8 | 1.1 | 6.1 | 107.22 | 110.4 | 110.1429 | 2 | 10.475 | 103.35 | 107.3 | 107.1 | 2.9 | 14.5 | 101.5 | 105.25 | 105
-8.3 | 0.8 | 3.4 | 112.4 | 113.3 | 112.8 | 0.9 | 4.2 | 110.3 | 112.1 | 111.7571 | 1.2 | 6.2 | 107.3167 | 110.3 | 109.9833 | 2 | 10.675 | 103.4333 | 107.3 | 107.075 | 2.9 | 14.7 | 101.7 | 105.2 | 104.9
-8.2 | 0.8 | 3.45 | 112.4 | 113.1 | 112.7 | 0.9 | 4.3 | 110.3 | 112 | 111.7 | 1.2 | 6.4 | 107.4 | 110.2 | 109.9 | 2.05 | 10.95 | 103.5333 | 107.2 | 107 | 3 | 15 | 101.7571 | 105.2 | 105
-8.1 | 0.8 | 3.5 | 112.3 | 112.9667 | 112.6 | 0.9 | 4.325 | 110.3 | 111.8833 | 111.5167 | 1.2 | 6.6 | 107.4 | 110.2 | 109.8 | 2.1 | 11.275 | 103.6 | 107.2 | 107 | 3 | 15.4 | 101.8167 | 105.1 | 105
-8 | 0.8 | 3.6 | 112.24 | 112.8 | 112.4 | 0.9 | 4.5 | 110.275 | 111.7286 | 111.4571 | 1.275 | 6.8 | 107.5 | 110.1 | 109.74 | 2.1 | 11.575 | 103.7167 | 107.2 | 106.875 | 3.1 | 15.72 | 101.9 | 105.1333 | 105
-7.9 | 0.8 | 3.7 | 112.2 | 112.7 | 112.3 | 0.9 | 4.6 | 110.2 | 111.6714 | 111.3333 | 1.3 | 7 | 107.5 | 110 | 109.6 | 2.2 | 11.76 | 103.8 | 107.2 | 106.8 | 3.2 | 16.1 | 102.05 | 105.2 | 104.9
-7.8 | 0.8 | 3.8 | 112.0833 | 112.575 | 112.2 | 0.9 | 4.7667 | 110.2 | 111.5167 | 111.1667 | 1.3 | 7.2 | 107.5 | 109.9 | 109.5 | 2.2 | 12.1333 | 103.9 | 107.1 | 106.8 | 3.2 | 16.5 | 102.1333 | 105.2 | 104.9
-7.7 | 0.8 | 3.9 | 111.98 | 112.4 | 112.08 | 0.9 | 4.9 | 110.2 | 111.4 | 111 | 1.3 | 7.4 | 107.5 | 109.84 | 109.4 | 2.2 | 12.45 | 103.9833 | 107 | 106.8 | 3.2 | 16.95 | 102.3 | 105.1 | 104.85
-7.6 | 0.8 | 4 | 111.9 | 112.3 | 111.9 | 0.9 | 5.1 | 110.1 | 111.2 | 110.9 | 1.3 | 7.625 | 107.525 | 109.72 | 109.3143 | 2.2 | 12.8 | 104.1 | 107 | 106.8 | 3.3 | 17.45 | 102.4 | 105.1 | 104.9
-7.5 | 0.9 | 4.1 | 111.8 | 112.1 | 111.8 | 0.9 | 5.2 | 110.08 | 111.1 | 110.8 | 1.3 | 7.825 | 107.55 | 109.6167 | 109.1571 | 2.3 | 13.05 | 104.15 | 107 | 106.775 | 3.3667 | 18.1333 | 102.4 | 105.1 | 104.8
-7.4 | 0.9 | 4.2 | 111.7 | 111.95 | 111.6 | 1 | 5.4 | 110 | 111 | 110.6333 | 1.4 | 8.0667 | 107.55 | 109.5 | 109.1 | 2.4 | 13.325 | 104.2 | 106.9 | 106.7 | 3.4 | 18.7 | 102.4 | 105 | 104.9
-7.3 | 0.9 | 4.275 | 111.68 | 111.82 | 111.5 | 1 | 5.5 | 110 | 110.9 | 110.54 | 1.4 | 8.325 | 107.6 | 109.4 | 109.0375 | 2.4 | 13.75 | 104.2 | 106.9 | 106.7 | 3.5 | 19.3 | 102.5 | 105 | 104.9
-7.2 | 0.9 | 4.4 | 111.5 | 111.7 | 111.3833 | 1 | 5.7 | 109.9 | 110.75 | 110.4667 | 1.4 | 8.625 | 107.6 | 109.3 | 109 | 2.425 | 14.15 | 104.3 | 106.84 | 106.6 | 3.42 | 20.6333 | 102.6 | 105 | 104.8
-7.1 | 0.9 | 4.5 | 111.4 | 111.5 | 111.2 | 1 | 5.8333 | 109.8 | 110.66 | 110.4 | 1.4 | 8.9 | 107.6 | 109.2 | 108.9167 | 2.5 | 14.6667 | 104.36 | 106.8 | 106.5167 | 3.44 | 21.9 | 102.6 | 105 | 104.8
-7 | 0.9 | 4.6 | 111.225 | 111.325 | 111.1 | 1 | 6.0667 | 109.7 | 110.5 | 110.22 | 1.4 | 9.1667 | 107.6 | 109.1 | 108.8625 | 2.6 | 15.15 | 104.4 | 106.7 | 106.5 | 3.5 | 23.9333 | 102.6 | 105 | 104.7
-6.9 | 0.9 | 4.8 | 111.1 | 111.24 | 111 | 1 | 6.3 | 109.65 | 110.4333 | 110.0333 | 1.5 | 9.5 | 107.5 | 109 | 108.7286 | 2.6 | 15.775 | 104.4 | 106.7 | 106.4 | 3.6 | 26.8 | 102.74 | 105 | 104.7
-6.8 | 0.9 | 4.9 | 111.02 | 111.1 | 110.8667 | 1 | 6.5333 | 109.5 | 110.3714 | 109.9 | 1.575 | 9.8 | 107.5 | 108.9 | 108.7 | 2.7 | 16.5 | 104.4 | 106.7 | 106.35 | 3.65 | 31 | 102.8 | 105 | 104.7
-6.7 | 0.9 | 5.1 | 110.8 | 110.9 | 110.68 | 1.1 | 6.8 | 109.5 | 110.2286 | 109.88 | 1.6 | 10.25 | 107.4 | 108.8833 | 108.5571 | 2.7 | 17.55 | 104.4 | 106.6 | 106.3 | 3.8 | 37.9 | 102.8 | 104.9 | 104.7
-6.6 | 0.9 | 5.2667 | 110.8 | 110.7667 | 110.5667 | 1.1 | 7.05 | 109.4 | 110.0667 | 109.7 | 1.6 | 10.7 | 107.4 | 108.8 | 108.5 | 2.8 | 19 | 104.5 | 106.6 | 106.3 | 4 | 45.125 | 102.9 | 104.9 | 104.6
-6.5 | 0.9 | 5.4 | 110.68 | 110.66 | 110.4 | 1.1 | 7.34 | 109.3 | 109.9 | 109.6 | 1.6 | 11.2333 | 107.4 | 108.64 | 108.3857 | 2.9 | 21.1 | 104.5 | 106.5 | 106.3 | 4.0333 | 54.3 | 102.95 | 104.9 | 104.6
-6.4 | 0.9 | 5.6 | 110.6 | 110.4667 | 110.3 | 1.1 | 7.7667 | 109.2 | 109.7667 | 109.4667 | 1.625 | 11.8 | 107.4 | 108.55 | 108.3 | 3 | 24.2 | 104.525 | 106.4 | 106.2 | 4.2 | 62.85 | 103 | 104.8 | 104.5667
-6.3 | 0.9 | 5.8 | 110.42 | 110.3 | 110.15 | 1.2 | 8.0667 | 109.175 | 109.6143 | 109.3 | 1.7 | 12.525 | 107.3 | 108.4 | 108.15 | 3 | 29.275 | 104.6 | 106.3167 | 106.1 | 4.3 | 71.375 | 103 | 104.8 | 104.5
-6.2 | 1 | 6.0667 | 110.3 | 110.25 | 110 | 1.2 | 8.6 | 109.06 | 109.4714 | 109.2333 | 1.7 | 13.5667 | 107.22 | 108.3 | 108.0833 | 3.125 | 36.2333 | 104.6 | 106.25 | 106.04 | 4.3 | 79.275 | 103 | 104.8 | 104.5
-6.1 | 1 | 6.4 | 110.18 | 110.1 | 109.8 | 1.2 | 9.2 | 108.9 | 109.4 | 109.0714 | 1.8 | 15 | 107.2 | 108.2 | 107.9167 | 3.2 | 44.4 | 104.5 | 106.1833 | 106 | 4.4 | 85.3 | 103 | 104.8 | 104.5
-6 | 1 | 6.75 | 110 | 109.9667 | 109.7 | 1.3 | 10.025 | 108.8 | 109.2625 | 109 | 1.8333 | 17.0333 | 107.1 | 108.1 | 107.85 | 3.3 | 54.2667 | 104.5 | 106.0667 | 105.94 | 4.55 | 89.2 | 103 | 104.7 | 104.4
-5.9 | 1 | 7.3333 | 109.86 | 109.8 | 109.5 | 1.3 | 11.2 | 108.7 | 109.1 | 108.9 | 1.9 | 20.6333 | 107.1 | 107.98 | 107.7 | 3.4 | 63.9 | 104.5 | 106 | 105.9 | 4.7 | 91.6 | 103.02 | 104.65 | 104.4
-5.8 | 1.05 | 8.1 | 109.8 | 109.7 | 109.4 | 1.3 | 12.925 | 108.6 | 109 | 108.74 | 2 | 26.2333 | 107 | 107.85 | 107.6286 | 3.4 | 73.3 | 104.5 | 105.9 | 105.8 | 4.725 | 93.5 | 103 | 104.58 | 104.3
-5.7 | 1.1 | 9.2 | 109.675 | 109.6 | 109.3 | 1.3 | 15.7 | 108.5 | 108.84 | 108.5833 | 2 | 34.4 | 107 | 107.7 | 107.5 | 3.55 | 81.5333 | 104.5 | 105.8 | 105.7 | 4.875 | 94.5667 | 103 | 104.5 | 104.3
-5.6 | 1.1 | 11.05 | 109.5 | 109.4 | 109.15 | 1.375 | 20.5 | 108.4 | 108.7 | 108.4 | 2.1 | 44.3 | 106.9 | 107.56 | 107.45 | 3.6 | 87.5 | 104.4 | 105.8 | 105.6 | 4.925 | 95.5 | 103 | 104.4 | 104.2
-5.5 | 1.1 | 14.0333 | 109.4 | 109.2 | 109 | 1.4 | 27.9 | 108.28 | 108.5571 | 108.25 | 2.1 | 54.9333 | 106.8 | 107.45 | 107.3143 | 3.7 | 91.3333 | 104.4 | 105.7 | 105.5667 | 5 | 96 | 102.9 | 104.3167 | 104.1
-5.4 | 1.1 | 20.0667 | 109.1667 | 109.05 | 108.875 | 1.475 | 38 | 108.14 | 108.4143 | 108.1143 | 2.2 | 66.4 | 106.7 | 107.3333 | 107.1429 | 3.8 | 93.9 | 104.34 | 105.62 | 105.5 | 5.125 | 96.6 | 102.9 | 104.3 | 104.1
-5.3 | 1.1 | 28.8667 | 109 | 108.9 | 108.74 | 1.5 | 49.8 | 108 | 108.3 | 107.9625 | 2.2333 | 76.7 | 106.6 | 107.2333 | 107.0857 | 3.9 | 95.3 | 104.3 | 105.6 | 105.4 | 5.275 | 97.075 | 102.9 | 104.2 | 104
-5.2 | 1.2 | 41.125 | 108.9 | 108.7 | 108.6 | 1.5 | 62.2 | 107.88 | 108.225 | 107.8286 | 2.3 | 84.7667 | 106.6 | 107.1 | 106.95 | 4 | 96.3 | 104.2 | 105.5 | 105.3 | 5.5 | 97.375 | 102.9 | 104.1 | 104
-5.1 | 1.2 | 54.05 | 108.8 | 108.5333 | 108.375 | 1.6 | 73.6 | 107.76 | 108.08 | 107.6857 | 2.4 | 91.325 | 106.4 | 106.95 | 106.8143 | 4.2 | 97.1333 | 104.2 | 105.3 | 105.1667 | 5.6 | 97.7 | 102.9 | 104.0333 | 103.9
-5 | 1.2 | 67.0333 | 108.65 | 108.4 | 108.2333 | 1.6 | 83.9 | 107.625 | 107.9143 | 107.6 | 2.5 | 94.8 | 106.3167 | 106.8 | 106.66 | 4.2333 | 97.65 | 104.2 | 105.2167 | 105.1 | 5.775 | 98 | 102.9 | 104 | 103.8
-4.9 | 1.3 | 78.9 | 108.45 | 108.275 | 108.1 | 1.65 | 91.7 | 107.55 | 107.76 | 107.3875 | 2.6 | 97.225 | 106.18 | 106.7 | 106.6 | 4.4 | 98.2 | 104.2 | 105.1 | 105 | 5.825 | 98.2 | 102.9 | 103.9 | 103.8
-4.8 | 1.3 | 89.05 | 108.2833 | 108.05 | 107.96 | 1.7 | 96.4667 | 107.38 | 107.4857 | 107.2571 | 2.7 | 98.3333 | 106.075 | 106.5833 | 106.4875 | 4.55 | 98.5 | 104.1 | 105.08 | 104.88 | 5.9667 | 98.4667 | 102.8 | 103.9 | 103.7
-4.7 | 1.3 | 96.5 | 108.1286 | 107.95 | 107.65 | 1.8 | 99.1 | 107.28 | 107.3429 | 107 | 2.8 | 99.3 | 105.98 | 106.4 | 106.3625 | 4.7 | 98.8 | 104 | 104.975 | 104.8 | 6.1 | 98.65 | 102.7167 | 103.8 | 103.7
-4.6 | 1.4 | 100.7 | 107.92 | 107.7 | 107.5 | 1.8 | 100.6667 | 107.08 | 107.2 | 106.8833 | 2.9 | 100.0333 | 105.875 | 106.26 | 106.1 | 4.8 | 99 | 103.9 | 104.88 | 104.7 | 6.3 | 98.9 | 102.7 | 103.7333 | 103.6
-4.5 | 1.4 | 102.7 | 107.675 | 107.54 | 107.375 | 1.9 | 101.675 | 106.9 | 106.9857 | 106.8 | 3 | 100.6333 | 105.7 | 106.1 | 105.9333 | 4.9 | 99.2 | 103.85 | 104.78 | 104.6 | 6.425 | 99.1 | 102.7 | 103.7 | 103.5
-4.4 | 1.4667 | 103.9 | 107.525 | 107.325 | 107.14 | 2 | 102.4 | 106.6 | 106.7833 | 106.54 | 3.075 | 100.8667 | 105.6 | 105.9833 | 105.78 | 5.05 | 99.4667 | 103.8 | 104.6 | 104.42 | 6.6 | 99.2 | 102.7 | 103.6 | 103.475
-4.3 | 1.5 | 104.4333 | 107.32 | 107.1 | 106.95 | 2.1 | 102.8 | 106.5 | 106.6 | 106.36 | 3.2 | 101.2 | 105.5 | 105.8167 | 105.6 | 5.3 | 99.6 | 103.725 | 104.5 | 104.4 | 6.825 | 99.3333 | 102.62 | 103.5 | 103.4
-4.2 | 1.575 | 104.7 | 107.16 | 107 | 106.825 | 2.2 | 103.0333 | 106.34 | 106.5 | 106.1667 | 3.325 | 101.3 | 105.3333 | 105.64 | 105.5 | 5.5 | 99.85 | 103.7 | 104.4 | 104.3 | 7 | 99.58 | 102.6 | 103.4167 | 103.3
-4.1 | 1.6 | 104.9 | 106.94 | 106.8333 | 106.6 | 2.225 | 103.25 | 106.2 | 106.3333 | 105.9833 | 3.5 | 101.4333 | 105.2 | 105.44 | 105.42 | 5.6 | 100 | 103.7 | 104.3 | 104.2 | 7.1 | 99.75 | 102.6 | 103.4 | 103.2
-4 | 1.66 | 105 | 106.72 | 106.55 | 106.3667 | 2.3667 | 103.4 | 106.04 | 106.1333 | 105.9 | 3.6667 | 101.575 | 105.1 | 105.3 | 105.2143 | 5.85 | 100.1 | 103.6 | 104.2 | 104.1 | 7.4 | 99.9 | 102.5 | 103.38 | 103.2
-3.9 | 1.7 | 105 | 106.52 | 106.28 | 106.18 | 2.4667 | 103.4 | 105.825 | 105.875 | 105.64 | 3.8 | 101.74 | 104.98 | 105.2 | 105.025 | 6.1 | 100.2 | 103.475 | 104.1 | 104.025 | 7.7 | 100 | 102.4 | 103.3 | 103.2
-3.8 | 1.8 | 105 | 106.35 | 106.1 | 105.98 | 2.6 | 103.5 | 105.58 | 105.65 | 105.5 | 4 | 101.9 | 104.8 | 105 | 104.875 | 6.2 | 100.3333 | 103.4 | 104 | 103.98 | 8 | 100.1 | 102.45 | 103.2833 | 103.1
-3.7 | 1.9 | 105 | 106.15 | 105.92 | 105.875 | 2.7 | 103.5 | 105.42 | 105.4833 | 105.3333 | 4.2 | 101.9333 | 104.7 | 104.9333 | 104.72 | 6.45 | 100.55 | 103.3 | 103.9 | 103.88 | 8.4 | 100.2 | 102.5 | 103.18 | 103
-3.6 | 2 | 104.9 | 105.95 | 105.75 | 105.625 | 2.8667 | 103.58 | 105.275 | 105.3286 | 105.1857 | 4.4 | 102.05 | 104.6 | 104.8 | 104.6 | 6.7667 | 100.7 | 103.3 | 103.9 | 103.8 | 8.9 | 100.4 | 102.5 | 103 | 103
-3.5 | 2 | 104.8 | 105.7 | 105.65 | 105.375 | 3.0333 | 103.6 | 105.06 | 105.15 | 104.9667 | 4.6667 | 102.2 | 104.5 | 104.65 | 104.55 | 7.0667 | 100.8 | 103.3 | 103.8 | 103.7 | 9.5 | 100.5 | 102.4 | 102.98 | 102.92
-3.4 | 2.125 | 104.8 | 105.575 | 105.4333 | 105.3 | 3.2 | 103.6 | 104.925 | 105 | 104.8714 | 4.9 | 102.2 | 104.4 | 104.48 | 104.4143 | 7.55 | 100.9 | 103.3 | 103.76 | 103.6 | 10.15 | 100.6 | 102.4 | 102.9 | 102.9
-3.3 | 2.3 | 104.7 | 105.4 | 105.26 | 105.2 | 3.4 | 103.6 | 104.88 | 104.94 | 104.7167 | 5.1333 | 102.3 | 104.3 | 104.36 | 104.3429 | 8.05 | 101 | 103.2 | 103.65 | 103.55 | 10.9 | 100.7 | 102.4 | 102.8 | 102.8667
-3.2 | 2.4 | 104.7 | 105.28 | 105.12 | 104.96 | 3.7 | 103.6 | 104.65 | 104.76 | 104.675 | 5.6 | 102.3 | 104.2 | 104.24 | 104.2125 | 8.725 | 101.1 | 103.2 | 103.5 | 103.46 | 12.3 | 100.7 | 102.325 | 102.7833 | 102.8
-3.1 | 2.575 | 104.6 | 105.05 | 104.94 | 104.8 | 4 | 103.6 | 104.5 | 104.6 | 104.5 | 6 | 102.35 | 104 | 104.14 | 104.04 | 9.7 | 101.1333 | 103.1143 | 103.5 | 103.3833 | 14.15 | 100.8 | 102.3 | 102.7 | 102.725
-3 | 2.875 | 104.6 | 104.9 | 104.74 | 104.7167 | 4.5 | 103.5 | 104.4 | 104.5333 | 104.4 | 6.6 | 102.4 | 103.9 | 104.06 | 103.9714 | 11 | 101.2 | 103.1 | 103.3333 | 103.3 | 16.95 | 100.8 | 102.3 | 102.7 | 102.6
-2.9 | 3.1667 | 104.5 | 104.8 | 104.625 | 104.52 | 5 | 103.5 | 104.3 | 104.4571 | 104.2571 | 7.6667 | 102.4 | 103.8 | 103.9 | 103.9 | 12.6 | 101.2667 | 103 | 103.2667 | 103.15 | 21.375 | 100.9 | 102.3 | 102.66 | 102.6
-2.8 | 3.6 | 104.375 | 104.625 | 104.475 | 104.4 | 5.95 | 103.5 | 104.24 | 104.3 | 104.05 | 8.9 | 102.4333 | 103.7 | 103.8 | 103.8 | 15.7 | 101.3 | 102.9 | 103.2 | 103.05 | 28.45 | 101 | 102.2 | 102.5 | 102.5
-2.7 | 4.4 | 104.3 | 104.5 | 104.3 | 104.3 | 7.1 | 103.4 | 104.1 | 104.1333 | 103.9714 | 10.6 | 102.5 | 103.68 | 103.7 | 103.6833 | 20.6 | 101.3 | 102.8 | 103.1 | 103 | 36.875 | 101.05 | 102.2 | 102.4 | 102.4
-2.6 | 5.6 | 104.2 | 104.36 | 104.2 | 104.15 | 9.1 | 103.4 | 103.9333 | 103.8714 | 103.9 | 13.6333 | 102.5 | 103.6 | 103.6 | 103.5167 | 27.2 | 101.4 | 102.7 | 102.9833 | 102.9 | 46.8 | 101.1 | 102.1667 | 102.3833 | 102.375
-2.5 | 7.6 | 104.1 | 104.225 | 104.05 | 104 | 11.9667 | 103.3 | 103.85 | 103.7286 | 103.7714 | 19 | 102.5 | 103.5 | 103.475 | 103.35 | 36.9 | 101.4 | 102.6 | 102.8 | 102.8 | 57.1 | 101.1 | 102.1 | 102.2167 | 102.3
-2.4 | 11.2 | 104 | 104.1 | 103.9 | 103.9 | 17.7 | 103.25 | 103.7 | 103.6857 | 103.6286 | 26.8 | 102.5 | 103.3 | 103.3833 | 103.2 | 47.1333 | 101.4 | 102.5 | 102.7 | 102.7 | 67.6 | 101.1 | 102 | 102.2 | 102.2
-2.3 | 18.0667 | 103.8 | 103.975 | 103.8 | 103.675 | 26.1 | 103.2 | 103.6 | 103.5429 | 103.4667 | 36.875 | 102.5 | 103.175 | 103.2 | 103.1167 | 58.7667 | 101.4 | 102.4667 | 102.6 | 102.6 | 76.8 | 101.2 | 102 | 102.1857 | 102.2
-2.2 | 28.4 | 103.7 | 103.76 | 103.56 | 103.54 | 37.1 | 103.1 | 103.5167 | 103.4222 | 103.3 | 49.025 | 102.5 | 103.08 | 103.1 | 103.0333 | 69.4667 | 101.4 | 102.4 | 102.5 | 102.5 | 85.1 | 101.2 | 101.9 | 102.1 | 102.1
-2.1 | 39.9 | 103.6 | 103.6 | 103.4 | 103.4 | 49.3 | 103 | 103.32 | 103.2714 | 103.1167 | 60.4 | 102.4 | 102.9 | 103 | 102.9857 | 79.775 | 101.5 | 102.3 | 102.4 | 102.5 | 90.7 | 101.2 | 101.84 | 102.0667 | 102.1
-2 | 52.6 | 103.4 | 103.44 | 103.3 | 103.275 | 61.1 | 102.9 | 103.1333 | 103.1167 | 103 | 72.3 | 102.2 | 102.8571 | 102.825 | 102.82 | 87.4667 | 101.5 | 102.2 | 102.3 | 102.3333 | 94.05 | 101.2 | 101.75 | 102 | 101.9833
-1.9 | 66.3 | 103.3 | 103.34 | 103.2 | 103.1 | 73 | 102.8333 | 102.94 | 102.9625 | 102.82 | 81.8667 | 102.2 | 102.7 | 102.7 | 102.65 | 92.05 | 101.5 | 102.15 | 102.3 | 102.2 | 96.175 | 101.1667 | 101.6 | 101.9 | 101.9
-1.8 | 78.3 | 103.1333 | 103.1 | 103.02 | 102.9167 | 83.3 | 102.7 | 102.82 | 102.8167 | 102.7 | 89.1 | 102.1 | 102.5333 | 102.55 | 102.46 | 95.25 | 101.4 | 101.9833 | 102.2 | 102.18 | 97.6 | 101.1 | 101.5143 | 101.8 | 101.7667
-1.7 | 87.925 | 102.9667 | 102.95 | 102.8333 | 102.8 | 90.56 | 102.6 | 102.58 | 102.6667 | 102.5667 | 94.05 | 102.0667 | 102.3667 | 102.4 | 102.38 | 97 | 101.32 | 101.8833 | 101.975 | 102 | 98.475 | 101.15 | 101.45 | 101.7 | 101.6
-1.6 | 94.4 | 102.8 | 102.825 | 102.66 | 102.65 | 95.3 | 102.425 | 102.5 | 102.5 | 102.4 | 96.5667 | 102 | 102.3 | 102.25 | 102.2143 | 98.3 | 101.3 | 101.8 | 101.825 | 101.9 | 99.1333 | 101.1 | 101.3833 | 101.54 | 101.6
-1.5 | 97.5333 | 102.7 | 102.6333 | 102.425 | 102.46 | 97.4667 | 102.3 | 102.3 | 102.28 | 102.3333 | 98 | 101.9 | 102.1 | 102.1 | 102.1 | 99.025 | 101.2 | 101.7 | 101.7 | 101.7429 | 99.55 | 101.1 | 101.3 | 101.5 | 101.4667
-1.4 | 99.4 | 102.5333 | 102.46 | 102.3167 | 102.24 | 98.9667 | 102.175 | 102.125 | 102.2 | 102.1667 | 99.1 | 101.7 | 101.96 | 101.9667 | 101.9875 | 99.6333 | 101.2 | 101.52 | 101.6 | 101.6 | 99.9 | 101.1 | 101.2 | 101.4 | 101.36
-1.3 | 100.2 | 102.4 | 102.2 | 102.12 | 102.06 | 99.65 | 102.0333 | 102.02 | 102 | 101.9286 | 99.6 | 101.6 | 101.825 | 101.8 | 101.85 | 100.05 | 101.1 | 101.42 | 101.5 | 101.5 | 100 | 101 | 101.1 | 101.3 | 101.24
-1.2 | 100.5 | 102.2 | 102.02 | 101.88 | 101.94 | 100.3 | 101.875 | 101.85 | 101.8167 | 101.85 | 100.0667 | 101.5 | 101.7 | 101.7 | 101.5875 | 100.1333 | 101 | 101.3 | 101.4 | 101.4 | 100.1 | 100.9 | 101.0571 | 101.2 | 101.2
-1.1 | 100.9 | 101.95 | 101.85 | 101.7 | 101.7 | 100.5 | 101.8 | 101.65 | 101.6429 | 101.6833 | 100.325 | 101.3 | 101.52 | 101.5 | 101.45 | 100.3 | 101 | 101.26 | 101.2833 | 101.3 | 100.15 | 100.8 | 101 | 101.15 | 101.05
-1 | 101 | 101.7 | 101.64 | 101.55 | 101.56 | 100.7 | 101.6 | 101.5 | 101.475 | 101.475 | 100.475 | 101.2 | 101.3667 | 101.36 | 101.3 | 100.4 | 100.9 | 101.2 | 101.2 | 101.2 | 100.2 | 100.625 | 100.9 | 101 | 100.9833
-0.9 | 101.1 | 101.5333 | 101.4 | 101.475 | 101.44 | 100.7 | 101.4 | 101.3 | 101.3 | 101.2833 | 100.55 | 101.1 | 101.3 | 101.2 | 101.1571 | 100.4 | 100.8 | 100.9667 | 101.0833 | 101.125 | 100.2 | 100.6 | 100.7333 | 100.9 | 100.8
-0.8 | 101.1 | 101.4 | 101.26 | 101.25 | 101.3 | 100.7 | 101.3 | 101.1667 | 101.1571 | 101.2 | 100.6 | 101 | 101.15 | 101.1 | 101 | 100.5 | 100.8 | 100.88 | 100.9 | 100.9333 | 100.2 | 100.6 | 100.6 | 100.8 | 100.775
-0.7 | 101 | 101.1 | 101.15 | 101.025 | 101.125 | 100.6 | 101.1 | 100.9 | 101.075 | 100.98 | 100.6 | 100.9 | 101 | 100.975 | 100.9429 | 100.4 | 100.7 | 100.75 | 100.8167 | 100.7714 | 100.3 | 100.6 | 100.5 | 100.7 | 100.7
-0.6 | 100.9 | 101 | 100.98 | 100.9 | 100.925 | 100.6 | 100.95 | 100.78 | 100.8667 | 100.8 | 100.55 | 100.7333 | 100.9 | 100.725 | 100.7 | 100.3 | 100.6 | 100.7 | 100.7571 | 100.7 | 100.3 | 100.6 | 100.475 | 100.62 | 100.6
-0.5 | 100.7 | 100.85 | 100.8 | 100.76 | 100.78 | 100.5 | 100.8 | 100.7 | 100.7 | 100.6286 | 100.5 | 100.625 | 100.8 | 100.65 | 100.65 | 100.2 | 100.4667 | 100.6 | 100.7 | 100.6 | 100.3 | 100.5 | 100.325 | 100.5333 | 100.5
-0.4 | 100.6 | 100.7 | 100.7 | 100.5 | 100.525 | 100.5 | 100.6 | 100.5 | 100.55 | 100.5 | 100.4 | 100.5 | 100.6 | 100.52 | 100.5833 | 100.2 | 100.4 | 100.5 | 100.5 | 100.5 | 100.3 | 100.4333 | 100.3 | 100.44 | 100.32
-0.3 | 100.4 | 100.45 | 100.5167 | 100.42 | 100.4 | 100.4 | 100.4 | 100.4 | 100.4 | 100.4 | 100.35 | 100.4 | 100.4333 | 100.4 | 100.4167 | 100.2 | 100.3 | 100.3 | 100.32 | 100.4 | 100.2333 | 100.3 | 100.2 | 100.3 | 100.2167
-0.2 | 100.3 | 100.3 | 100.375 | 100.2333 | 100.25 | 100.2 | 100.3 | 100.3 | 100.2625 | 100.2333 | 100.3 | 100.3 | 100.3 | 100.2 | 100.2571 | 100.2 | 100.2 | 100.2 | 100.2 | 100.25 | 100.1 | 100.2 | 100.1 | 100.2667 | 100.1
-0.1 | 100.1 | 100.1 | 100.24 | 100.1 | 100.14 | 100.1 | 100.2 | 100.1 | 100.2 | 100.0833 | 100.1 | 100.1 | 100.175 | 100.1167 | 100.1143 | 100.1 | 100.1 | 100.18 | 100.1 | 100.12 | 100.025 | 100.1 | 100.1 | 100.04 | 100.1
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.8333 | 99.9 | 99.825 | 99.9 | 99.9 | 99.8 | 99.9 | 99.875 | 99.8667 | 99.8333 | 99.825 | 99.9 | 99.8833 | 99.9 | 99.9 | 99.9 | 99.825 | 99.9 | 99.9 | 99.9 | 99.9 | 100 | 99.9 | 99.9167 | 99.9
0.2 | 99.7 | 99.7 | 99.7 | 99.7333 | 99.8 | 99.7 | 99.7667 | 99.7 | 99.8 | 99.75 | 99.7 | 99.8 | 99.8 | 99.8 | 99.7667 | 99.8 | 99.725 | 99.825 | 99.8 | 99.9 | 99.8 | 99.9 | 99.8 | 99.84 | 99.9
0.3 | 99.5 | 99.6 | 99.6 | 99.6 | 99.7 | 99.6 | 99.7 | 99.6 | 99.7 | 99.6 | 99.6 | 99.7 | 99.7 | 99.675 | 99.7 | 99.7 | 99.6 | 99.7 | 99.7 | 99.8 | 99.7 | 99.825 | 99.7 | 99.7333 | 99.8
0.4 | 99.3 | 99.4 | 99.5 | 99.5 | 99.6 | 99.4667 | 99.5 | 99.48 | 99.55 | 99.5 | 99.4667 | 99.5667 | 99.6 | 99.56 | 99.6 | 99.5 | 99.55 | 99.56 | 99.6833 | 99.7 | 99.56 | 99.7 | 99.6 | 99.7 | 99.7167
0.5 | 99.2 | 99.325 | 99.4 | 99.4 | 99.5 | 99.3 | 99.45 | 99.4 | 99.46 | 99.4286 | 99.3333 | 99.4 | 99.42 | 99.4 | 99.5 | 99.3 | 99.5 | 99.4667 | 99.6 | 99.6333 | 99.3 | 99.6 | 99.575 | 99.5 | 99.7
0.6 | 99.1 | 99.2 | 99.3 | 99.325 | 99.4 | 99.0333 | 99.35 | 99.3 | 99.38 | 99.3 | 99.2 | 99.3333 | 99.4 | 99.4 | 99.4 | 99.1 | 99.4 | 99.4 | 99.5 | 99.6 | 99.1667 | 99.48 | 99.4667 | 99.4 | 99.5833
0.7 | 98.9 | 99.14 | 99.1667 | 99.2 | 99.28 | 98.875 | 99.2667 | 99.3 | 99.28 | 99.3 | 99 | 99.2 | 99.3 | 99.3 | 99.3 | 99 | 99.375 | 99.3 | 99.4 | 99.5 | 99.1 | 99.4 | 99.3667 | 99.36 | 99.5
0.8 | 98.675 | 99 | 99.02 | 99.2 | 99.175 | 98.675 | 99.1 | 99.12 | 99.2 | 99.2 | 98.8 | 99.1 | 99.18 | 99.2 | 99.2 | 98.8 | 99.2333 | 99.2167 | 99.3 | 99.36 | 98.95 | 99.3 | 99.275 | 99.3 | 99.42
0.9 | 98.4 | 99 | 98.98 | 99 | 99.1 | 98.3333 | 98.975 | 99.0333 | 99.1 | 99.2 | 98.6 | 99 | 99.0833 | 99.1 | 99.1 | 98.7 | 99.1 | 99.15 | 99.2167 | 99.3 | 98.8667 | 99.25 | 99.1667 | 99.2 | 99.3
1 | 98.2333 | 98.9 | 98.86 | 98.9 | 99.1 | 97.95 | 98.825 | 98.94 | 99 | 99.0429 | 98.2333 | 98.85 | 99 | 99.075 | 99.0375 | 98.4 | 99 | 99.05 | 99.1 | 99.2 | 98.6 | 99.1 | 99 | 99.2 | 99.225
1.1 | 97.8667 | 98.8 | 98.8 | 98.8667 | 99 | 97.7 | 98.7333 | 98.9 | 98.8667 | 98.9125 | 97.85 | 98.8 | 98.9 | 98.94 | 98.8875 | 98.2 | 98.875 | 99 | 99.1 | 99.1 | 98.35 | 98.9 | 98.9 | 99.1667 | 99.2
1.2 | 97.3667 | 98.7 | 98.7 | 98.8 | 98.9 | 97.05 | 98.7 | 98.7 | 98.7833 | 98.85 | 97.375 | 98.675 | 98.85 | 98.8 | 98.8 | 97.75 | 98.7 | 98.9 | 98.96 | 99 | 98 | 98.76 | 98.8 | 99.1 | 99.1
1.3 | 96.7 | 98.6 | 98.7 | 98.7 | 98.8 | 96.4 | 98.6 | 98.6167 | 98.7 | 98.8 | 96.875 | 98.5 | 98.75 | 98.7 | 98.7125 | 97.425 | 98.5333 | 98.8 | 98.8 | 98.88 | 97.75 | 98.625 | 98.7167 | 99.04 | 99
1.4 | 95.6 | 98.45 | 98.575 | 98.62 | 98.76 | 95.4 | 98.45 | 98.5571 | 98.6 | 98.7 | 95.975 | 98.4 | 98.6667 | 98.6833 | 98.6667 | 97.05 | 98.3 | 98.7 | 98.8 | 98.8 | 97.525 | 98.5 | 98.7 | 98.9 | 99
1.5 | 93.625 | 98.3 | 98.44 | 98.5 | 98.7 | 94.1 | 98.3333 | 98.5 | 98.5833 | 98.5833 | 94.9 | 98.2333 | 98.6 | 98.6 | 98.6 | 96.4 | 98.2 | 98.62 | 98.7 | 98.7167 | 97.1 | 98.3333 | 98.6 | 98.85 | 98.9
1.6 | 90.0333 | 98.2 | 98.26 | 98.4 | 98.575 | 91.5 | 98.3 | 98.3667 | 98.4375 | 98.5 | 93.4 | 98.1 | 98.4333 | 98.5 | 98.5 | 95.65 | 98.1 | 98.5 | 98.66 | 98.7 | 96.5 | 98.2 | 98.6 | 98.7 | 98.82
1.7 | 83.575 | 98.1 | 98.125 | 98.3 | 98.45 | 87.375 | 98.1 | 98.225 | 98.4 | 98.4 | 90.8 | 97.9 | 98.3333 | 98.4 | 98.4167 | 94.5 | 98 | 98.45 | 98.54 | 98.6 | 96 | 98.1 | 98.5 | 98.6 | 98.7167
1.8 | 74.25 | 98 | 98.02 | 98.26 | 98.4 | 81 | 97.9333 | 98.12 | 98.3 | 98.24 | 87 | 97.8 | 98.2 | 98.3 | 98.3714 | 92.65 | 97.9 | 98.325 | 98.425 | 98.5 | 95.3667 | 98 | 98.4 | 98.5167 | 98.64
1.9 | 63.5 | 97.775 | 98 | 98.125 | 98.3 | 71.375 | 97.8 | 98 | 98.2 | 98.2 | 80.4 | 97.6 | 98.0167 | 98.2 | 98.3 | 90.3 | 97.675 | 98.2 | 98.3 | 98.42 | 94.15 | 97.9 | 98.3 | 98.4667 | 98.56
2 | 50.9 | 97.7 | 97.9 | 98 | 98.2 | 60.2667 | 97.6 | 97.95 | 98.15 | 98.0333 | 71.4 | 97.45 | 97.9167 | 98.1 | 98.2 | 86.3 | 97.5 | 98.1 | 98.2833 | 98.3 | 92.4333 | 97.8 | 98.15 | 98.36 | 98.5
2.1 | 38.6333 | 97.575 | 97.8 | 97.9 | 98.05 | 48.075 | 97.425 | 97.84 | 98.025 | 97.9 | 60.6667 | 97.35 | 97.8 | 98.1 | 98.1 | 80.3 | 97.4 | 98.05 | 98.2 | 98.2 | 90.025 | 97.675 | 98.04 | 98.225 | 98.4167
2.2 | 27.2667 | 97.3667 | 97.7 | 97.8 | 97.95 | 36.8 | 97.275 | 97.72 | 97.9 | 97.8 | 49.6667 | 97.2 | 97.7 | 97.9 | 97.9714 | 71.9 | 97.2 | 98 | 98.1 | 98.1 | 85.6667 | 97.6 | 97.9 | 98.2 | 98.3
2.3 | 17.5667 | 97.1667 | 97.6 | 97.7 | 97.9 | 26.4 | 97.075 | 97.6 | 97.7625 | 97.7167 | 38.5333 | 97 | 97.6 | 97.8 | 97.8375 | 62.2 | 96.9667 | 97.9 | 97.9 | 98 | 80.2667 | 97.45 | 97.8 | 98.075 | 98.2
2.4 | 10.8 | 97.1 | 97.4 | 97.6 | 97.8 | 17.6 | 97 | 97.5 | 97.7 | 97.6714 | 28 | 96.8 | 97.525 | 97.68 | 97.7125 | 50.95 | 96.8 | 97.65 | 97.9 | 97.9 | 72.1 | 97.2 | 97.74 | 98 | 98.1
2.5 | 7.6 | 96.9 | 97.26 | 97.5 | 97.65 | 12.2333 | 96.8 | 97.4 | 97.55 | 97.5286 | 20.1333 | 96.55 | 97.4 | 97.575 | 97.5875 | 40.825 | 96.6667 | 97.6 | 97.8 | 97.8333 | 62.9 | 97 | 97.7 | 97.9 | 98
2.6 | 5.5 | 96.7 | 97.14 | 97.3 | 97.5 | 8.8 | 96.6 | 97.225 | 97.4857 | 97.4857 | 14.3667 | 96.32 | 97.3 | 97.425 | 97.4571 | 31 | 96.4667 | 97.4 | 97.625 | 97.7714 | 52.625 | 96.85 | 97.6 | 97.8 | 97.9
2.7 | 4.3 | 96.4 | 97.025 | 97.18 | 97.44 | 6.9 | 96.4333 | 97.1 | 97.32 | 97.3333 | 11.0333 | 96.05 | 97.1 | 97.375 | 97.3143 | 23.2667 | 96.1333 | 97.2667 | 97.525 | 97.6167 | 42.76 | 96.6 | 97.5 | 97.7 | 97.8
2.8 | 3.62 | 96.2 | 96.875 | 97.0714 | 97.3 | 5.75 | 96.1667 | 96.9833 | 97.14 | 97.175 | 9 | 95.7667 | 97.06 | 97.25 | 97.2 | 17.4 | 95.925 | 97.1667 | 97.45 | 97.45 | 33.6 | 96.4333 | 97.3 | 97.6 | 97.7
2.9 | 3.1 | 96.025 | 96.7 | 96.9167 | 97.1 | 4.9 | 96 | 96.7667 | 96.96 | 97.02 | 7.7 | 95.5333 | 96.9 | 97.1 | 97.1 | 13.7 | 95.75 | 97 | 97.3 | 97.3 | 25.7 | 96.175 | 97.16 | 97.5 | 97.7
3 | 2.8 | 95.8 | 96.6 | 96.8 | 97 | 4.35 | 95.7667 | 96.64 | 96.9 | 96.95 | 6.8 | 95.2667 | 96.7333 | 97 | 96.9667 | 11.5 | 95.5 | 96.8333 | 97.2 | 97.2286 | 19.9 | 95.85 | 97 | 97.5 | 97.5167
3.1 | 2.5 | 95.7 | 96.4833 | 96.7 | 96.9 | 3.9 | 95.5333 | 96.48 | 96.8 | 96.8143 | 6.1 | 95.0667 | 96.54 | 96.9 | 96.825 | 10 | 95.2 | 96.6667 | 97.1 | 97.1333 | 15.8 | 95.6 | 96.9 | 97.4 | 97.4
3.2 | 2.3 | 95.4667 | 96.3 | 96.56 | 96.7 | 3.6 | 95.2667 | 96.3667 | 96.6714 | 96.7 | 5.5333 | 94.7667 | 96.4 | 96.78 | 96.7857 | 8.8333 | 94.9333 | 96.575 | 96.96 | 96.975 | 13.3 | 95.45 | 96.7333 | 97.3 | 97.2
3.3 | 2.15 | 95.2 | 96.2 | 96.425 | 96.6 | 3.36 | 94.9667 | 96.2 | 96.5167 | 96.6167 | 5.15 | 94.4333 | 96.2833 | 96.68 | 96.6333 | 7.975 | 94.6 | 96.425 | 96.8 | 96.9 | 11.72 | 95.3 | 96.66 | 97.175 | 97.16
3.4 | 2.1 | 95 | 96.1 | 96.4 | 96.525 | 3.1333 | 94.7 | 96.1 | 96.4714 | 96.4429 | 4.8333 | 94.1 | 96.1 | 96.6 | 96.5571 | 7.45 | 94.3667 | 96.3 | 96.72 | 96.8 | 10.65 | 94.95 | 96.44 | 97.04 | 97.08
3.5 | 1.9667 | 94.8 | 95.98 | 96.3 | 96.4667 | 2.95 | 94.5 | 96.0429 | 96.3286 | 96.3 | 4.6333 | 93.8667 | 95.9833 | 96.46 | 96.4 | 7.15 | 94.1333 | 96.2 | 96.6 | 96.7 | 9.75 | 94.7667 | 96.38 | 96.92 | 96.9
3.6 | 1.9 | 94.6 | 95.8 | 96.1 | 96.35 | 2.8 | 94.15 | 95.9 | 96.2 | 96.2333 | 4.4 | 93.6 | 95.9 | 96.3333 | 96.3 | 6.8 | 93.8667 | 96.1 | 96.5 | 96.58 | 9.0667 | 94.5 | 96.25 | 96.8 | 96.8
3.7 | 1.8 | 94.4 | 95.7 | 96.0167 | 96.2 | 2.6 | 93.9 | 95.75 | 96.1 | 96.15 | 4.175 | 93.25 | 95.76 | 96.1833 | 96.2167 | 6.5 | 93.5 | 95.9714 | 96.4 | 96.48 | 8.6333 | 94.2 | 96.05 | 96.7333 | 96.7
3.8 | 1.7 | 94.2 | 95.6 | 95.92 | 96.1 | 2.5 | 93.4667 | 95.64 | 96 | 96.075 | 3.9667 | 92.9 | 95.65 | 96.1 | 96.14 | 6.175 | 93.175 | 95.72 | 96.275 | 96.3 | 8.15 | 93.95 | 95.9 | 96.675 | 96.58
3.9 | 1.6 | 93.95 | 95.5167 | 95.9 | 96.1 | 2.4 | 93.2 | 95.525 | 95.9 | 96 | 3.75 | 92.65 | 95.5 | 96 | 95.94 | 6 | 92.9 | 95.5 | 96.175 | 96.225 | 7.7667 | 93.6 | 95.8 | 96.5 | 96.4
4 | 1.6 | 93.7 | 95.5 | 95.85 | 96 | 2.3 | 92.925 | 95.4 | 95.9 | 95.9 | 3.6 | 92.2333 | 95.34 | 95.84 | 95.86 | 5.75 | 92.5 | 95.4 | 96.0833 | 96.0833 | 7.4 | 93.2667 | 95.6 | 96.35 | 96.3
4.1 | 1.5 | 93.3 | 95.36 | 95.8 | 95.9 | 2.2 | 92.6 | 95.3167 | 95.7714 | 95.8 | 3.4667 | 92 | 95.24 | 95.8 | 95.78 | 5.5 | 92.15 | 95.3 | 95.9167 | 96 | 7.18 | 92.9 | 95.46 | 96.3 | 96.2
4.2 | 1.5 | 92.9 | 95.3 | 95.72 | 95.9 | 2.1 | 92.15 | 95.2 | 95.7 | 95.7286 | 3.3 | 91.725 | 95.1 | 95.7 | 95.7 | 5.4 | 91.8333 | 95.2 | 95.8 | 95.92 | 7.05 | 92.625 | 95.325 | 96.1 | 96.04
4.3 | 1.4 | 92.3 | 95.24 | 95.6 | 95.8 | 2 | 91.7 | 95.15 | 95.6 | 95.6667 | 3.2 | 91.125 | 95 | 95.58 | 95.6 | 5.25 | 91.5667 | 95.02 | 95.6833 | 95.8 | 7 | 92.45 | 95.2 | 96.0333 | 95.9
4.4 | 1.4 | 91.45 | 95.14 | 95.56 | 95.7333 | 1.9 | 91 | 95 | 95.5 | 95.6 | 3.1 | 90.7 | 94.9 | 95.5 | 95.5 | 5 | 91.225 | 94.8333 | 95.5167 | 95.7 | 6.7667 | 92.125 | 95.08 | 95.8 | 95.8
4.5 | 1.3 | 89.9667 | 95 | 95.5 | 95.62 | 1.8 | 90.2 | 94.96 | 95.5 | 95.4333 | 3 | 90.2667 | 94.8 | 95.4 | 95.375 | 4.85 | 90.65 | 94.66 | 95.45 | 95.6 | 6.5 | 91.7667 | 94.8667 | 95.7333 | 95.7833
4.6 | 1.3 | 87.8667 | 94.9 | 95.4 | 95.5 | 1.8 | 89.1 | 94.85 | 95.4 | 95.4 | 2.9 | 89.65 | 94.7 | 95.36 | 95.3 | 4.7 | 90.38 | 94.56 | 95.35 | 95.45 | 6.4 | 91.35 | 94.7 | 95.66 | 95.7
4.7 | 1.3 | 83.7333 | 94.86 | 95.3 | 95.4333 | 1.7 | 87.5 | 94.725 | 95.3 | 95.3 | 2.75 | 88.9 | 94.5167 | 95.25 | 95.2 | 4.6 | 90.0667 | 94.425 | 95.2667 | 95.3 | 6.175 | 90.975 | 94.6 | 95.55 | 95.64
4.8 | 1.2 | 76.775 | 94.75 | 95.2 | 95.3667 | 1.675 | 84.8 | 94.6 | 95.2143 | 95.2 | 2.6 | 87.8333 | 94.4 | 95.12 | 95.1333 | 4.5 | 89.475 | 94.4 | 95.2 | 95.2 | 6 | 90.725 | 94.4333 | 95.46 | 95.4571
4.9 | 1.2 | 67.225 | 94.6 | 95.075 | 95.2833 | 1.6 | 80.4 | 94.5 | 95.15 | 95.1 | 2.5 | 86.2667 | 94.2667 | 95 | 95 | 4.3 | 89.1 | 94.2 | 95.05 | 95.1 | 5.8 | 90.35 | 94.3 | 95.3833 | 95.275
5 | 1.2 | 56.8 | 94.5 | 95 | 95.1167 | 1.5 | 73.45 | 94.45 | 95 | 95.02 | 2.4 | 84.4 | 94.2 | 94.9167 | 95 | 4.2 | 88.5667 | 94.1 | 95 | 95 | 5.7 | 90.02 | 94.2333 | 95.2167 | 95.18
5.1 | 1.1 | 45.65 | 94.3333 | 94.9 | 95.025 | 1.5 | 64.1333 | 94.2833 | 94.9 | 94.94 | 2.4 | 81.1667 | 94.0667 | 94.85 | 94.8833 | 4.075 | 87.7 | 93.9 | 94.9 | 94.9 | 5.5 | 89.5333 | 94.125 | 95.125 | 95.1
5.2 | 1.1 | 34.9 | 94.22 | 94.8 | 94.9167 | 1.4 | 54.3 | 94.1167 | 94.8 | 94.84 | 2.3 | 75.7 | 93.9 | 94.7857 | 94.8 | 3.9333 | 87 | 93.8 | 94.8 | 94.82 | 5.5 | 89.25 | 94 | 95 | 95
5.3 | 1.1 | 24.7 | 94.04 | 94.7 | 94.75 | 1.4 | 43.3667 | 94 | 94.7 | 94.76 | 2.2 | 68.7 | 93.8 | 94.7 | 94.6667 | 3.8 | 85.9 | 93.7 | 94.7 | 94.74 | 5.38 | 88.6333 | 93.9333 | 94.9 | 94.9
5.4 | 1.025 | 17 | 93.9 | 94.6 | 94.7 | 1.4 | 33.5333 | 93.9 | 94.56 | 94.5667 | 2.125 | 59.2 | 93.65 | 94.6 | 94.6 | 3.7 | 84.5667 | 93.56 | 94.575 | 94.65 | 5.2 | 88.225 | 93.6667 | 94.8286 | 94.8
5.5 | 1 | 12 | 93.8 | 94.45 | 94.5 | 1.3 | 25.3333 | 93.76 | 94.4833 | 94.5 | 2.1 | 49.7333 | 93.52 | 94.5 | 94.4286 | 3.6333 | 82.7667 | 93.4 | 94.48 | 94.5833 | 5.1 | 87.7 | 93.5 | 94.6667 | 94.675
5.6 | 1 | 9.5 | 93.7 | 94.3333 | 94.44 | 1.3 | 18.2 | 93.6 | 94.3375 | 94.3 | 2 | 40.6 | 93.38 | 94.3 | 94.3 | 3.5 | 79.45 | 93.2333 | 94.3 | 94.4 | 4.9 | 87.05 | 93.3 | 94.6 | 94.6
5.7 | 1 | 8.1 | 93.56 | 94.2667 | 94.32 | 1.2 | 14.0667 | 93.42 | 94.2125 | 94.2286 | 2 | 31.8 | 93.2 | 94.2 | 94.2 | 3.4 | 74.8 | 93 | 94.225 | 94.3 | 4.7 | 86.3 | 93.1667 | 94.4333 | 94.4667
5.8 | 1 | 7.1333 | 93.38 | 94.1 | 94.22 | 1.2 | 11.6 | 93.3 | 94.0625 | 94.0714 | 1.9 | 24.325 | 93.08 | 94 | 94.1333 | 3.3 | 68 | 92.7333 | 94.1 | 94.175 | 4.6 | 85.2 | 93 | 94.35 | 94.375
5.9 | 0.9 | 6.5333 | 93.2 | 94 | 94.0833 | 1.2 | 10.1 | 93.08 | 93.9375 | 93.9286 | 1.9 | 19 | 92.9 | 93.9 | 93.95 | 3.25 | 60.8 | 92.6 | 94 | 94.05 | 4.5 | 83.7667 | 92.82 | 94.15 | 94.3
6 | 0.9 | 6.05 | 93.1 | 93.8167 | 93.975 | 1.2 | 9.1 | 92.9 | 93.8667 | 93.8 | 1.8 | 15.8 | 92.74 | 93.8 | 93.7833 | 3.2 | 52 | 92.55 | 93.8571 | 93.925 | 4.4667 | 81.7 | 92.7167 | 94.1 | 94.1833
6.1 | 0.9 | 5.7 | 92.94 | 93.7 | 93.82 | 1.1 | 8.3667 | 92.8 | 93.7143 | 93.6714 | 1.7667 | 13.95 | 92.55 | 93.6667 | 93.7 | 3.1 | 43.6 | 92.35 | 93.7 | 93.8 | 4.4 | 79.3 | 92.58 | 93.95 | 94.0167
6.2 | 0.9 | 5.4333 | 92.7333 | 93.5 | 93.625 | 1.1 | 7.7667 | 92.6714 | 93.5571 | 93.5 | 1.7 | 12.5 | 92.38 | 93.5 | 93.575 | 3.1 | 36 | 92.14 | 93.6333 | 93.7333 | 4.3 | 74.4667 | 92.4667 | 93.78 | 93.95
6.3 | 0.9 | 5.2333 | 92.6 | 93.4 | 93.5167 | 1.1 | 7.3333 | 92.5 | 93.425 | 93.4 | 1.7 | 11.6 | 92.175 | 93.4167 | 93.45 | 3 | 29.1 | 91.94 | 93.5 | 93.5667 | 4.25 | 68.7 | 92.3 | 93.66 | 93.8
6.4 | 0.9 | 5 | 92.5 | 93.26 | 93.4 | 1.1 | 7 | 92.32 | 93.2833 | 93.3 | 1.6 | 10.85 | 91.96 | 93.24 | 93.3 | 2.9 | 23.85 | 91.7 | 93.3333 | 93.45 | 4.1667 | 61.6 | 92.08 | 93.5 | 93.7
6.5 | 0.9 | 4.8 | 92.26 | 93.125 | 93.275 | 1.1 | 6.7 | 92.2 | 93.1167 | 93.2 | 1.6 | 10.3 | 91.76 | 93.1 | 93.1625 | 2.9 | 20.2 | 91.56 | 93.2 | 93.32 | 4 | 54 | 91.85 | 93.4833 | 93.5
6.6 | 0.9 | 4.7 | 92.1 | 93 | 93.1 | 1 | 6.4 | 92.025 | 93.05 | 93.1 | 1.6 | 9.8 | 91.6 | 92.9 | 93.0286 | 2.8 | 18.25 | 91.35 | 93.08 | 93.1167 | 3.9667 | 46.5667 | 91.6667 | 93.3167 | 93.4
6.7 | 0.8 | 4.6 | 91.92 | 92.8 | 92.96 | 1 | 6.1667 | 91.8 | 92.9 | 93.0333 | 1.52 | 9.45 | 91.38 | 92.8 | 92.8857 | 2.7 | 16.8 | 91.14 | 92.9 | 93 | 3.9 | 38.9 | 91.5143 | 93.3 | 93.24
6.8 | 0.8 | 4.4 | 91.78 | 92.675 | 92.82 | 1 | 5.9 | 91.6 | 92.82 | 92.84 | 1.5 | 9.1 | 91.2 | 92.7 | 92.7333 | 2.7 | 15.74 | 91.02 | 92.75 | 92.9 | 3.8 | 32.2 | 91.35 | 93.2 | 93.12
6.9 | 0.8 | 4.3 | 91.66 | 92.54 | 92.7 | 1 | 5.7 | 91.4333 | 92.6571 | 92.7333 | 1.5 | 8.775 | 91 | 92.6 | 92.6667 | 2.6 | 14.95 | 90.8 | 92.625 | 92.72 | 3.8 | 27.0667 | 91.1667 | 93.12 | 93.02
7 | 0.8 | 4.2 | 91.4 | 92.42 | 92.6 | 1 | 5.5 | 91.2 | 92.5 | 92.4857 | 1.5 | 8.425 | 90.75 | 92.4 | 92.4167 | 2.6 | 14.375 | 90.6 | 92.58 | 92.6 | 3.7 | 23.6667 | 90.9167 | 92.92 | 92.9
7.1 | 0.8 | 4.06 | 91.3 | 92.3 | 92.4 | 1 | 5.3333 | 91.06 | 92.3429 | 92.3333 | 1.4 | 8.2 | 90.6 | 92.3 | 92.3 | 2.5 | 13.825 | 90.4333 | 92.4 | 92.5 | 3.6 | 21.4 | 90.6167 | 92.8 | 92.68
7.2 | 0.8 | 4 | 91.175 | 92.15 | 92.24 | 1 | 5.2 | 90.925 | 92.3 | 92.1833 | 1.4 | 7.925 | 90.4 | 92.2 | 92.1667 | 2.5 | 13.35 | 90.2 | 92.1 | 92.3667 | 3.425 | 19.8 | 90.5 | 92.64 | 92.52
7.3 | 0.8 | 3.9 | 90.975 | 92.02 | 92.1 | 0.975 | 5 | 90.62 | 92.1333 | 92.1 | 1.3 | 7.7 | 90.12 | 92.02 | 92.0143 | 2.4 | 12.95 | 89.9667 | 91.9667 | 92.1 | 3.4 | 18.8 | 90.3 | 92.55 | 92.4
7.4 | 0.8 | 3.7667 | 90.84 | 91.9 | 91.94 | 0.9 | 4.9 | 90.5 | 91.9667 | 91.9714 | 1.3 | 7.4 | 89.92 | 91.9167 | 91.8333 | 2.4 | 12.7 | 89.75 | 91.875 | 92 | 3.4 | 17.9 | 90.1 | 92.42 | 92.3
7.5 | 0.8 | 3.675 | 90.7 | 91.75 | 91.84 | 0.9 | 4.7 | 90.36 | 91.8143 | 91.8286 | 1.3 | 7.3 | 89.7 | 91.7 | 91.75 | 2.3 | 12.25 | 89.4833 | 91.76 | 91.82 | 3.2333 | 17.2 | 89.875 | 92.2 | 92.125
7.6 | 0.75 | 3.6 | 90.5 | 91.6 | 91.74 | 0.9 | 4.6 | 90.1667 | 91.6714 | 91.6833 | 1.3 | 7.05 | 89.56 | 91.6 | 91.5667 | 2.3 | 11.95 | 89.2333 | 91.5667 | 91.675 | 3.2 | 16.7333 | 89.7 | 92.06 | 92
7.7 | 0.7 | 3.5 | 90.275 | 91.4 | 91.6 | 0.9 | 4.4667 | 90 | 91.6 | 91.5167 | 1.3 | 6.8 | 89.25 | 91.4 | 91.5 | 2.2 | 11.65 | 89.05 | 91.5 | 91.55 | 3.2 | 16.3667 | 89.5 | 91.875 | 91.85
7.8 | 0.7 | 3.4 | 90.08 | 91.3667 | 91.4667 | 0.9 | 4.3 | 89.76 | 91.4 | 91.3667 | 1.3 | 6.6 | 89.1 | 91.3 | 91.3857 | 2.2 | 11.35 | 88.76 | 91.3 | 91.35 | 3.125 | 15.875 | 89.275 | 91.7 | 91.675
7.9 | 0.7 | 3.3667 | 89.875 | 91.2 | 91.275 | 0.9 | 4.25 | 89.48 | 91.2333 | 91.275 | 1.2 | 6.5 | 88.8333 | 91.14 | 91.1857 | 2.2 | 11.1 | 88.4667 | 91.125 | 91.22 | 3.1 | 15.45 | 88.92 | 91.5 | 91.4667
8 | 0.7 | 3.275 | 89.7 | 91.025 | 91.15 | 0.8 | 4.125 | 89.3333 | 91.14 | 91.0857 | 1.2 | 6.3 | 88.6 | 91.025 | 91.06 | 2.1 | 10.75 | 88.2 | 91 | 91.1 | 3.05 | 15.1333 | 88.6667 | 91.375 | 91.4
8.1 | 0.7 | 3.2 | 89.5333 | 90.9 | 91 | 0.875 | 4 | 89.15 | 90.98 | 91 | 1.2 | 6.1 | 88.38 | 90.9 | 90.8833 | 2.1 | 10.475 | 88 | 90.84 | 90.9833 | 3 | 14.675 | 88.5 | 91.2 | 91.2333
8.2 | 0.7 | 3.1333 | 89.325 | 90.76 | 90.9 | 0.875 | 3.9333 | 88.82 | 90.7833 | 90.8 | 1.1667 | 6 | 88.1 | 90.7 | 90.7286 | 2.0667 | 10.3 | 87.68 | 90.725 | 90.8 | 2.9 | 14.4 | 88.3 | 91.1 | 91.06
8.3 | 0.7 | 3.1 | 89.14 | 90.575 | 90.7333 | 0.8 | 3.8 | 88.5 | 90.6167 | 90.7 | 1.1 | 5.7667 | 87.82 | 90.52 | 90.5714 | 2 | 10 | 87.3667 | 90.5 | 90.625 | 2.9 | 14.2 | 88.08 | 91 | 90.95
8.4 | 0.7 | 3.05 | 88.92 | 90.45 | 90.65 | 0.8 | 3.7 | 88.28 | 90.46 | 90.5667 | 1.1 | 5.6 | 87.5667 | 90.325 | 90.4375 | 2 | 9.775 | 87.2333 | 90.3333 | 90.5167 | 2.9 | 13.95 | 87.7667 | 90.7667 | 90.82
8.5 | 0.7 | 3 | 88.56 | 90.3 | 90.425 | 0.8 | 3.6 | 88.06 | 90.3 | 90.4 | 1.1 | 5.5 | 87.225 | 90.22 | 90.2 | 1.95 | 9.6 | 87 | 90.16 | 90.275 | 2.9 | 13.5333 | 87.4833 | 90.6 | 90.575
8.6 | 0.7 | 2.9 | 88.34 | 90.125 | 90.275 | 0.8 | 3.6 | 87.8 | 90.2143 | 90.2333 | 1.1 | 5.375 | 86.9833 | 90.1 | 90.0625 | 1.9 | 9.25 | 86.7 | 90 | 90.0833 | 2.9 | 13.4 | 87.1833 | 90.4 | 90.425
8.7 | 0.6 | 2.9 | 88.14 | 90 | 90.1 | 0.8 | 3.5 | 87.6 | 90.0833 | 90.1 | 1.1 | 5.2 | 86.7 | 89.96 | 90 | 1.875 | 9.1 | 86.4667 | 89.9 | 90 | 2.75 | 13 | 87.0286 | 90.2333 | 90.26
8.8 | 0.65 | 2.8 | 87.84 | 89.9 | 90 | 0.8 | 3.4 | 87.2 | 89.9286 | 89.92 | 1.1 | 5.1 | 86.45 | 89.78 | 89.8 | 1.8 | 8.975 | 86.18 | 89.7833 | 89.84 | 2.7 | 12.7 | 86.8333 | 90.1 | 90.14
8.9 | 0.7 | 2.8 | 87.65 | 89.825 | 89.9 | 0.8 | 3.3 | 87 | 89.8 | 89.8 | 1 | 5 | 86.125 | 89.6 | 89.7 | 1.8 | 8.8 | 85.8 | 89.6167 | 89.7 | 2.6 | 12.5 | 86.6 | 90 | 90
9 | 0.7 | 2.7 | 87.3333 | 89.7 | 89.8 | 0.8 | 3.3 | 86.75 | 89.65 | 89.6833 | 1 | 4.8333 | 85.82 | 89.4833 | 89.5167 | 1.8 | 8.625 | 85.55 | 89.5 | 89.55 | 2.6 | 12.2333 | 86.2667 | 89.825 | 89.8167
9.1 | 0.7 | 2.675 | 87.05 | 89.5 | 89.7 | 0.8 | 3.2 | 86.32 | 89.5 | 89.5 | 1 | 4.7 | 85.4 | 89.2833 | 89.3667 | 1.8 | 8.4 | 85.4 | 89.3 | 89.38 | 2.6 | 12.05 | 85.98 | 89.7 | 89.65
9.2 |  | 2.6 | 86.7 | 89.46 | 89.6 | 0.8 | 3.125 | 86.05 | 89.34 | 89.45 | 1 | 4.6 | 85.075 | 89.1167 | 89.1833 | 1.7 | 8.2 | 84.9667 | 89.1 | 89.1667 | 2.6 | 11.8333 | 85.75 | 89.5333 | 89.4833
9.3 |  | 2.6 | 86.3 | 89.4 | 89.5 | 0.7 | 3.1 | 85.66 | 89.3 | 89.2714 | 1 | 4.5 | 84.84 | 88.9333 | 89.0167 | 1.7 | 7.975 | 84.7167 | 89 | 89 | 2.6 | 11.6 | 85.425 | 89.325 | 89.2333
9.4 |  | 2.5 | 85.85 | 89.3 | 89.5 | 0.7 | 3 | 85.2 | 89.14 | 89.2 | 1 | 4.4 | 84.44 | 88.825 | 88.85 | 1.7 | 7.775 | 84.4667 | 88.8333 | 88.85 | 2.6 | 11.3333 | 85.22 | 89.2 | 89.15
9.5 |  | 2.5 | 85.1 | 89.3 | 89.4 |  | 2.9667 | 84.9 | 89.0714 | 89.0875 | 0.9 | 4.3 | 84.05 | 88.7 | 88.7 | 1.7 | 7.6 | 84.1 | 88.5 | 88.64 | 2.4 | 11.1 | 85 | 88.95 | 89.0667
9.6 |  | 2.4 | 83.9667 | 89.225 | 89.4 |  | 2.9 | 84.45 | 89 | 89 | 0.9 | 4.2 | 83.675 | 88.55 | 88.6 | 1.7 | 7.5 | 83.8 | 88.4 | 88.5667 | 2.4 | 10.875 | 84.65 | 88.8857 | 88.88
9.7 |  | 2.4 | 82.1 | 89.2 | 89.3 |  | 2.9 | 83.75 | 88.96 | 88.9 | 0.9 | 4.2 | 83.3 | 88.5 | 88.4833 | 1.7 | 7.3 | 83.45 | 88.2167 | 88.3833 | 2.35 | 10.725 | 84.3 | 88.7333 | 88.7
9.8 |  | 2.4 | 78.4 | 89.175 | 89.3 |  | 2.8 | 83.02 | 88.9 | 88.82 | 0.9 | 4.1 | 83.025 | 88.3 | 88.3375 | 1.6 | 7.2 | 83.1 | 88.04 | 88.2 | 2.3 | 10.5 | 84.05 | 88.52 | 88.58
9.9 |  | 2.4 | 72.05 | 89.1 | 89.2 |  | 2.8 | 81.96 | 88.8 | 88.8 |  | 4 | 82.56 | 88.1857 | 88.2 | 1.6 | 7.1 | 82.7 | 87.85 | 88 | 2.3 | 10.35 | 83.7 | 88.3 | 88.4143
10 |  | 2.3 | 63.9 | 89.1 | 89.2 |  | 2.7 | 80.4 | 88.72 | 88.7833 |  | 3.9 | 82.04 | 88.0333 | 88.12 | 1.6 | 6.9 | 82.36 | 87.725 | 87.76 | 2.275 | 10.1 | 83.375 | 88.1333 | 88.1333
10.1 |  | 2.3 | 55.12 | 89 | 89.125 |  | 2.7 | 77.5 | 88.62 | 88.6286 |  | 3.8 | 81.58 | 87.96 | 87.9667 | 1.6 | 6.8 | 82.0333 | 87.5 | 87.6 | 2.2 | 10 | 83.0667 | 88 | 87.98
10.2 |  | 2.3 | 45.075 | 89 | 89.1 |  | 2.6 | 73.15 | 88.5714 | 88.54 |  | 3.8 | 80.86 | 87.8667 | 87.8857 | 1.575 | 6.64 | 81.675 | 87.3833 | 87.46 | 2.2 | 9.8667 | 82.86 | 87.8714 | 87.775
10.3 |  | 2.2 | 35.1 | 88.925 | 89 |  | 2.6 | 66.55 | 88.5 | 88.5 |  | 3.7 | 80 | 87.78 | 87.8 | 1.6 | 6.5 | 81.22 | 87.2167 | 87.25 | 2.2 | 9.65 | 82.5 | 87.6333 | 87.65
10.4 |  | 2.2 | 26.1 | 88.9 | 88.9 |  | 2.6 | 58.275 | 88.48 | 88.4286 |  | 3.6667 | 78.9 | 87.7 | 87.7 | 1.5 | 6.32 | 80.9 | 87.025 | 87.125 | 2.1 | 9.5 | 82.3333 | 87.52 | 87.475
10.5 |  | 2.2 | 19.9 | 88.86 | 88.9 |  | 2.5 | 49.475 | 88.4 | 88.3 |  | 3.5667 | 77.35 | 87.6 | 87.6571 | 1.5 | 6.2667 | 80.55 | 86.9 | 87 | 2.1 | 9.26 | 81.9333 | 87.3 | 87.325
10.6 |  | 2.2 | 15.44 | 88.8 | 88.85 |  | 2.5 | 40.3667 | 88.3286 | 88.3 |  | 3.5 | 74.92 | 87.5 | 87.5 | 1.5 | 6.2 | 80.22 | 86.8 | 86.9 | 2.08 | 9.2 | 81.6 | 87.1 | 87.1667
10.7 |  | 2.1 | 12.9 | 88.7 | 88.8 |  | 2.4667 | 31.56 | 88.2667 | 88.2125 |  | 3.5 | 70.88 | 87.4 | 87.4429 | 1.5 | 6.1 | 79.76 | 86.6333 | 86.7 | 2.1 | 9.075 | 81.24 | 86.95 | 87.0667
10.8 |  | 2.1 | 11.6 | 88.7 | 88.8 |  | 2.4 | 24.85 | 88.2 | 88.1 |  | 3.4 | 65.1 | 87.3 | 87.3857 |  | 6 | 79.3 | 86.4667 | 86.5 | 2.1 | 8.9 | 80.86 | 86.8 | 86.88
10.9 |  | 2.1 | 10.75 | 88.6 | 88.72 |  | 2.4 | 19.55 | 88.1 | 88.0571 |  | 3.3 | 57.5 | 87.225 | 87.2333 |  | 5.9 | 78.85 | 86.3 | 86.38 | 2.1 | 8.6 | 80.525 | 86.6 | 86.6
11 |  | 2.1 | 10.1 | 88.6 | 88.68 |  | 2.4 | 16.3333 | 88.1 | 88 |  | 3.3 | 49.55 | 87.1 | 87.1833 |  | 5.8 | 78.4 | 86.1667 | 86.3 | 2.1 | 8.5 | 80.1 | 86.3833 | 86.45
11.1 |  | 2.025 | 9.7 | 88.5 | 88.6 |  | 2.3 | 14.3 | 88 | 88 |  | 3.2 | 41.4 | 87 | 87.0167 |  | 5.6667 | 77.9 | 86.08 | 86.1 | 2.1 | 8.4 | 79.75 | 86.2167 | 86.2333
11.2 |  | 2 | 9.3 | 88.5 | 88.6 |  | 2.3 | 13.2 | 87.9429 | 87.9 |  | 3.2 | 34.1 | 86.92 | 87 |  | 5.55 | 77.3 | 85.8833 | 86.02 | 2.1 | 8.3 | 79.425 | 86.0571 | 86
11.3 |  | 2 | 9.08 | 88.4 | 88.5 |  | 2.25 | 12.3 | 87.9 | 87.9 |  | 3.1 | 27.28 | 86.8 | 86.9 |  | 5.5 | 76.4333 | 85.7 | 85.9 | 2.1 | 8.15 | 79 | 85.9 | 85.9
11.4 |  | 2 | 8.85 | 88.4 | 88.4 |  | 2.2 | 11.7 | 87.8 | 87.8 |  | 3.1 | 22.725 | 86.7 | 86.8 |  | 5.4 | 75.3 | 85.6 | 85.75 | 1.9 | 8 | 78.7 | 85.72 | 85.72
11.5 |  | 2 | 8.64 | 88.3 | 88.325 |  | 2.2 | 11.22 | 87.7 | 87.7286 |  | 3 | 19.42 | 86.62 | 86.65 |  | 5.3 | 73.68 | 85.52 | 85.56 | 1.9 | 7.8 | 78.3 | 85.6 | 85.54
11.6 |  | 2 | 8.44 | 88.275 | 88.3 |  | 2.2 | 10.825 | 87.7 | 87.6 |  | 3 | 17.55 | 86.5667 | 86.6 |  | 5.2 | 71.5 | 85.4 | 85.42 | 1.9 | 7.75 | 77.8667 | 85.4333 | 85.3429
11.7 |  | 1.9 | 8.22 | 88.2 | 88.24 |  | 2.2 | 10.44 | 87.6 | 87.5143 |  | 3 | 16.325 | 86.48 | 86.45 |  | 5.1 | 68.26 | 85.25 | 85.3167 |  | 7.7 | 77.5333 | 85.3 | 85.1833
11.8 |  | 1.9 | 8.1 | 88.1 | 88.15 |  | 2.1 | 10.22 | 87.5 | 87.5 |  | 2.9 | 15.475 | 86.4 | 86.3 |  | 5.1 | 62.95 | 85.125 | 85.14 |  | 7.6 | 77.1 | 85.18 | 85
11.9 |  | 1.9 | 8 | 88.04 | 88.1 |  | 2.1 | 9.925 | 87.4 | 87.4 |  | 2.9 | 14.8 | 86.2 | 86.225 |  | 5 | 57.2167 | 84.9 | 85 |  | 7.4 | 76.7 | 84.88 | 84.92
12 |  | 1.9 | 7.9 | 87.95 | 88 |  | 2.1 | 9.6833 | 87.3571 | 87.3 |  | 2.85 | 14.15 | 86.1167 | 86.0571 |  | 4.9 | 50.2167 | 84.8 | 84.82 |  | 7.3 | 76.18 | 84.775 | 84.8
12.1 |  | 1.9 | 7.65 | 87.9 | 87.9 |  | 2.1 | 9.5 | 87.2143 | 87.2 |  | 2.8 | 13.8 | 86.04 | 86 |  | 4.8 | 43.45 | 84.7 | 84.7 |  | 7.2 | 75.5667 | 84.6 | 84.6
12.2 |  |  | 7.54 | 87.8 | 87.8667 |  | 2.1 | 9.2 | 87.2 | 87.0429 |  | 2.8 | 13.35 | 85.92 | 85.9 |  | 4.8 | 37.24 | 84.6 | 84.5 |  | 7.1 | 74.7 | 84.46 | 84.4
12.3 |  |  | 7.4 | 87.8 | 87.8 |  | 2 | 9.1 | 87.1 | 87 |  | 2.7 | 13.025 | 85.8 | 85.8 |  | 4.7 | 31.7667 | 84.38 | 84.4 |  | 7 | 74.05 | 84.3 | 84.3333
12.4 |  |  | 7.275 | 87.76 | 87.7 |  | 2 | 8.8 | 86.9857 | 86.9 |  | 2.7 | 12.675 | 85.7 | 85.62 |  | 4.6 | 27.125 | 84.18 | 84.3 |  | 6.85 | 73.05 | 84.2 | 84.125
12.5 |  |  | 7.1 | 87.7 | 87.68 |  | 2 | 8.66 | 86.9 | 86.8 |  | 2.6 | 12.375 | 85.52 | 85.5 |  | 4.5 | 24.2333 | 84.1 | 84.2 |  | 6.7 | 71.45 | 84 | 84.0333
12.6 |  |  | 6.975 | 87.6 | 87.6 |  |  | 8.52 | 86.7833 | 86.6833 |  | 2.6 | 12.125 | 85.4 | 85.3667 |  | 4.44 | 22.3 | 83.96 | 84 |  | 6.7 | 69.4 | 83.875 | 83.925
12.7 |  |  | 6.875 | 87.5 | 87.52 |  |  | 8.3 | 86.7 | 86.6 |  | 2.6 | 11.9 | 85.3 | 85.2143 |  | 4.4 | 21 | 83.8 | 83.9 |  | 6.6 | 66.3167 | 83.68 | 83.72
12.8 |  |  | 6.8 | 87.4 | 87.5 |  |  | 8.175 | 86.5333 | 86.4429 |  | 2.5 | 11.65 | 85.1667 | 85.0857 |  | 4.3 | 19.98 | 83.64 | 83.7 |  | 6.5 | 62.6333 | 83.5 | 83.5333
12.9 |  |  | 6.7 | 87.32 | 87.375 |  |  | 8 | 86.44 | 86.3857 |  | 2.5 | 11.375 | 85 | 84.92 |  | 4.3 | 19.275 | 83.425 | 83.65 |  | 6.4 | 57.5 | 83.3 | 83.34
13 |  |  | 6.54 | 87.2833 | 87.3 |  |  | 7.8 | 86.3667 | 86.3 |  | 2.5 | 11.14 | 84.9 | 84.8571 |  | 4.2 | 18.78 | 83.3 | 83.4 |  | 6.35 | 51.3 | 83.2 | 83.1333
13.1 |  |  | 6.425 | 87.2 | 87.2 |  |  | 7.7 | 86.2143 | 86.16 |  | 2.5 | 10.9 | 84.72 | 84.7143 |  | 4.2 | 18.28 | 83.2 | 83.275 |  | 6.3 | 45.6 | 83.125 | 82.925
13.2 |  |  | 6.375 | 87.1 | 87.175 |  |  | 7.5 | 86.1333 | 86.1 |  |  | 10.6 | 84.58 | 84.55 |  | 4.15 | 17.82 | 82.95 | 83.075 |  | 6.2 | 39.58 | 82.9667 | 82.725
13.3 |  |  | 6.3 | 87.075 | 87.04 |  |  | 7.4 | 86 | 85.9625 |  |  | 10.45 | 84.46 | 84.3714 |  | 4.075 | 17.4 | 82.8 | 82.9 |  | 6.1 | 34.5667 | 82.7429 | 82.6833
13.4 |  |  | 6.14 | 87 | 87 |  |  | 7.24 | 85.8833 | 85.8167 |  |  | 10.2 | 84.26 | 84.2167 |  | 4 | 17.08 | 82.6 | 82.8 |  | 5.9 | 30.5 | 82.5 | 82.6
13.5 |  |  | 6.1 | 86.9 | 86.9 |  |  | 7.18 | 85.7167 | 85.75 |  |  | 10 | 84.1 | 84.075 |  | 3.9 | 16.75 | 82.45 | 82.65 |  | 5.8 | 27.9 | 82.3286 | 82.4
13.6 |  |  | 6 | 86.8 | 86.8 |  |  | 7.04 | 85.62 | 85.6125 |  |  | 9.8167 | 83.95 | 83.9333 |  | 3.9 | 16.4 | 82.3 | 82.36 |  | 5.7 | 25.7 | 82.125 | 82.2333
13.7 |  |  | 5.9 | 86.7 | 86.7 |  |  | 6.9 | 85.45 | 85.45 |  |  | 9.65 | 83.725 | 83.7714 |  | 3.8 | 16.1 | 82.12 | 82.3 |  | 5.7 | 24.2 | 81.9833 | 82.0667
13.8 |  |  | 5.8 | 86.56 | 86.52 |  |  | 6.8 | 85.3857 | 85.275 |  |  | 9.525 | 83.58 | 83.7 |  | 3.8 | 15.8 | 81.975 | 82.16 |  | 5.7 | 23.26 | 81.8 | 81.9
13.9 |  |  | 5.75 | 86.4 | 86.4 |  |  | 6.7 | 85.2429 | 85.15 |  |  | 9.3 | 83.4667 | 83.42 |  | 3.8 | 15.5333 | 81.75 | 81.9 |  | 5.6 | 22.6333 | 81.6 | 81.66
14 |  |  | 5.65 | 86.375 | 86.275 |  |  | 6.6 | 85.1125 | 85.025 |  |  | 9.16 | 83.25 | 83.24 |  | 3.7 | 15.3 | 81.6333 | 81.7 |  | 5.5 | 21.8833 | 81.5 | 81.55
14.1 |  |  | 5.6 | 86.12 | 86.06 |  |  | 6.5 | 84.88 | 84.84 |  |  | 8.95 | 83.1 | 83 |  | 3.6 | 15 | 81.475 | 81.56 |  | 5.4 | 21.45 | 81.3333 | 81.3833
14.2 |  |  | 5.5 | 85.9 | 85.8333 |  |  | 6.4 | 84.6833 | 84.5833 |  |  | 8.775 | 82.9333 | 82.9167 |  | 3.6 | 14.74 | 81.22 | 81.2833 |  | 5.3 | 20.96 | 81.175 | 81.1333
14.3 |  |  | 5.4 | 85.65 | 85.6 |  |  | 6.2 | 84.5286 | 84.4167 |  |  | 8.66 | 82.7 | 82.725 |  | 3.5 | 14.4 | 80.975 | 81.1 |  | 5.3 | 20.55 | 81 | 80.95
14.4 |  |  | 5.4 | 85.4 | 85.25 |  |  | 6.2 | 84.2667 | 84.22 |  |  | 8.5 | 82.5 | 82.4333 |  |  | 14.26 | 80.7833 | 81 |  | 5.2 | 20.2 | 80.8 | 80.7333
14.5 |  |  | 5.3 | 84.78 | 84.72 |  |  | 6.0667 | 84.0667 | 83.8857 |  |  | 8.4 | 82.2667 | 82.1625 |  |  | 14.025 | 80.6 | 80.74 |  | 5.2 | 19.96 | 80.5857 | 80.575
14.6 |  |  | 5.2 | 84.06 | 84 |  |  | 5.98 | 83.7571 | 83.7 |  |  | 8.2667 | 82.1333 | 82.0167 |  |  | 13.7 | 80.44 | 80.4667 |  | 5.2 | 19.5 | 80.4 | 80.44
14.7 |  |  | 5.12 | 82.9 | 82.92 |  |  | 5.88 | 83.4143 | 83.42 |  |  | 8.1 | 81.875 | 81.7889 |  |  | 13.5 | 80.2333 | 80.22 |  | 5.1 | 19.26 | 80.24 | 80.225
14.8 |  |  | 5.1 | 80.9 | 80.82 |  |  | 5.8 | 83.1 | 82.95 |  |  | 8 | 81.625 | 81.6667 |  |  | 13.32 | 79.975 | 80.1 |  | 5.1 | 19 | 80 | 80.05
14.9 |  |  | 5 | 77 | 77.2 |  |  | 5.7 | 82.4667 | 82.45 |  |  | 7.8 | 81.4 | 81.2833 |  |  | 13.05 | 79.84 | 79.96 |  | 5.075 | 18.6667 | 79.88 | 79.9
15 |  |  | 4.975 | 70.65 | 70.9 |  |  | 5.625 | 81.96 | 81.8429 |  |  | 7.7 | 81.125 | 81 |  |  | 12.8 | 79.64 | 79.675 |  | 4.9 | 18.3667 | 79.6 | 79.64
15.1 |  |  | 4.9 | 62.26 | 62.9 |  |  | 5.6 | 81.1 | 81 |  |  | 7.5667 | 80.85 | 80.7667 |  |  | 12.6 | 79.32 | 79.5 |  | 4.9 | 18.12 | 79.5 | 79.5
15.2 |  |  | 4.875 | 53.92 | 53.8 |  |  | 5.5 | 79.9 | 79.7571 |  |  | 7.4 | 80.48 | 80.4429 |  |  | 12.45 | 79.1 | 79.3 |  | 4.9 | 17.82 | 79.375 | 79.2333
15.3 |  |  | 4.8 | 43.66 | 44.2 |  |  | 5.44 | 77.8857 | 77.8167 |  |  | 7.3 | 80.2 | 80.1333 |  |  | 12.2 | 78.8 | 78.9 |  | 4.725 | 17.54 | 79.0833 | 78.96
15.4 |  |  | 4.7 | 35 | 35.9333 |  |  | 5.4 | 74.5667 | 74.64 |  |  | 7.2 | 79.8667 | 79.8 |  |  | 11.98 | 78.6 | 78.75 |  | 4.675 | 17.32 | 78.9 | 78.75
15.5 |  |  | 4.7 | 26.95 | 28.3 |  |  | 5.3 | 69.5429 | 69.8 |  |  | 7.1 | 79.5 | 79.4 |  |  | 11.8 | 78.44 | 78.5 |  | 4.6 | 17.1 | 78.75 | 78.5714
15.6 |  |  | 4.6 | 20.6 | 21.9 |  |  | 5.2 | 62.45 | 62.4 |  |  | 7 | 78.98 | 78.9429 |  |  | 11.7 | 78.22 | 78.25 |  |  | 16.86 | 78.55 | 78.3333
15.7 |  |  | 4.6 | 17.05 | 18.325 |  |  | 5.18 | 54.3 | 54.52 |  |  | 6.9 | 78.3 | 78.36 |  |  | 11.5 | 77.9333 | 77.98 |  |  | 16.625 | 78.26 | 78.15
15.8 |  |  | 4.6 | 15.15 | 16.3 |  |  | 5.1 | 45.94 | 46.4571 |  |  | 6.8 | 77.66 | 77.55 |  |  | 11.3 | 77.7 | 77.76 |  |  | 16.44 | 77.9667 | 77.95
15.9 |  |  | 4.5 | 13.925 | 15.075 |  |  | 5 | 37.55 | 38.0429 |  |  | 6.7 | 76.6 | 76.8 |  |  | 11.1571 | 77.4 | 77.45 |  |  | 16.1 | 77.78 | 77.76
16 |  |  | 4.5 | 13.1 | 14.36 |  |  | 5 | 29.9571 | 30.9 |  |  | 6.6 | 75.18 | 75.2143 |  |  | 11.075 | 77.2 | 77.3 |  |  | 15.88 | 77.5 | 77.56
16.1 |  |  | 4.42 | 12.6667 | 13.86 |  |  | 4.9 | 23.9 | 25.2375 |  |  | 6.5 | 72.92 | 73.1 |  |  | 10.8333 | 76.9667 | 76.95 |  |  | 15.6 | 77.28 | 77.2667
16.2 |  |  | 4.4 | 12.25 | 13.4333 |  |  | 4.9 | 20.0667 | 21.05 |  |  | 6.44 | 69.5 | 69.54 |  |  | 10.68 | 76.65 | 76.72 |  |  | 15.42 | 77.1 | 77.06
16.3 |  |  | 4.4 | 11.95 | 13.1333 |  |  | 4.8 | 17.6857 | 18.7286 |  |  | 6.3 | 64.15 | 64.3667 |  |  | 10.46 | 76.32 | 76.425 |  |  | 15.2 | 76.86 | 76.94
16.4 |  |  | 4.3333 | 11.7 | 12.9 |  |  | 4.8 | 16.2714 | 17.3 |  |  | 6.2 | 57.16 | 58.06 |  |  | 10.325 | 76.1 | 76.2 |  |  | 15.1 | 76.66 | 76.65
16.5 |  |  | 4.3 | 11.46 | 12.66 |  |  | 4.7 | 15.3429 | 16.4 |  |  | 6.1167 | 50.425 | 50.7 |  |  | 10.2 | 75.76 | 75.75 |  |  | 14.76 | 76.425 | 76.38
16.6 |  |  | 4.3 | 11.26 | 12.5 |  |  | 4.7 | 14.5857 | 15.76 |  |  | 6.1 | 43.04 | 43.08 |  |  | 9.9833 | 75.44 | 75.4 |  |  | 14.54 | 76.2 | 76.16
16.7 |  |  | 4.2 | 11.1 | 12.3333 |  |  | 4.6 | 14.1333 | 15.2286 |  |  | 6.0143 | 35.5 | 36.44 |  |  | 9.8 | 75.02 | 75.0667 |  |  | 14.4 | 75.9667 | 75.86
16.8 |  |  | 4.2 | 11 | 12.2 |  |  | 4.6 | 13.6714 | 14.8 |  |  | 5.95 | 29.475 | 30.68 |  |  | 9.7 | 74.66 | 74.66 |  |  | 14.2143 | 75.6667 | 75.65
16.9 |  |  | 4.1 | 10.84 | 12 |  |  | 4.575 | 13.3667 | 14.46 |  |  | 5.9 | 24.925 | 25.8 |  |  | 9.6 | 74.3 | 74.3 |  |  | 14 | 75.45 | 75.45
17 |  |  | 4.1 | 10.7 | 11.9 |  |  | 4.5 | 13 | 14.2 |  |  | 5.8 | 21.9 | 22.9167 |  |  | 9.48 | 73.8 | 73.85 |  |  | 13.8167 | 75.15 | 75.18
17.1 |  |  | 4 | 10.6 | 11.8 |  |  | 4.5 | 12.7833 | 13.95 |  |  | 5.7 | 19.925 | 21.1667 |  |  | 9.36 | 73.35 | 73.34 |  |  | 13.6 | 74.9 | 74.84
17.2 |  |  | 4 | 10.48 | 11.65 |  |  | 4.4167 | 12.5 | 13.68 |  |  | 5.7 | 18.775 | 19.9 |  |  | 9.22 | 72.8 | 72.78 |  |  | 13.46 | 74.6667 | 74.48
17.3 |  |  |  | 10.3667 | 11.54 |  |  | 4.3 | 12.32 | 13.4833 |  |  | 5.6 | 17.925 | 19.1 |  |  | 9.15 | 72 | 72 |  |  | 13.2 | 74.32 | 74.24
17.4 |  |  |  | 10.225 | 11.4 |  |  | 4.3 | 12.075 | 13.3167 |  |  | 5.5667 | 17.225 | 18.4143 |  |  | 8.9667 | 71.08 | 71.14 |  |  | 13.0571 | 74.0333 | 74.04
17.5 |  |  |  | 10.1 | 11.35 |  |  | 4.3 | 11.9333 | 13.14 |  |  | 5.475 | 16.7333 | 18 |  |  | 8.8 | 69.6333 | 69.66 |  |  | 12.9 | 73.75 | 73.75
17.6 |  |  |  | 10 | 11.22 |  |  | 4.2 | 11.76 | 12.8857 |  |  | 5.4 | 16.25 | 17.5667 |  |  | 8.7 | 67.575 | 67.8 |  |  | 12.74 | 73.4667 | 73.5
17.7 |  |  |  | 9.9 | 11.1 |  |  | 4.2 | 11.5833 | 12.7429 |  |  | 5.35 | 15.94 | 17.2 |  |  | 8.6 | 64.5 | 64.78 |  |  | 12.56 | 73.2 | 73.1667
17.8 |  |  |  | 9.8 | 11 |  |  | 4.2 | 11.4 | 12.58 |  |  | 5.3 | 15.6 | 16.86 |  |  | 8.525 | 60.1667 | 60.64 |  |  | 12.36 | 72.92 | 72.85
17.9 |  |  |  | 9.7 | 10.86 |  |  |  | 11.2333 | 12.3833 |  |  | 5.2 | 15.24 | 16.48 |  |  | 8.45 | 55.075 | 55.42 |  |  | 12.225 | 72.5 | 72.5667
18 |  |  |  | 9.6 | 10.7333 |  |  |  | 11.06 | 12.2 |  |  | 5.2 | 14.9 | 16.1833 |  |  | 8.28 | 48.7 | 49.6667 |  |  | 12.1 | 72.1833 | 72.24
18.1 |  |  |  | 9.5 | 10.68 |  |  |  | 10.9857 | 12.1333 |  |  | 5.1 | 14.68 | 15.9 |  |  | 8.1833 | 43.1 | 43.65 |  |  | 11.875 | 72 | 71.8667
18.2 |  |  |  | 9.4 | 10.6 |  |  |  | 10.82 | 11.9833 |  |  | 5.1 | 14.5 | 15.6667 |  |  | 8.075 | 37.2167 | 38.025 |  |  | 11.75 | 71.65 | 71.4
18.3 |  |  |  | 9.3 | 10.5 |  |  |  | 10.64 | 11.8 |  |  | 5 | 14.2 | 15.475 |  |  | 7.9833 | 31.8 | 33.15 |  |  | 11.64 | 71.1 | 71.06
18.4 |  |  |  | 9.2 | 10.35 |  |  |  | 10.48 | 11.7 |  |  | 5 | 14.05 | 15.2 |  |  | 7.8167 | 28.12 | 29.4 |  |  | 11.5 | 70.78 | 70.66
18.5 |  |  |  | 9.1 | 10.275 |  |  |  | 10.4 | 11.5333 |  |  | 4.92 | 13.74 | 15.0333 |  |  | 7.7 | 25.7 | 26.775 |  |  | 11.3 | 70.25 | 70.25
18.6 |  |  |  | 9 | 10.1 |  |  |  | 10.2333 | 11.44 |  |  | 4.875 | 13.54 | 14.74 |  |  | 7.7 | 23.86 | 25.075 |  |  | 11.2 | 69.9 | 69.8
18.7 |  |  |  | 8.9 | 10.05 |  |  |  | 10.15 | 11.3 |  |  |  | 13.4 | 14.56 |  |  | 7.6 | 22.7 | 24 |  |  | 11.1 | 69.25 | 69.3
18.8 |  |  |  | 8.8 | 10 |  |  |  | 10 | 11.16 |  |  |  | 13.2 | 14.4 |  |  | 7.52 | 21.86 | 23.3 |  |  | 10.9 | 68.44 | 68.5
18.9 |  |  |  | 8.7 | 9.875 |  |  |  | 9.86 | 11.0125 |  |  |  | 12.975 | 14.2333 |  |  | 7.475 | 21.26 | 22.6 |  |  | 10.8 | 67.6 | 67.8
19 |  |  |  | 8.625 | 9.8 |  |  |  | 9.7667 | 10.94 |  |  |  | 12.775 | 14.0333 |  |  | 7.4 | 20.675 | 22.1 |  |  | 10.7 | 66.3333 | 66.5
19.1 |  |  |  | 8.5 | 9.7 |  |  |  | 9.6 | 10.7667 |  |  |  | 12.56 | 13.7667 |  |  | 7.3 | 20.3 | 21.65 |  |  | 10.5 | 64.54 | 64.6
19.2 |  |  |  | 8.425 | 9.6 |  |  |  | 9.5167 | 10.6 |  |  |  | 12.36 | 13.58 |  |  | 7.2 | 19.8333 | 21.22 |  |  | 10.36 | 62 | 62.28
19.3 |  |  |  | 8.3 | 9.46 |  |  |  | 9.4 | 10.55 |  |  |  | 12.2 | 13.4 |  |  | 7.1 | 19.575 | 20.825 |  |  | 10.3 | 58.14 | 58.8
19.4 |  |  |  | 8.2 | 9.4 |  |  |  | 9.3 | 10.4 |  |  |  | 12.02 | 13.3 |  |  | 7.06 | 19.25 | 20.54 |  |  | 10.2 | 54 | 54.5
19.5 |  |  |  | 8.2 | 9.3 |  |  |  | 9.1333 | 10.3333 |  |  |  | 11.9 | 13.1333 |  |  | 6.9 | 18.84 | 20.25 |  |  | 10.1 | 49.32 | 49.95
19.6 |  |  |  | 8 | 9.2 |  |  |  | 9.06 | 10.2 |  |  |  | 11.8 | 12.94 |  |  | 6.85 | 18.6 | 19.85 |  |  | 10 | 43.84 | 45
19.7 |  |  |  | 8 | 9.1 |  |  |  | 8.9833 | 10.1 |  |  |  | 11.54 | 12.8 |  |  | 6.72 | 18.26 | 19.675 |  |  | 9.9 | 38.7667 | 40.2667
19.8 |  |  |  | 7.9 | 9 |  |  |  | 8.9 | 10 |  |  |  | 11.4 | 12.6833 |  |  | 6.7 | 18.025 | 19.34 |  |  | 9.7333 | 34.8 | 36
19.9 |  |  |  | 7.8167 | 8.9167 |  |  |  | 8.7167 | 9.85 |  |  |  | 11.225 | 12.5286 |  |  | 6.7 | 17.775 | 19 |  |  | 9.625 | 31.1667 | 32.44
20 |  |  |  | 7.6833 | 8.82 |  |  |  | 8.6571 | 9.7833 |  |  |  | 11.1 | 12.34 |  |  | 6.6 | 17.5 | 18.86 |  |  | 9.525 | 28.5 | 29.9
20.1 |  |  |  | 7.6 | 8.7 |  |  |  | 8.58 | 9.7 |  |  |  | 11 | 12.14 |  |  | 6.5 | 17.3167 | 18.55 |  |  | 9.5 | 26.7 | 28.1
20.2 |  |  |  | 7.5 | 8.65 |  |  |  | 8.5 | 9.5857 |  |  |  | 10.86 | 12 |  |  | 6.4 | 17 | 18.3 |  |  | 9.375 | 25.6 | 26.85
20.3 |  |  |  | 7.4 | 8.54 |  |  |  | 8.3167 | 9.5 |  |  |  | 10.7 | 11.8833 |  |  | 6.3 | 16.78 | 18.1333 |  |  | 9.24 | 24.65 | 25.98
20.4 |  |  |  | 7.36 | 8.4571 |  |  |  | 8.2875 | 9.3714 |  |  |  | 10.56 | 11.8 |  |  |  | 16.575 | 17.85 |  |  | 9.2 | 24 | 25.35
20.5 |  |  |  | 7.2 | 8.3 |  |  |  | 8.1571 | 9.2375 |  |  |  | 10.4 | 11.6429 |  |  |  | 16.3667 | 17.62 |  |  | 9.1 | 23.5429 | 24.9
20.6 |  |  |  | 7.175 | 8.225 |  |  |  | 8.1 | 9.2 |  |  |  | 10.3 | 11.46 |  |  |  | 16.15 | 17.4 |  |  | 9 | 23.0857 | 24.46
20.7 |  |  |  | 7 | 8.1 |  |  |  | 7.9333 | 9.0571 |  |  |  | 10.2 | 11.3 |  |  |  | 15.9 | 17.1333 |  |  | 8.9 | 22.6667 | 24
20.8 |  |  |  | 6.9 | 8 |  |  |  | 7.9 | 8.925 |  |  |  | 10.075 | 11.2 |  |  |  | 15.7 | 17 |  |  | 8.8 | 22.32 | 23.8
20.9 |  |  |  | 6.8 | 7.9 |  |  |  | 7.7667 | 8.9 |  |  |  | 9.9 | 11.1 |  |  |  | 15.52 | 16.75 |  |  | 8.7 | 22.1 | 23.4333
21 |  |  |  | 6.775 | 7.9 |  |  |  | 7.625 | 8.7 |  |  |  | 9.8 | 10.9167 |  |  |  | 15.3 | 16.6 |  |  | 8.6 | 21.66 | 23.1
21.1 |  |  |  | 6.7 | 7.8 |  |  |  | 7.5667 | 8.6375 |  |  |  | 9.65 | 10.8429 |  |  |  | 15.05 | 16.375 |  |  | 8.5 | 21.35 | 22.76
21.2 |  |  |  | 6.6 | 7.7 |  |  |  | 7.5 | 8.58 |  |  |  | 9.6 | 10.7 |  |  |  | 14.825 | 16.14 |  |  | 8.35 | 21.12 | 22.475
21.3 |  |  |  | 6.5 | 7.6 |  |  |  | 7.4 | 8.5 |  |  |  | 9.48 | 10.6333 |  |  |  | 14.725 | 15.92 |  |  | 8.3 | 20.8571 | 22.15
21.4 |  |  |  | 6.4833 | 7.55 |  |  |  | 7.3125 | 8.3571 |  |  |  | 9.38 | 10.4667 |  |  |  | 14.475 | 15.8 |  |  | 8.2 | 20.5 | 21.92
21.5 |  |  |  | 6.4 | 7.46 |  |  |  | 7.1571 | 8.2 |  |  |  | 9.2 | 10.4 |  |  |  | 14.325 | 15.58 |  |  | 8.1 | 20.3 | 21.6
21.6 |  |  |  | 6.3 | 7.4 |  |  |  | 7.1 | 8.1 |  |  |  | 9.1 | 10.2333 |  |  |  | 14.1 | 15.375 |  |  | 8.08 | 20 | 21.46
21.7 |  |  |  | 6.275 | 7.3 |  |  |  | 7 | 8.0667 |  |  |  | 9 | 10.1 |  |  |  | 13.9667 | 15.22 |  |  | 7.9667 | 19.8 | 21.125
21.8 |  |  |  | 6.14 | 7.2167 |  |  |  | 6.8875 | 7.95 |  |  |  | 8.875 | 10.0333 |  |  |  | 13.8 | 15.02 |  |  | 7.9 | 19.625 | 20.92
21.9 |  |  |  | 6.1 | 7.2 |  |  |  | 6.8 | 7.9 |  |  |  | 8.74 | 9.9 |  |  |  | 13.6667 | 14.8 |  |  | 7.8 | 19.3 | 20.6333
22 |  |  |  | 6 | 7.1 |  |  |  | 6.7 | 7.8 |  |  |  | 8.64 | 9.7571 |  |  |  | 13.5 | 14.675 |  |  | 7.7 | 19.05 | 20.5
22.1 |  |  |  | 6 | 7 |  |  |  | 6.6667 | 7.7 |  |  |  | 8.5 | 9.7 |  |  |  | 13.3 | 14.52 |  |  |  | 18.8 | 20.1
22.2 |  |  |  | 5.9 | 7 |  |  |  | 6.6 | 7.6143 |  |  |  | 8.4 | 9.5429 |  |  |  | 13.1 | 14.325 |  |  |  | 18.625 | 19.9
22.3 |  |  |  | 5.9 | 6.9 |  |  |  | 6.5 | 7.5 |  |  |  | 8.32 | 9.4125 |  |  |  | 12.98 | 14.175 |  |  |  | 18.3 | 19.7
22.4 | [8 empty] | 6.4 | 7.5 |  |  |  | 8.2 | 9.2875 |  |  |  | 12.775 | 14 |  |  |  | 18.18 | 19.46
22.5 | [8 empty] | 6.36 | 7.4 |  |  |  | 8 | 9.1625 |  |  |  | 12.6 | 13.85 |  |  |  | 17.96 | 19.25
22.6 | [8 empty] | 6.3 | 7.3 |  |  |  | 7.9 | 9.0286 |  |  |  | 12.5 | 13.7 |  |  |  | 17.725 | 19.04
22.7 | [8 empty] | 6.2 | 7.2429 |  |  |  | 7.9 | 8.975 |  |  |  | 12.4 | 13.5 |  |  |  | 17.54 | 18.9
22.8 | [8 empty] | 6.2 | 7.2 |  |  |  | 7.74 | 8.85 |  |  |  | 12.2 | 13.3667 |  |  |  | 17.34 | 18.675
22.9 | [8 empty] | 6.1 | 7.1 |  |  |  | 7.64 | 8.8 |  |  |  | 12.04 | 13.2 |  |  |  | 17.1 | 18.4667

## Wedge 15 Longitudinal Profiles
Longitudinal Profiles
Wedge 15
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-34.9 |  |  |  |  | 11.7
-34.8 |  |  |  |  | 11.7
-34.7 |  |  |  |  | 11.8111
-34.6 |  |  |  |  | 11.92
-34.5 |  |  |  |  | 12.02
-34.4 |  |  |  |  | 12.12
-34.3 |  |  |  |  | 12.22
-34.2 |  |  |  |  | 12.32
-34.1 |  |  |  |  | 12.4286
-34 |  |  |  |  | 12.5556
-33.9 |  |  |  |  | 12.6545
-33.8 |  |  |  |  | 12.75
-33.7 |  |  |  |  | 12.85
-33.6 |  |  |  |  | 12.95
-33.5 |  |  |  |  | 13.0556
-33.4 |  |  |  |  | 13.25
-33.3 |  |  |  |  | 13.4
-33.2 |  |  |  |  | 13.5125
-33.1 |  |  |  |  | 13.6333
-33 |  |  |  |  | 13.74
-32.9 |  |  |  |  | 13.84
-32.8 |  |  |  |  | 13.94
-32.7 |  |  |  |  | 14.04
-32.6 |  |  |  |  | 14.1444
-32.5 |  |  |  |  | 14.2625
-32.4 |  |  |  |  | 14.44
-32.3 |  |  |  | 10.0143 | 14.57
-32.2 |  |  |  | 10.1 | 14.67
-32.1 |  |  |  | 10.2 | 14.77
-32 |  |  |  | 10.28 | 14.9273
-31.9 |  |  |  | 10.4 | 15.12
-31.8 |  |  |  | 10.4 | 15.26
-31.7 |  |  |  | 10.5125 | 15.36
-31.6 |  |  |  | 10.6 | 15.5333
-31.5 |  |  |  | 10.6714 | 15.6778
-31.4 |  |  |  | 10.8333 | 15.8143
-31.3 |  |  |  | 10.9857 | 15.9571
-31.2 |  |  |  | 11.1 | 16.1556
-31.1 |  |  |  | 11.15 | 16.28
-31 |  |  |  | 11.2 | 16.4125
-30.9 |  |  |  | 11.3429 | 16.5429
-30.8 |  |  |  | 11.475 | 16.6857
-30.7 |  |  |  | 11.6 | 16.9
-30.6 |  |  |  | 11.7 | 17.1429
-30.5 |  |  |  | 11.7857 | 17.26
-30.4 |  |  |  | 11.9125 | 17.42
-30.3 |  |  |  | 12 | 17.575
-30.2 |  |  |  | 12.1125 | 17.76
-30.1 |  |  |  | 12.2375 | 17.96
-30 |  |  |  | 12.3625 | 18.0727
-29.9 |  |  |  | 12.4875 | 18.24
-29.8 |  |  |  | 12.6167 | 18.4556
-29.7 |  |  |  | 12.7714 | 18.5889
-29.6 |  |  | 8.4 | 12.8333 | 18.8455
-29.5 |  |  | 8.4833 | 13.0714 | 19.0778
-29.4 |  |  | 8.5 | 13.1333 | 19.2571
-29.3 |  |  | 8.6125 | 13.3 | 19.4625
-29.2 |  |  | 8.7 | 13.4 | 19.6556
-29.1 |  |  | 8.7714 | 13.55 | 19.8286
-29 |  |  | 8.84 | 13.7167 | 20.0667
-28.9 |  |  | 8.9 | 13.8714 | 20.4125
-28.8 |  |  | 8.9375 | 14.0167 | 20.5429
-28.7 |  |  | 9.0714 | 14.1833 | 20.7714
-28.6 |  |  | 9.1333 | 14.3143 | 21
-28.5 |  |  | 9.2143 | 14.5143 | 21.2286
-28.4 |  |  | 9.35 | 14.7 | 21.4429
-28.3 |  | 7.6 | 9.4 | 14.85 | 21.7286
-28.2 |  | 7.7 | 9.5 | 14.92 | 22
-28.1 |  | 7.7 | 9.5875 | 15.1714 | 22.25
-28 |  | 7.75 | 9.7 | 15.3333 | 22.5
-27.9 |  | 7.8 | 9.7875 | 15.5 | 22.8444
-27.8 |  | 7.9 | 9.9 | 15.6667 | 23.075
-27.7 |  | 7.9875 | 9.9667 | 15.94 | 23.3111
-27.6 |  | 8 | 10.1 | 16.12 | 23.6
-27.5 |  | 8 | 10.2 | 16.3167 | 23.9333
-27.4 |  | 8.1 | 10.3 | 16.4714 | 24.2667
-27.3 | 7.3286 | 8.2 | 10.4 | 16.6429 | 24.57
-27.2 | 7.4 | 8.2714 | 10.4833 | 16.9286 | 24.9
-27.1 | 7.5 | 8.3143 | 10.6286 | 17.1 | 25.3286
-27 | 7.5 | 8.4 | 10.7714 | 17.375 | 25.8
-26.9 | 7.5571 | 8.42 | 10.9143 | 17.5833 | 26.2714
-26.8 | 7.6 | 8.5 | 11 | 17.7429 | 26.95
-26.7 | 7.6833 | 8.6286 | 11.1143 | 18.0571 | 27.8
-26.6 | 7.7 | 8.7 | 11.2667 | 18.2857 | 28.85
-26.5 | 7.8 | 8.8 | 11.4 | 18.55 | 30.275
-26.4 | 7.8 | 8.9 | 11.5 | 18.7 | 32.3778
-26.3 | 7.9 | 8.9714 | 11.7 | 18.95 | 35.12
-26.2 | 7.9 | 9.1 | 11.7714 | 19.2 | 38.2
-26.1 | 8.0143 | 9.1 | 11.92 | 19.45 | 41.9286
-26 | 8.1 | 9.2 | 12.1 | 19.8 | 46.2143
-25.9 | 8.15 | 9.35 | 12.2429 | 20.0286 | 50.825
-25.8 | 8.2 | 9.4 | 12.3857 | 20.3333 | 55.6
-25.7 | 8.2625 | 9.5 | 12.5333 | 20.5778 | 59.8
-25.6 | 8.3 | 9.6429 | 12.6857 | 20.9375 | 63.2545
-25.5 | 8.4 | 9.775 | 12.85 | 21.3 | 66.6
-25.4 | 8.5 | 9.825 | 13.0125 | 21.575 | 68.8
-25.3 | 8.5286 | 9.9667 | 13.1429 | 21.9375 | 70.5222
-25.2 | 8.6625 | 10 | 13.3714 | 22.3125 | 71.8333
-25.1 | 8.7 | 10.1571 | 13.5286 | 22.75 | 72.9
-25 | 8.8 | 10.3 | 13.7667 | 23.2111 | 73.6636
-24.9 | 8.9 | 10.3667 | 13.9333 | 23.7 | 74.3
-24.8 | 8.98 | 10.5 | 14.1714 | 24.45 | 74.8
-24.7 | 9.0429 | 10.5833 | 14.35 | 25.35 | 75.3
-24.6 | 9.1143 | 10.7375 | 14.6 | 26.625 | 75.76
-24.5 | 9.2 | 10.8833 | 14.7667 | 28.3286 | 76.08
-24.4 | 9.3571 | 10.98 | 14.9667 | 30.9 | 76.5
-24.3 | 9.4286 | 11.1667 | 15.1857 | 34.3 | 76.9091
-24.2 | 9.5 | 11.32 | 15.46 | 38.2 | 77.19
-24.1 | 9.6 | 11.4 | 15.68 | 43.3333 | 77.52
-24 | 9.7 | 11.5571 | 15.9286 | 48.42 | 77.89
-23.9 | 9.7667 | 11.7167 | 16.2143 | 53.8 | 78.16
-23.8 | 9.9 | 11.8833 | 16.4333 | 59.4667 | 78.4
-23.7 | 10 | 12.08 | 16.7 | 64.4667 | 78.6571
-23.6 | 10.1 | 12.175 | 16.9667 | 68.6 | 78.9333
-23.5 | 10.2 | 12.3429 | 17.3 | 71.7 | 79.2
-23.4 | 10.28 | 12.5714 | 17.575 | 74.3 | 79.575
-23.3 | 10.45 | 12.7333 | 17.9429 | 76.1 | 79.88
-23.2 | 10.5 | 12.9 | 18.35 | 77.4333 | 80.15
-23.1 | 10.6429 | 13.1143 | 18.725 | 78.4 | 80.4333
-23 | 10.8 | 13.3 | 19.1667 | 79.1143 | 80.675
-22.9 | 10.9833 | 13.55 | 19.8333 | 79.7 | 80.9
-22.8 | 11.1 | 13.74 | 20.4714 | 80.25 | 81.1222
-22.7 | 11.1857 | 13.975 | 21.4 | 80.6857 | 81.36
-22.6 | 11.3286 | 14.24 | 22.6667 | 81.15 | 81.66
-22.5 | 11.4833 | 14.5 | 24.7143 | 81.5625 | 81.9364
-22.4 | 11.6286 | 14.725 | 27.9 | 81.875 | 82.14
-22.3 | 11.8 | 15 | 31.9 | 82.2143 | 82.51
-22.2 | 11.98 | 15.3667 | 37 | 82.6 | 82.63
-22.1 | 12.14 | 15.75 | 43.3 | 82.85 | 82.9182
-22 | 12.3167 | 16.1667 | 50 | 83.15 | 83.1
-21.9 | 12.4714 | 16.6714 | 57.3667 | 83.525 | 83.3
-21.8 | 12.6333 | 17.4 | 64.8 | 83.8 | 83.6
-21.7 | 12.9143 | 18.4 | 71.4 | 83.9667 | 83.7
-21.6 | 13.0667 | 19.9 | 76.2286 | 84.2714 | 83.9
-21.5 | 13.3167 | 22.2 | 79.95 | 84.625 | 84.2
-21.4 | 13.6 | 25.6286 | 82.6 | 84.8857 | 84.4
-21.3 | 13.92 | 30.6 | 84.4286 | 85.1714 | 84.7
-21.2 | 14.4143 | 36.9143 | 85.6429 | 85.425 | 84.8818
-21.1 | 15.075 | 44.2 | 86.6 | 85.675 | 85.08
-21 | 15.9 | 52.1 | 87.3286 | 85.8833 | 85.19
-20.9 | 17.3 | 60.3 | 87.9 | 86.1667 | 85.38
-20.8 | 19.7429 | 68.4143 | 88.3857 | 86.45 | 85.58
-20.7 | 23.7 | 76.0333 | 88.7667 | 86.6667 | 85.87
-20.6 | 29.2857 | 81.92 | 89.2286 | 86.9 | 85.99
-20.5 | 36.2714 | 86 | 89.6 | 87.05 | 86.18
-20.4 | 44.1286 | 88.4833 | 89.86 | 87.3 | 86.38
-20.3 | 53.65 | 90.175 | 90.08 | 87.5333 | 86.5636
-20.2 | 63.3 | 91.3714 | 90.4 | 87.6875 | 86.76
-20.1 | 72.1 | 92.12 | 90.6429 | 87.925 | 86.88
-20 | 80.4 | 92.7667 | 90.9571 | 88.1429 | 87.06
-19.9 | 86.1667 | 93.3571 | 91.1286 | 88.3714 | 87.3143
-19.8 | 90.1 | 93.7 | 91.3667 | 88.56 | 87.48
-19.7 | 92.9 | 93.96 | 91.62 | 88.7286 | 87.65
-19.6 | 94.46 | 94.2857 | 91.825 | 88.925 | 87.8714
-19.5 | 95.4 | 94.56 | 92.0375 | 89.1 | 88
-19.4 | 96.05 | 94.8167 | 92.2667 | 89.2375 | 88.1125
-19.3 | 96.4667 | 95.0667 | 92.4375 | 89.425 | 88.2375
-19.2 | 96.82 | 95.2286 | 92.6667 | 89.62 | 88.4
-19.1 | 97.0571 | 95.4 | 92.8375 | 89.8 | 88.5909
-19 | 97.2833 | 95.6 | 93.0667 | 89.9286 | 88.7444
-18.9 | 97.4429 | 95.7167 | 93.2333 | 90.1667 | 88.8556
-18.8 | 97.5857 | 95.8714 | 93.375 | 90.3286 | 89.02
-18.7 | 97.74 | 96 | 93.5 | 90.4625 | 89.22
-18.6 | 97.9 | 96.1167 | 93.6857 | 90.625 | 89.4333
-18.5 | 97.9714 | 96.2833 | 93.9 | 90.8 | 89.5875
-18.4 | 98.1 | 96.4 | 93.9857 | 90.9 | 89.6286
-18.3 | 98.1714 | 96.5333 | 94.1 | 91.0667 | 89.7556
-18.2 | 98.2333 | 96.6 | 94.3429 | 91.2571 | 89.92
-18.1 | 98.3 | 96.7571 | 94.4857 | 91.4 | 90.075
-18 | 98.3833 | 96.9 | 94.6167 | 91.5 | 90.2
-17.9 | 98.55 | 97 | 94.7833 | 91.6143 | 90.31
-17.8 | 98.6 | 97.1167 | 94.9 | 91.75 | 90.4125
-17.7 | 98.6833 | 97.2833 | 94.9857 | 91.8667 | 90.5333
-17.6 | 98.76 | 97.35 | 95.1 | 92 | 90.64
-17.5 | 98.8143 | 97.4143 | 95.15 | 92.1667 | 90.78
-17.4 | 98.9 | 97.5667 | 95.2857 | 92.325 | 90.9444
-17.3 | 99 | 97.6 | 95.4333 | 92.5 | 91.0556
-17.2 | 99 | 97.7 | 95.6 | 92.6 | 91.22
-17.1 | 99.1 | 97.8 | 95.7 | 92.7375 | 91.36
-17 | 99.15 | 97.9 | 95.7857 | 92.8625 | 91.5
-16.9 | 99.2 | 97.9857 | 95.84 | 93.0125 | 91.6222
-16.8 | 99.3 | 98.0667 | 96.0143 | 93.1 | 91.7375
-16.7 | 99.3833 | 98.1167 | 96.15 | 93.2 | 91.8625
-16.6 | 99.4 | 98.2 | 96.2 | 93.3 | 92
-16.5 | 99.475 | 98.3 | 96.3 | 93.4 | 92.14
-16.4 | 99.5167 | 98.3 | 96.3857 | 93.5 | 92.32
-16.3 | 99.6 | 98.4 | 96.5 | 93.6 | 92.42
-16.2 | 99.6 | 98.5 | 96.5571 | 93.725 | 92.5
-16.1 | 99.7 | 98.6 | 96.7 | 93.8571 | 92.53
-16 | 99.7 | 98.6625 | 96.7714 | 93.9875 | 92.63
-15.9 | 99.8 | 98.7 | 96.9 | 94.1125 | 92.7429
-15.8 | 99.8 | 98.7429 | 96.9143 | 94.2375 | 92.8667
-15.7 | 99.8 | 98.8 | 97.0571 | 94.3 | 92.9875
-15.6 | 99.9 | 98.9 | 97.1 | 94.4 | 93.1429
-15.5 | 99.9333 | 98.95 | 97.2 | 94.4875 | 93.2
-15.4 | 100 | 99 | 97.2 | 94.6 | 93.2
-15.3 | 100 | 99 | 97.3 | 94.6857 | 93.4333
-15.2 | 100.0333 | 99.1 | 97.4 | 94.8 | 93.5444
-15.1 | 100.1 | 99.1 | 97.45 | 94.8714 | 93.6556
-15 | 100.1 | 99.1 | 97.5 | 94.9286 | 93.76
-14.9 | 100.1 | 99.2 | 97.5889 | 95.0833 | 93.92
-14.8 | 100.2 | 99.2 | 97.6143 | 95.1286 | 94.0857
-14.7 | 100.2 | 99.3 | 97.7 | 95.2625 | 94.1375
-14.6 | 100.25 | 99.3333 | 97.76 | 95.3875 | 94.2
-14.5 | 100.3 | 99.4 | 97.8 | 95.5 | 94.27
-14.4 | 100.3 | 99.4 | 97.9 | 95.56 | 94.3
-14.3 | 100.3143 | 99.4 | 97.9 | 95.6333 | 94.4
-14.2 | 100.4 | 99.4 | 98 | 95.7 | 94.475
-14.1 | 100.4 | 99.4571 | 98.1 | 95.8333 | 94.6625
-14 | 100.4 | 99.5333 | 98.1 | 95.9 | 94.7875
-13.9 | 100.4 | 99.6 | 98.175 | 95.9857 | 94.8818
-13.8 | 100.4 | 99.6571 | 98.2 | 96.0429 | 94.9
-13.7 | 100.44 | 99.7 | 98.3 | 96.175 | 95
-13.6 | 100.5 | 99.7 | 98.3 | 96.3 | 95.1
-13.5 | 100.5 | 99.7 | 98.4 | 96.35 | 95.2
-13.4 | 100.5333 | 99.7143 | 98.4286 | 96.4 | 95.32
-13.3 | 100.6 | 99.8 | 98.5 | 96.4833 | 95.5
-13.2 | 100.6 | 99.8143 | 98.6 | 96.6 | 95.5286
-13.1 | 100.6 | 99.9 | 98.6 | 96.6 | 95.6
-13 | 100.6 | 99.9 | 98.7 | 96.7 | 95.6889
-12.9 | 100.6 | 99.9143 | 98.7 | 96.7714 | 95.7
-12.8 | 100.6 | 100 | 98.7571 | 96.9 | 95.8143
-12.7 | 100.6833 | 100 | 98.8 | 96.9 | 95.95
-12.6 | 100.7333 | 100.0833 | 98.9 | 97 | 96
-12.5 | 100.8 | 100.1 | 99 | 97.1 | 96.1
-12.4 | 100.8 | 100.1 | 99 | 97.1 | 96.225
-12.3 | 100.8 | 100.1 | 99 | 97.1571 | 96.3
-12.2 | 100.8 | 100.2 | 99.1 | 97.2 | 96.3571
-12.1 | 100.8833 | 100.2 | 99.1 | 97.2714 | 96.4
-12 | 100.9 | 100.225 | 99.1 | 97.3 | 96.5
-11.9 | 100.9 | 100.3 | 99.2 | 97.4667 | 96.5
-11.8 | 100.9 | 100.3375 | 99.2429 | 97.55 | 96.59
-11.7 | 100.9 | 100.4 | 99.3 | 97.6 | 96.69
-11.6 | 101 | 100.4429 | 99.3 | 97.6 | 96.7
-11.5 | 101 | 100.5 | 99.4 | 97.6833 | 96.8
-11.4 | 101.0667 | 100.5 | 99.4 | 97.7333 | 96.8375
-11.3 | 101.1 | 100.5 | 99.4571 | 97.8 | 96.9
-11.2 | 101.1 | 100.5 | 99.5 | 97.9 | 96.9875
-11.1 | 101.1 | 100.5 | 99.5375 | 97.9667 | 97.09
-11 | 101.1 | 100.5 | 99.6 | 98 | 97.2
-10.9 | 101.1 | 100.5 | 99.6 | 98.025 | 97.32
-10.8 | 101.1 | 100.5 | 99.6 | 98.1667 | 97.4
-10.7 | 101.1667 | 100.6 | 99.6 | 98.2 | 97.4429
-10.6 | 101.1 | 100.6 | 99.6 | 98.2 | 97.5
-10.5 | 101.1 | 100.6833 | 99.7 | 98.3 | 97.57
-10.4 | 101.1 | 100.7 | 99.7 | 98.4 | 97.6
-10.3 | 101.2 | 100.7 | 99.7 | 98.4167 | 97.7
-10.2 | 101.2 | 100.7 | 99.8 | 98.5 | 97.7375
-10.1 | 101.2 | 100.7 | 99.8333 | 98.5286 | 97.8
-10 | 101.2 | 100.7 | 99.9 | 98.6 | 97.8556
-9.9 | 101.2 | 100.7625 | 99.9 | 98.6429 | 97.975
-9.8 | 101.2 | 100.8 | 100 | 98.7 | 98.025
-9.7 | 101.2 | 100.8 | 100 | 98.7 | 98.1
-9.6 | 101.2 | 100.8 | 100.0714 | 98.7857 | 98.1
-9.5 | 101.225 | 100.9 | 100.1 | 98.8 | 98.2
-9.4 | 101.3 | 100.9 | 100.1 | 98.9 | 98.24
-9.3 | 101.3 | 100.9 | 100.15 | 98.9 | 98.3
-9.2 | 101.3 | 100.9 | 100.2 | 99 | 98.3364
-9.1 | 101.3 | 100.9 | 100.2 | 99.0875 | 98.4333
-9 | 101.3 | 100.9 | 100.275 | 99.1286 | 98.5
-8.9 | 101.3 | 100.9 | 100.3 | 99.2 | 98.5333
-8.8 | 101.3 | 100.9 | 100.3 | 99.2 | 98.65
-8.7 | 101.4 | 100.9 | 100.3 | 99.26 | 98.7
-8.6 | 101.4 | 100.975 | 100.3 | 99.3 | 98.7875
-8.5 | 101.4 | 101 | 100.3125 | 99.3 | 98.8
-8.4 | 101.4 | 101.1 | 100.4 | 99.3 | 98.8333
-8.3 | 101.4 | 101.1 | 100.4 | 99.3167 | 98.9
-8.2 | 101.4 | 101.1 | 100.4 | 99.4 | 98.9
-8.1 | 101.5 | 101.1 | 100.5 | 99.4286 | 98.96
-8 | 101.45 | 101.1 | 100.5 | 99.5 | 99
-7.9 | 101.4 | 101.1 | 100.5 | 99.5429 | 99.075
-7.8 | 101.4857 | 101.1 | 100.5 | 99.6 | 99.1
-7.7 | 101.4667 | 101.1 | 100.5 | 99.7 | 99.1
-7.6 | 101.4 | 101.1 | 100.5 | 99.7 | 99.1286
-7.5 | 101.5 | 101.1 | 100.6 | 99.7 | 99.2
-7.4 | 101.5 | 101.1 | 100.6 | 99.7625 | 99.2
-7.3 | 101.5 | 101.1833 | 100.6 | 99.8 | 99.2
-7.2 | 101.5 | 101.2 | 100.6 | 99.8 | 99.2125
-7.1 | 101.5 | 101.2 | 100.6 | 99.8857 | 99.3
-7 | 101.5 | 101.2 | 100.6571 | 99.9 | 99.3556
-6.9 | 101.5 | 101.2 | 100.7 | 99.9 | 99.46
-6.8 | 101.5 | 101.2 | 100.7 | 99.9 | 99.5
-6.7 | 101.5 | 101.2 | 100.7 | 100 | 99.5
-6.6 | 101.5 | 101.2 | 100.8 | 100 | 99.6
-6.5 | 101.5 | 101.3 | 100.8 | 100 | 99.6
-6.4 | 101.5 | 101.3 | 100.8 | 100 | 99.6
-6.3 | 101.5 | 101.3 | 100.8 | 100 | 99.6625
-6.2 | 101.5 | 101.3 | 100.8 | 100 | 99.7
-6.1 | 101.5 | 101.25 | 100.8 | 100 | 99.7
-6 | 101.5 | 101.2 | 100.8 | 100.025 | 99.7
-5.9 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
-5.8 | 101.5 | 101.2 | 100.7333 | 100.1 | 99.7
-5.7 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
-5.6 | 101.5 | 101.2 | 100.8 | 100.1 | 99.775
-5.5 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-5.4 | 101.45 | 101.2 | 100.8 | 100.1 | 99.8
-5.3 | 101.4 | 101.2 | 100.8 | 100.1667 | 99.8
-5.2 | 101.4 | 101.1 | 100.8 | 100.2 | 99.8
-5.1 | 101.4 | 101.1 | 100.7571 | 100.12 | 99.8
-5 | 101.4 | 101.1 | 100.7 | 100.2 | 99.8
-4.9 | 101.3 | 101.1 | 100.6167 | 100.1 | 99.8
-4.8 | 101.3 | 101.1 | 100.6 | 100.1 | 99.7429
-4.7 | 101.3 | 101.1 | 100.6 | 100.1 | 99.7778
-4.6 | 101.2 | 101.1 | 100.5143 | 100.1 | 99.8
-4.5 | 101.2 | 101 | 100.5 | 100.1 | 99.8375
-4.4 | 101.1 | 101 | 100.5 | 100.1 | 99.9
-4.3 | 101.0833 | 100.9 | 100.4889 | 100.1 | 99.9
-4.2 | 101 | 100.85 | 100.4 | 100.1 | 99.9667
-4.1 | 101 | 100.8 | 100.4 | 100.1 | 100
-4 | 100.9 | 100.8 | 100.4 | 100.1 | 100
-3.9 | 100.85 | 100.7 | 100.4 | 100.1 | 100
-3.8 | 100.8 | 100.7 | 100.4 | 100.1 | 100
-3.7 | 100.8 | 100.7 | 100.4 | 100.1 | 100
-3.6 | 100.74 | 100.6667 | 100.4 | 100.1 | 100
-3.5 | 100.7 | 100.6 | 100.4 | 100.1875 | 100
-3.4 | 100.7 | 100.6 | 100.4 | 100.2 | 100
-3.3 | 100.7 | 100.6 | 100.4 | 100.2 | 100
-3.2 | 100.7 | 100.6 | 100.4 | 100.2 | 100
-3.1 | 100.6 | 100.6 | 100.4 | 100.2 | 100.075
-3 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
-2.9 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
-2.8 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
-2.7 | 100.6 | 100.6 | 100.4 | 100.2 | 100.2
-2.6 | 100.6 | 100.5833 | 100.4 | 100.2 | 100.2
-2.5 | 100.5286 | 100.5 | 100.4 | 100.2 | 100.2
-2.4 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
-2.3 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
-2.2 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
-2.1 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
-2 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
-1.9 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
-1.8 | 100.5 | 100.5 | 100.3 | 100.1833 | 100.2
-1.7 | 100.5 | 100.4429 | 100.3 | 100.1 | 100.13
-1.6 | 100.4 | 100.4 | 100.2714 | 100.1 | 100.1
-1.5 | 100.325 | 100.4 | 100.2 | 100.1 | 100.2
-1.4 | 100.3 | 100.4 | 100.1 | 100.1 | 100.2
-1.3 | 100.3 | 100.3 | 100.1 | 100.1 | 100.2
-1.2 | 100.2 | 100.3 | 100.1 | 100.1 | 100.2
-1.1 | 100.2 | 100.2167 | 100.1 | 100.1 | 100.1667
-1 | 100.2 | 100.1625 | 100.1 | 100.1 | 100.1
-0.9 | 100.2 | 100.1 | 100 | 100.1 | 100.1
-0.8 | 100.1 | 100.1 | 100 | 100 | 100.1
-0.7 | 100.1 | 100.1 | 100 | 100 | 100.1
-0.6 | 100 | 100.1 | 100 | 100 | 100.0375
-0.5 | 100 | 100 | 100 | 100 | 100
-0.4 | 100 | 100 | 100 | 99.9167 | 100.1
-0.3 | 100 | 100 | 100 | 100 | 100.1
-0.2 | 100 | 100 | 100 | 100 | 100.07
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100.07
0.3 | 100 | 100 | 100 | 100 | 100.1
0.4 | 100 | 100 | 100 | 100 | 100.1
0.5 | 100 | 100.075 | 100 | 100 | 100
0.6 | 100 | 100.1 | 100 | 100 | 100.0375
0.7 | 100.1 | 100.1 | 100 | 100 | 100.1
0.8 | 100.1 | 100.1 | 100 | 100 | 100.1
0.9 | 100.2 | 100.1 | 100 | 100.0833 | 100.1
1 | 100.2 | 100.1375 | 100.1 | 100.1 | 100.1
1.1 | 100.2 | 100.2 | 100.1 | 100.1 | 100.1667
1.2 | 100.2 | 100.3 | 100.1 | 100.1 | 100.2
1.3 | 100.3 | 100.3 | 100.1 | 100.1 | 100.2
1.4 | 100.3 | 100.4 | 100.1 | 100.1 | 100.2
1.5 | 100.325 | 100.4 | 100.2 | 100.1 | 100.2
1.6 | 100.4 | 100.4 | 100.2714 | 100.1 | 100.1
1.7 | 100.5 | 100.4167 | 100.3 | 100.1 | 100.13
1.8 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
1.9 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
2 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
2.1 | 100.5 | 100.5 | 100.3 | 100.2 | 100.2
2.2 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
2.3 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
2.4 | 100.5 | 100.5 | 100.4 | 100.2 | 100.2
2.5 | 100.5286 | 100.5 | 100.4 | 100.2 | 100.2
2.6 | 100.6 | 100.6 | 100.4 | 100.2 | 100.2
2.7 | 100.6 | 100.6 | 100.4 | 100.2 | 100.2
2.8 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
2.9 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
3 | 100.6 | 100.6 | 100.4 | 100.2 | 100.1
3.1 | 100.6 | 100.6 | 100.4 | 100.2 | 100.075
3.2 | 100.7 | 100.6 | 100.4 | 100.2 | 100
3.3 | 100.7 | 100.6 | 100.4 | 100.2 | 100
3.4 | 100.7 | 100.6 | 100.4 | 100.2 | 100
3.5 | 100.7 | 100.6 | 100.4 | 100.2 | 100
3.6 | 100.74 | 100.7 | 100.4 | 100.1 | 100
3.7 | 100.8 | 100.7 | 100.4 | 100.1 | 100
3.8 | 100.8 | 100.7 | 100.4 | 100.1 | 100
3.9 | 100.85 | 100.7 | 100.4 | 100.1 | 100
4 | 100.9 | 100.8 | 100.4 | 100.1 | 100
4.1 | 101 | 100.8 | 100.4 | 100.1 | 100
4.2 | 101 | 100.8333 | 100.4 | 100.1 | 99.9667
4.3 | 101.0833 | 100.9 | 100.4889 | 100.1 | 99.9
4.4 | 101.1 | 100.9667 | 100.5 | 100.1 | 99.9
4.5 | 101.2 | 101 | 100.5 | 100.1 | 99.8375
4.6 | 101.2 | 101.1 | 100.5143 | 100.1 | 99.8
4.7 | 101.3 | 101.1 | 100.6 | 100.1 | 99.7778
4.8 | 101.3 | 101.1 | 100.6 | 100.1 | 99.7429
4.9 | 101.3 | 101.1 | 100.6167 | 100.1 | 99.8
5 | 101.4 | 101.1 | 100.7 | 100.1857 | 99.8
5.1 | 101.4 | 101.1 | 100.7571 | 100.1667 | 99.8
5.2 | 101.4 | 101.1 | 100.8 | 100.2 | 99.8
5.3 | 101.4 | 101.2 | 100.8 | 100.2 | 99.8
5.4 | 101.45 | 101.2 | 100.8 | 100.1 | 99.8
5.5 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
5.6 | 101.5 | 101.2 | 100.8 | 100.1 | 99.775
5.7 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
5.8 | 101.5 | 101.2 | 100.7333 | 100.1 | 99.7
5.9 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
6 | 101.5 | 101.2 | 100.8 | 100.0857 | 99.7
6.1 | 101.5 | 101.2167 | 100.8 | 100 | 99.7
6.2 | 101.5 | 101.3 | 100.8 | 100 | 99.7
6.3 | 101.5 | 101.3 | 100.8 | 100 | 99.6625
6.4 | 101.5 | 101.3 | 100.8 | 100 | 99.6
6.5 | 101.5 | 101.3 | 100.8 | 100 | 99.6
6.6 | 101.5 | 101.22 | 100.8 | 100 | 99.6
6.7 | 101.5 | 101.2 | 100.7 | 100 | 99.5
6.8 | 101.5 | 101.2 | 100.7 | 99.9 | 99.5
6.9 | 101.5 | 101.2 | 100.7 | 99.9 | 99.46
7 | 101.5 | 101.2 | 100.6571 | 99.9 | 99.3556
7.1 | 101.5 | 101.2 | 100.6 | 99.8429 | 99.3
7.2 | 101.5 | 101.2 | 100.6 | 99.8 | 99.2125
7.3 | 101.5 | 101.2 | 100.6 | 99.8 | 99.2
7.4 | 101.5 | 101.1 | 100.6 | 99.8 | 99.2
7.5 | 101.5 | 101.1 | 100.6 | 99.7 | 99.2
7.6 | 101.4 | 101.1 | 100.5 | 99.7 | 99.1286
7.7 | 101.4667 | 101.1 | 100.5 | 99.6714 | 99.1
7.8 | 101.4857 | 101.1 | 100.5 | 99.6 | 99.1
7.9 | 101.4 | 101.1 | 100.5 | 99.6 | 99.075
8 | 101.45 | 101.1 | 100.5 | 99.5 | 99
8.1 | 101.5 | 101.1 | 100.5 | 99.4 | 98.96
8.2 | 101.4 | 101.1 | 100.4 | 99.4 | 98.9
8.3 | 101.4 | 101.1 | 100.4 | 99.3 | 98.9
8.4 | 101.4 | 101.1 | 100.4 | 99.3 | 98.8333
8.5 | 101.4 | 101 | 100.3125 | 99.3 | 98.8
8.6 | 101.4 | 101 | 100.3 | 99.3 | 98.7875
8.7 | 101.4 | 100.9 | 100.3 | 99.2167 | 98.7
8.8 | 101.3 | 100.9 | 100.3 | 99.2 | 98.65
8.9 | 101.3 | 100.9 | 100.3 | 99.2 | 98.5333
9 | 101.3 | 100.9 | 100.275 | 99.1 | 98.5
9.1 | 101.3 | 100.9 | 100.2 | 99.0375 | 98.4333
9.2 | 101.3 | 100.9 | 100.2 | 99 | 98.3364
9.3 | 101.3 | 100.9 | 100.15 | 98.9 | 98.3
9.4 | 101.3 | 100.9 | 100.1 | 98.88 | 98.24
9.5 | 101.225 | 100.9 | 100.1 | 98.8 | 98.2
9.6 | 101.2 | 100.8143 | 100.0714 | 98.8 | 98.1
9.7 | 101.2 | 100.8 | 100 | 98.7 | 98.1
9.8 | 101.2 | 100.8 | 100 | 98.7 | 98.025
9.9 | 101.2 | 100.7875 | 99.9 | 98.6875 | 97.975
10 | 101.2 | 100.7 | 99.9 | 98.6 | 97.8556
10.1 | 101.2 | 100.7 | 99.8333 | 98.5833 | 97.8
10.2 | 101.2 | 100.7 | 99.8 | 98.5 | 97.7375
10.3 | 101.2 | 100.7 | 99.7 | 98.4714 | 97.7
10.4 | 101.1 | 100.6167 | 99.7 | 98.4 | 97.6
10.5 | 101.1 | 100.6 | 99.7 | 98.3 | 97.57
10.6 | 101.1 | 100.6 | 99.6 | 98.26 | 97.5
10.7 | 101.1667 | 100.6 | 99.6 | 98.2 | 97.4429
10.8 | 101.1 | 100.5 | 99.6 | 98.1143 | 97.4
10.9 | 101.1 | 100.5 | 99.6 | 98.075 | 97.32
11 | 101.1 | 100.5 | 99.6 | 98 | 97.2
11.1 | 101.1 | 100.5 | 99.5375 | 97.9143 | 97.09
11.2 | 101.1 | 100.5 | 99.5 | 97.8714 | 96.9875
11.3 | 101.1 | 100.5 | 99.4571 | 97.8 | 96.9
11.4 | 101.0667 | 100.5 | 99.4 | 97.7833 | 96.8375
11.5 | 101 | 100.5 | 99.4 | 97.7 | 96.8
11.6 | 101 | 100.4 | 99.3 | 97.6 | 96.7
11.7 | 100.9 | 100.4 | 99.3 | 97.6 | 96.69
11.8 | 100.9 | 100.3 | 99.2429 | 97.5 | 96.59
11.9 | 100.9 | 100.3 | 99.2 | 97.42 | 96.5
12 | 100.9 | 100.25 | 99.1 | 97.3571 | 96.5
12.1 | 100.8833 | 100.2 | 99.1 | 97.225 | 96.4
12.2 | 100.8 | 100.2 | 99.1 | 97.2 | 96.3571
12.3 | 100.8 | 100.1 | 99 | 97.1167 | 96.3
12.4 | 100.8 | 100.1 | 99 | 97.1 | 96.225
12.5 | 100.8 | 100.1 | 99 | 97.05 | 96.1
12.6 | 100.7333 | 100.1 | 98.9 | 97 | 96
12.7 | 100.6833 | 100 | 98.8 | 96.9 | 95.95
12.8 | 100.6 | 100 | 98.7571 | 96.8625 | 95.8143
12.9 | 100.6 | 99.9429 | 98.7 | 96.8 | 95.7
13 | 100.6 | 99.9 | 98.7 | 96.7 | 95.6889
13.1 | 100.6 | 99.9 | 98.6 | 96.65 | 95.6
13.2 | 100.6 | 99.8333 | 98.6 | 96.6 | 95.5286
13.3 | 100.6 | 99.8 | 98.5 | 96.4333 | 95.5
13.4 | 100.5333 | 99.7333 | 98.4286 | 96.4 | 95.32
13.5 | 100.5 | 99.7 | 98.4 | 96.3 | 95.2
13.6 | 100.5 | 99.7 | 98.3 | 96.25 | 95.1
13.7 | 100.44 | 99.7 | 98.3 | 96.2 | 95
13.8 | 100.4 | 99.6857 | 98.2 | 96.1 | 94.9
13.9 | 100.4 | 99.6 | 98.175 | 96 | 94.8818
14 | 100.4 | 99.56 | 98.1 | 95.9 | 94.7875
14.1 | 100.4 | 99.4857 | 98.1 | 95.8 | 94.6625
14.2 | 100.4 | 99.4 | 98 | 95.75 | 94.475
14.3 | 100.3143 | 99.4 | 97.9 | 95.6 | 94.4
14.4 | 100.3 | 99.4 | 97.9 | 95.5 | 94.3
14.5 | 100.3 | 99.4 | 97.8 | 95.4625 | 94.27
14.6 | 100.25 | 99.35 | 97.76 | 95.3375 | 94.2
14.7 | 100.2 | 99.3 | 97.7 | 95.2125 | 94.1375
14.8 | 100.2 | 99.2167 | 97.6143 | 95.1 | 94.0857
14.9 | 100.1 | 99.2 | 97.5889 | 95.1 | 93.92
15 | 100.1 | 99.1286 | 97.5 | 94.9 | 93.76
15.1 | 100.1 | 99.1 | 97.45 | 94.8167 | 93.6556
15.2 | 100.0333 | 99.1 | 97.4 | 94.775 | 93.5444
15.3 | 100 | 99 | 97.3 | 94.6333 | 93.4333
15.4 | 100 | 99 | 97.2 | 94.5833 | 93.2
15.5 | 99.9333 | 98.9714 | 97.2 | 94.5 | 93.2
15.6 | 99.9 | 98.8286 | 97.1 | 94.4 | 93.1429
15.7 | 99.8 | 98.8 | 97.0571 | 94.3 | 92.9875
15.8 | 99.8 | 98.7714 | 96.9143 | 94.2 | 92.8667
15.9 | 99.8 | 98.7 | 96.9 | 94.125 | 92.7429
16 | 99.7 | 98.6875 | 96.7714 | 93.9375 | 92.63
16.1 | 99.7 | 98.6 | 96.7 | 93.8125 | 92.53
16.2 | 99.6 | 98.5 | 96.5571 | 93.7 | 92.5
16.3 | 99.6 | 98.4 | 96.5 | 93.65 | 92.42
16.4 | 99.5167 | 98.32 | 96.3857 | 93.525 | 92.32
16.5 | 99.475 | 98.3 | 96.3 | 93.4 | 92.14
16.6 | 99.4 | 98.2143 | 96.2 | 93.275 | 92
16.7 | 99.3833 | 98.1 | 96.15 | 93.15 | 91.8625
16.8 | 99.3 | 98.1 | 96.0143 | 93.1 | 91.7375
16.9 | 99.2 | 98 | 95.84 | 92.9571 | 91.6222
17 | 99.15 | 97.9 | 95.7857 | 92.9 | 91.5
17.1 | 99.1 | 97.8286 | 95.7 | 92.6875 | 91.36
17.2 | 99 | 97.7 | 95.6 | 92.6 | 91.22
17.3 | 99 | 97.6286 | 95.4333 | 92.5 | 91.0556
17.4 | 98.9 | 97.4833 | 95.2857 | 92.3143 | 90.9444
17.5 | 98.8143 | 97.4 | 95.15 | 92.2 | 90.78
17.6 | 98.76 | 97.3833 | 95.1 | 91.9571 | 90.64
17.7 | 98.6833 | 97.3 | 94.9857 | 91.9 | 90.5333
17.8 | 98.6 | 97.15 | 94.9 | 91.8 | 90.4125
17.9 | 98.55 | 96.9714 | 94.7833 | 91.6667 | 90.31
18 | 98.3833 | 96.9 | 94.6167 | 91.5143 | 90.2
18.1 | 98.3 | 96.775 | 94.4857 | 91.4 | 90.075
18.2 | 98.2333 | 96.6333 | 94.3429 | 91.2833 | 89.92
18.3 | 98.1714 | 96.5 | 94.1 | 91.1167 | 89.7556
18.4 | 98.1 | 96.3833 | 93.9857 | 90.95 | 89.6286
18.5 | 97.9714 | 96.3 | 93.9 | 90.8 | 89.5875
18.6 | 97.9 | 96.1333 | 93.6857 | 90.6571 | 89.4333
18.7 | 97.74 | 96 | 93.5 | 90.5143 | 89.22
18.8 | 97.5857 | 95.8857 | 93.375 | 90.35 | 89.02
18.9 | 97.4429 | 95.7429 | 93.2333 | 90.1333 | 88.8556
19 | 97.2833 | 95.5833 | 93.0667 | 89.975 | 88.7444
19.1 | 97.0571 | 95.4167 | 92.8375 | 89.7857 | 88.5909
19.2 | 96.82 | 95.2429 | 92.6667 | 89.5667 | 88.4
19.3 | 96.4667 | 95 | 92.4375 | 89.4143 | 88.2375
19.4 | 96.05 | 94.7857 | 92.2667 | 89.1875 | 88.1125
19.5 | 95.4 | 94.56 | 92.0375 | 89.05 | 88
19.6 | 94.46 | 94.3143 | 91.825 | 88.9125 | 87.8714
19.7 | 92.9 | 94 | 91.62 | 88.7667 | 87.65
19.8 | 90.1 | 93.68 | 91.3667 | 88.6 | 87.48
19.9 | 86.1667 | 93.3429 | 91.1286 | 88.3857 | 87.3143
20 | 80.4 | 92.8 | 90.9571 | 88.175 | 87.06
20.1 | 72.1 | 92.14 | 90.6429 | 87.925 | 86.88
20.2 | 63.3 | 91.4 | 90.4 | 87.7375 | 86.76
20.3 | 53.65 | 90.25 | 90.08 | 87.5444 | 86.5636
20.4 | 44.1286 | 88.5 | 89.86 | 87.3 | 86.38
20.5 | 36.2714 | 86.0333 | 89.6 | 87.075 | 86.18
20.6 | 29.2857 | 81.78 | 89.2286 | 86.8857 | 85.99
20.7 | 23.7 | 76.18 | 88.7667 | 86.6667 | 85.87
20.8 | 19.7429 | 68.7714 | 88.3857 | 86.4286 | 85.58
20.9 | 17.3 | 60.5333 | 87.9 | 86.1667 | 85.38
21 | 15.9 | 52.2 | 87.3286 | 85.8571 | 85.19
21.1 | 15.075 | 43.95 | 86.6 | 85.675 | 85.08
21.2 | 14.4143 | 37.1429 | 85.6429 | 85.425 | 84.8818
21.3 | 13.92 | 30.35 | 84.4286 | 85.1571 | 84.7
21.4 | 13.6 | 25.7667 | 82.6 | 84.9 | 84.4
21.5 | 13.3167 | 22.1 | 79.95 | 84.5875 | 84.2
21.6 | 13.0667 | 19.8143 | 76.2286 | 84.275 | 83.9
21.7 | 12.9143 | 18.3 | 71.4 | 83.9857 | 83.7
21.8 | 12.6333 | 17.3667 | 64.8 | 83.8 | 83.6
21.9 | 12.4714 | 16.6857 | 57.3667 | 83.55 | 83.3
22 | 12.3167 | 16.2 | 50 | 83.2 | 83.1
22.1 | 12.14 | 15.7667 | 43.3 | 82.85 | 82.9182
22.2 | 11.98 | 15.3333 | 37 | 82.55 | 82.63
22.3 | 11.8 | 15.0286 | 31.9 | 82.175 | 82.51
22.4 | 11.6286 | 14.6857 | 27.9 | 81.8714 | 82.14
22.5 | 11.4833 | 14.4667 | 24.7143 | 81.5286 | 81.9364
22.6 | 11.3286 | 14.22 | 22.6667 | 81.225 | 81.66
22.7 | 11.1857 | 14 | 21.4 | 80.6571 | 81.36
22.8 | 11.1 | 13.75 | 20.4714 | 80.2 | 81.1222
22.9 | 10.9833 | 13.5429 | 19.8333 | 79.65 | 80.9
23 | 10.8 | 13.2571 | 19.1667 | 79.1286 | 80.675
23.1 | 10.6429 | 13.0833 | 18.725 | 78.4 | 80.4333
23.2 | 10.5 | 12.9286 | 18.35 | 77.4143 | 80.15
23.3 | 10.45 | 12.75 | 17.9429 | 76.1 | 79.88
23.4 | 10.28 | 12.525 | 17.575 | 74.35 | 79.575
23.5 | 10.2 | 12.3714 | 17.3 | 71.6833 | 79.2
23.6 | 10.1 | 12.2 | 16.9667 | 68.5 | 78.9333
23.7 | 10 | 12 | 16.7 | 64.5 | 78.6571
23.8 | 9.9 | 11.9167 | 16.4333 | 59.3571 | 78.4
23.9 | 9.7667 | 11.7333 | 16.2143 | 53.9 | 78.16
24 | 9.7 | 11.575 | 15.9286 | 48.3333 | 77.89
24.1 | 9.6 | 11.42 | 15.68 | 43.1143 | 77.52
24.2 | 9.5 | 11.3 | 15.46 | 38.3 | 77.19
24.3 | 9.4286 | 11.1 | 15.1857 | 34.2429 | 76.9091
24.4 | 9.3571 | 11.0167 | 14.9667 | 30.7 | 76.5
24.5 | 9.2 | 10.9 | 14.7667 | 28.3667 | 76.08
24.6 | 9.1143 | 10.7571 | 14.6 | 26.6 | 75.76
24.7 | 9.0429 | 10.6 | 14.35 | 25.4 | 75.3
24.8 | 8.98 | 10.45 | 14.1714 | 24.45 | 74.8
24.9 | 8.9 | 10.3857 | 13.9333 | 23.75 | 74.3
25 | 8.8 | 10.2333 | 13.7667 | 23.2 | 73.6636
25.1 | 8.7 | 10.1 | 13.5286 | 22.7444 | 72.9
25.2 | 8.6625 | 10.0286 | 13.3714 | 22.2625 | 71.8333
25.3 | 8.5286 | 9.9 | 13.1429 | 21.925 | 70.5222
25.4 | 8.5 | 9.8429 | 13.0125 | 21.6125 | 68.8
25.5 | 8.4 | 9.7 | 12.85 | 21.3 | 66.6
25.6 | 8.3 | 9.6 | 12.6857 | 20.95 | 63.2545
25.7 | 8.2625 | 9.5 | 12.5333 | 20.5889 | 59.8
25.8 | 8.2 | 9.4167 | 12.3857 | 20.3286 | 55.6
25.9 | 8.15 | 9.3667 | 12.2429 | 20.05 | 50.825
26 | 8.1 | 9.2 | 12.1 | 19.75 | 46.2143
26.1 | 8.0143 | 9.1167 | 11.92 | 19.45 | 41.9286
26.2 | 7.9 | 9.05 | 11.7714 | 19.2 | 38.2
26.3 | 7.9 | 8.9857 | 11.7 | 18.95 | 35.12
26.4 | 7.8 | 8.9 | 11.5 | 18.75 | 32.3778
26.5 | 7.8 | 8.8 | 11.4 | 18.55 | 30.275
26.6 | 7.7 | 8.7 | 11.2667 | 18.2857 | 28.85
26.7 | 7.6833 | 8.6 | 11.1143 | 18.0286 | 27.8
26.8 | 7.6 | 8.5143 | 11 | 17.6857 | 26.95
26.9 | 7.5571 | 8.44 | 10.9143 | 17.5333 | 26.2714
27 | 7.5 | 8.4 | 10.7714 | 17.375 | 25.8
27.1 | 7.5 | 8.3429 | 10.6286 | 17.15 | 25.3286
27.2 | 7.4 | 8.2 | 10.4833 | 16.88 | 24.9
27.3 | 7.3286 | 8.2 | 10.4 | 16.6857 | 24.57
27.4 |  | 8.1 | 10.3 | 16.52 | 24.2667
27.5 |  | 8.0143 | 10.2 | 16.26 | 23.9333
27.6 |  | 8 | 10.1 | 16.08 | 23.6
27.7 |  | 7.9143 | 9.9667 | 15.8833 | 23.3111
27.8 |  | 7.9 | 9.9 | 15.7 | 23.075
27.9 |  | 7.8286 | 9.7875 | 15.55 | 22.8444
28 |  | 7.7714 | 9.7 | 15.3667 | 22.5
28.1 |  | 7.7 | 9.5875 | 15.1286 | 22.25
28.2 |  | 7.6667 | 9.5 | 14.98 | 22
28.3 |  |  | 9.4 | 14.8 | 21.7286
28.4 |  |  | 9.35 | 14.65 | 21.4429
28.5 |  |  | 9.2143 | 14.5125 | 21.2286
28.6 |  |  | 9.1333 | 14.3625 | 21
28.7 |  |  | 9.0714 | 14.2286 | 20.7714
28.8 |  |  | 8.9375 | 14 | 20.5429
28.9 |  |  | 8.9 | 13.9143 | 20.4125
29 |  |  | 8.84 | 13.7714 | 20.0667
29.1 |  |  | 8.7714 | 13.6 | 19.8286
29.2 |  |  | 8.7 | 13.4167 | 19.6556
29.3 |  |  | 8.6125 | 13.35 | 19.4625
29.4 |  |  | 8.5 | 13.1833 | 19.2571
29.5 |  |  | 8.4833 | 13.0375 | 19.0778
29.6 |  |  | 8.4 | 12.8833 | 18.8455
29.7 |  |  |  | 12.7167 | 18.5889
29.8 |  |  |  | 12.5714 | 18.4556
29.9 |  |  |  | 12.4375 | 18.24
30 |  |  |  | 12.3125 | 18.0727
30.1 |  |  |  | 12.2 | 17.96
30.2 |  |  |  | 12.1625 | 17.76
30.3 |  |  |  | 12 | 17.575
30.4 |  |  |  | 11.8625 | 17.42
30.5 |  |  |  | 11.8 | 17.26
30.6 |  |  |  | 11.6857 | 17.1429
30.7 |  |  |  | 11.55 | 16.9
30.8 |  |  |  | 11.425 | 16.6857
30.9 |  |  |  | 11.3 | 16.5429
31 |  |  |  | 11.25 | 16.4125
31.1 |  |  |  | 11.1 | 16.28
31.2 |  |  |  | 11.0625 | 16.1556
31.3 |  |  |  | 11 | 15.9571
31.4 |  |  |  | 10.8833 | 15.8143
31.5 |  |  |  | 10.7167 | 15.6778
31.6 |  |  |  | 10.6 | 15.5333
31.7 |  |  |  | 10.5714 | 15.36
31.8 |  |  |  | 10.4375 | 15.26
31.9 |  |  |  | 10.4 | 15.12
32 |  |  |  | 10.3 | 14.9273
32.1 |  |  |  | 10.175 | 14.77
32.2 |  |  |  | 10.1 | 14.67
32.3 |  |  |  | 10 | 14.57
32.4 |  |  |  |  | 14.44
32.5 |  |  |  |  | 14.2625
32.6 |  |  |  |  | 14.1444
32.7 |  |  |  |  | 14.04
32.8 |  |  |  |  | 13.94
32.9 |  |  |  |  | 13.84
33 |  |  |  |  | 13.74
33.1 |  |  |  |  | 13.6333
33.2 |  |  |  |  | 13.5125
33.3 |  |  |  |  | 13.4
33.4 |  |  |  |  | 13.25
33.5 |  |  |  |  | 13.0556
33.6 |  |  |  |  | 12.95
33.7 |  |  |  |  | 12.85
33.8 |  |  |  |  | 12.75
33.9 |  |  |  |  | 12.6545
34 |  |  |  |  | 12.5556
34.1 |  |  |  |  | 12.4286
34.2 |  |  |  |  | 12.32
34.3 |  |  |  |  | 12.22
34.4 |  |  |  |  | 12.12
34.5 |  |  |  |  | 12.02
34.6 |  |  |  |  | 11.92
34.7 |  |  |  |  | 11.8111
34.8 |  |  |  |  | 11.7
34.9 |  |  |  |  | 11.7
35 |  |  |  |  | 11.675

## Wedge 15 Output Factors
Output factors
Wedge 15 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.879 | 0.89 | 0.897 | 0.906 | 0.914 | 0.92 | 0.924 | 0.928
| 4 | 0.89 | 0.904 | 0.913 | 0.925 | 0.934 | 0.943 | 0.948 | 0.954
| 5 | 0.899 | 0.915 | 0.926 | 0.94 | 0.951 | 0.961 | 0.967 | 0.975
| 7 | 0.911 | 0.93 | 0.944 | 0.962 | 0.976 | 0.99 | 0.999 | 1.009
| 10 | 0.922 | 0.944 | 0.96 | 0.982 | 1 | 1.019 | 1.031 | 1.045
| 15 | 0.931 | 0.955 | 0.973 | 0.999 | 1.023 | 1.047 | 1.065 | 1.085
| 20 | 0.936 | 0.962 | 0.981 | 1.009 | 1.037 | 1.066 | 1.087 | 1.112
| 30 | 0.942 | 0.97 | 0.991 | 1.023 | 1.055 | 1.091 | 1.118 | 1.151
| 40 | 0.946 | 0.975 | 0.997 | 1.031 | 1.066 | 1.105 | 1.136 | 1.173

## Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 42.5 | 48.1 | 56.8 | 65 | 68.1
0.1 | 47.7667 | 53 | 61.55 | 69.4 | 72.54
0.2 | 56.6 | 61.2 | 68.72 | 75.85 | 78.125
0.3 | 66.66 | 71.05 | 77.35 | 82.5 | 84.42
0.4 | 74.95 | 78.325 | 83.5 | 87.775 | 89.225
0.5 | 81.6 | 84.12 | 88.5 | 91.7 | 93
0.6 | 86.5 | 88.675 | 92.05 | 94.64 | 95.5
0.7 | 90.65 | 92.26 | 94.725 | 96.65 | 97.25
0.8 | 93.42 | 94.7 | 96.55 | 98.125 | 98.48
0.9 | 95.7 | 96.625 | 97.9 | 98.9 | 99.1
1 | 97.3 | 97.9 | 98.925 | 99.5333 | 99.6
1.1 | 98.46 | 98.775 | 99.5 | 99.8 | 99.9
1.2 | 99.225 | 99.525 | 99.825 | 100 | 100
1.3 | 99.72 | 99.825 | 99.94 | 100 | 100
1.4 | 99.9 | 99.96 | 100 | 99.9833 | 99.9
1.5 | 100 | 100 | 100 | 99.9 | 99.8
1.6 | 99.9667 | 99.95 | 99.88 | 99.575 | 99.5333
1.7 | 99.86 | 99.7 | 99.55 | 99.4 | 99.18
1.8 | 99.6333 | 99.58 | 99.3 | 99 | 98.96
1.9 | 99.36 | 99.32 | 99.05 | 98.8 | 98.58
2 | 99 | 98.88 | 98.66 | 98.475 | 98.35
2.1 | 98.65 | 98.5 | 98.42 | 98.1333 | 97.975
2.2 | 98.175 | 98.1 | 98.05 | 97.775 | 97.62
2.3 | 97.62 | 97.8 | 97.72 | 97.45 | 97.24
2.4 | 97.1 | 97.45 | 97.325 | 97 | 96.9
2.5 | 96.64 | 97 | 96.9667 | 96.7333 | 96.55
2.6 | 96.2 | 96.66 | 96.6 | 96.375 | 96.26
2.7 | 95.85 | 96.25 | 96.175 | 95.95 | 95.85
2.8 | 95.4 | 95.85 | 95.75 | 95.58 | 95.4333
2.9 | 94.7 | 95.35 | 95.4 | 95.2 | 95.175
3 | 94.3 | 95.0333 | 95.1 | 94.84 | 94.78
3.1 | 93.78 | 94.575 | 94.65 | 94.46 | 94.45
3.2 | 93.3 | 94.15 | 94.28 | 94.1 | 94.15
3.3 | 92.9 | 93.76 | 93.9 | 93.75 | 93.6667
3.4 | 92.35 | 93.3333 | 93.54 | 93.4 | 93.38
3.5 | 91.9 | 92.95 | 93.1 | 93 | 92.975
3.6 | 91.4333 | 92.46 | 92.7333 | 92.72 | 92.62
3.7 | 90.9 | 92.08 | 92.4333 | 92.3667 | 92.225
3.8 | 90.4 | 91.72 | 91.9 | 91.9 | 91.9
3.9 | 89.82 | 91.24 | 91.575 | 91.55 | 91.6
4 | 89.45 | 90.75 | 91.15 | 91.3 | 91.1667
4.1 | 88.95 | 90.35 | 90.775 | 90.85 | 90.8
4.2 | 88.5 | 90 | 90.48 | 90.56 | 90.46
4.3 | 87.94 | 89.5333 | 90.1 | 90.15 | 90.05
4.4 | 87.475 | 89.1 | 89.725 | 89.825 | 89.7
4.5 | 87 | 88.72 | 89.3333 | 89.4 | 89.35
4.6 | 86.5667 | 88.25 | 88.9 | 89.04 | 88.96
4.7 | 86.02 | 87.9 | 88.55 | 88.6 | 88.65
4.8 | 85.66 | 87.5 | 88.1 | 88.38 | 88.25
4.9 | 85.24 | 87.14 | 87.82 | 88.05 | 87.9333
5 | 84.7 | 86.64 | 87.425 | 87.68 | 87.5
5.1 | 83.9667 | 86.25 | 87.2 | 87.2 | 87.1
5.2 | 83.5667 | 85.8 | 86.75 | 86.825 | 86.775
5.3 | 83.2 | 85.475 | 86.4 | 86.5429 | 86.375
5.4 | 82.675 | 85.02 | 85.96 | 86.15 | 86.08
5.5 | 82.175 | 84.5 | 85.65 | 85.725 | 85.775
5.6 | 81.6 | 84.25 | 85.1667 | 85.55 | 85.46
5.7 | 81.28 | 83.98 | 84.8667 | 85.2 | 85.12
5.8 | 80.95 | 83.44 | 84.56 | 84.84 | 84.75
5.9 | 80.3 | 83.04 | 84.2 | 84.36 | 84.35
6 | 79.86 | 82.74 | 83.6667 | 83.95 | 84
6.1 | 79.5 | 82.2333 | 83.34 | 83.7 | 83.7
6.2 | 78.95 | 81.7 | 82.975 | 83.3667 | 83.35
6.3 | 78.6 | 81.425 | 82.65 | 82.95 | 82.95
6.4 | 78.08 | 81 | 82.325 | 82.675 | 82.65
6.5 | 77.7 | 80.65 | 81.9 | 82.38 | 82.4
6.6 | 77.1 | 80.3 | 81.6 | 82.02 | 82.02
6.7 | 76.8 | 79.8 | 81.2 | 81.66 | 81.64
6.8 | 76.2 | 79.5 | 80.86 | 81.25 | 81.3333
6.9 | 75.925 | 79.125 | 80.5 | 81.02 | 80.95
7 | 75.46 | 78.75 | 80.14 | 80.625 | 80.6
7.1 | 75.05 | 78.2333 | 79.88 | 80.2667 | 80.35
7.2 | 74.52 | 77.975 | 79.58 | 80.0167 | 79.98
7.3 | 74.08 | 77.65 | 79.1333 | 79.6 | 79.65
7.4 | 73.75 | 77.24 | 78.86 | 79.32 | 79.325
7.5 | 73.35 | 76.8286 | 78.4 | 78.9333 | 78.95
7.6 | 72.95 | 76.3667 | 78.16 | 78.675 | 78.625
7.7 | 72.5 | 76.0429 | 77.78 | 78.35 | 78.4
7.8 | 72.0667 | 75.6667 | 77.4571 | 78 | 78
7.9 | 71.6 | 75.3 | 77.04 | 77.65 | 77.64
8 | 71.25 | 74.95 | 76.6667 | 77.3 | 77.325
8.1 | 70.88 | 74.48 | 76.4 | 76.95 | 76.94
8.2 | 70.42 | 74.16 | 76.08 | 76.62 | 76.675
8.3 | 69.92 | 73.7 | 75.65 | 76.24 | 76.36
8.4 | 69.5 | 73.4 | 75.34 | 76 | 76.075
8.5 | 69.125 | 73.05 | 75.05 | 75.64 | 75.75
8.6 | 68.7333 | 72.7 | 74.56 | 75.3 | 75.3
8.7 | 68.36 | 72.3 | 74.3 | 74.9714 | 75.0333
8.8 | 68.0333 | 71.95 | 73.92 | 74.6667 | 74.7667
8.9 | 67.7 | 71.65 | 73.5333 | 74.325 | 74.38
9 | 67.15 | 71.2 | 73.2 | 74.05 | 74.075
9.1 | 66.7667 | 70.875 | 72.8 | 73.64 | 73.75
9.2 | 66.36 | 70.5 | 72.6 | 73.44 | 73.375
9.3 | 65.9333 | 70.14 | 72.2 | 73.0667 | 73.1
9.4 | 65.5333 | 69.775 | 71.9 | 72.7333 | 72.75
9.5 | 65.12 | 69.4 | 71.6 | 72.38 | 72.38
9.6 | 64.725 | 69 | 71.32 | 72.12 | 72.1714
9.7 | 64.3 | 68.5667 | 71.0667 | 71.7333 | 71.86
9.8 | 64 | 68.3 | 70.7 | 71.4 | 71.45
9.9 | 63.68 | 67.95 | 70.3 | 71.1 | 71.14
10 | 63.25 | 67.58 | 69.975 | 70.78 | 70.8
10.1 | 62.8 | 67.3 | 69.65 | 70.52 | 70.56
10.2 | 62.3 | 66.8 | 69.2 | 70.15 | 70.28
10.3 | 62.05 | 66.5 | 69 | 69.8 | 70.1
10.4 | 61.66 | 66.2 | 68.7 | 69.5667 | 69.6
10.5 | 61.3 | 65.8 | 68.4 | 69.22 | 69.2
10.6 | 61.025 | 65.45 | 67.92 | 68.9 | 69
10.7 | 60.58 | 65.1 | 67.65 | 68.575 | 68.7
10.8 | 60.3 | 64.92 | 67.4 | 68.275 | 68.38
10.9 | 59.9 | 64.5333 | 67.1333 | 67.98 | 68.15
11 | 59.58 | 64.225 | 66.85 | 67.7 | 67.85
11.1 | 59.3 | 63.9 | 66.5 | 67.36 | 67.52
11.2 | 58.84 | 63.625 | 66.16 | 67.1 | 67.24
11.3 | 58.5 | 63.14 | 65.76 | 66.8 | 66.9
11.4 | 58.2167 | 62.74 | 65.45 | 66.4667 | 66.6
11.5 | 57.84 | 62.4 | 65.12 | 66.2 | 66.3167
11.6 | 57.55 | 62.12 | 64.8 | 65.9667 | 66.1
11.7 | 57.15 | 61.7 | 64.5 | 65.64 | 65.8
11.8 | 56.7667 | 61.4667 | 64.1667 | 65.35 | 65.575
11.9 | 56.48 | 61.2 | 63.9333 | 65.0333 | 65.22
12 | 56.1 | 60.9 | 63.6667 | 64.725 | 64.9
12.1 | 55.85 | 60.6 | 63.4 | 64.44 | 64.65
12.2 | 55.4667 | 60.24 | 63.08 | 64.16 | 64.46
12.3 | 55.16 | 59.92 | 62.82 | 63.8429 | 64.05
12.4 | 54.9 | 59.62 | 62.5333 | 63.6 | 63.68
12.5 | 54.5667 | 59.4 | 62.2 | 63.3 | 63.45
12.6 | 54.225 | 58.95 | 61.92 | 63 | 63.22
12.7 | 53.95 | 58.65 | 61.66 | 62.775 | 62.94
12.8 | 53.65 | 58.425 | 61.4 | 62.5 | 62.75
12.9 | 53.325 | 58.0667 | 61.15 | 62.225 | 62.38
13 | 52.9333 | 57.8 | 60.85 | 61.95 | 62.15
13.1 | 52.6 | 57.4667 | 60.6 | 61.64 | 61.82
13.2 | 52.325 | 57.1 | 60.3 | 61.3 | 61.525
13.3 | 52 | 56.8333 | 59.96 | 61.1 | 61.3
13.4 | 51.85 | 56.5667 | 59.64 | 60.7833 | 61.0167
13.5 | 51.54 | 56.2 | 59.325 | 60.6 | 60.82
13.6 | 51.26 | 55.88 | 59.02 | 60.26 | 60.6
13.7 | 50.9 | 55.56 | 58.8 | 59.95 | 60.28
13.8 | 50.44 | 55.325 | 58.525 | 59.68 | 59.9333
13.9 | 50.14 | 55.04 | 58.2 | 59.45 | 59.62
14 | 49.82 | 54.725 | 57.86 | 59.16 | 59.35
14.1 | 49.6 | 54.525 | 57.64 | 58.9 | 59.1167
14.2 | 49.28 | 54.25 | 57.4 | 58.65 | 58.85
14.3 | 49.05 | 53.88 | 57.25 | 58.45 | 58.6
14.4 | 48.7 | 53.6333 | 56.825 | 58.125 | 58.32
14.5 | 48.4667 | 53.32 | 56.4833 | 57.9 | 58.04
14.6 | 48.2333 | 53.08 | 56.3 | 57.65 | 57.8
14.7 | 47.94 | 52.825 | 56 | 57.35 | 57.4
14.8 | 47.64 | 52.5 | 55.8 | 57.1 | 57.28
14.9 | 47.35 | 52.25 | 55.4 | 56.8 | 57.075
15 | 47 | 51.9 | 55.2 | 56.5667 | 56.78
15.1 | 46.7 | 51.575 | 54.96 | 56.3 | 56.48
15.2 | 46.6 | 51.36 | 54.675 | 56.075 | 56.22
15.3 | 46.26 | 51.05 | 54.4 | 55.82 | 55.925
15.4 | 45.95 | 50.825 | 54.16 | 55.54 | 55.7
15.5 | 45.68 | 50.5 | 53.925 | 55.2 | 55.4667
15.6 | 45.4 | 50.25 | 53.675 | 55.08 | 55.22
15.7 | 45.15 | 49.9667 | 53.42 | 54.85 | 54.98
15.8 | 44.92 | 49.7 | 53.0833 | 54.5333 | 54.78
15.9 | 44.675 | 49.45 | 52.9 | 54.3 | 54.4667
16 | 44.35 | 49.14 | 52.6 | 54.0429 | 54.2
16.1 | 44.1667 | 48.88 | 52.3 | 53.775 | 54
16.2 | 43.88 | 48.62 | 52.1 | 53.525 | 53.775
16.3 | 43.65 | 48.375 | 51.84 | 53.3 | 53.52
16.4 | 43.24 | 48.1333 | 51.6 | 53.025 | 53.26
16.5 | 43.0667 | 47.85 | 51.35 | 52.82 | 53.05
16.6 | 42.8 | 47.6 | 51.12 | 52.6 | 52.8
16.7 | 42.54 | 47.5 | 50.9167 | 52.3 | 52.62
16.8 | 42.3 | 47.18 | 50.75 | 52.1 | 52.38
16.9 | 42.1 | 46.975 | 50.425 | 51.88 | 52.14
17 | 41.8333 | 46.66 | 50.16 | 51.6 | 51.8
17.1 | 41.6167 | 46.36 | 49.9 | 51.44 | 51.575
17.2 | 41.4 | 46.125 | 49.64 | 51.2 | 51.35
17.3 | 41.1 | 45.9 | 49.46 | 50.975 | 51.0667
17.4 | 40.8667 | 45.68 | 49.24 | 50.72 | 50.9
17.5 | 40.6 | 45.375 | 48.95 | 50.5 | 50.7333
17.6 | 40.4667 | 45.16 | 48.74 | 50.26 | 50.5
17.7 | 40.175 | 44.95 | 48.54 | 50.05 | 50.3
17.8 | 39.925 | 44.72 | 48.3 | 49.76 | 49.9
17.9 | 39.72 | 44.5 | 48.05 | 49.64 | 49.82
18 | 39.52 | 44.2 | 47.75 | 49.32 | 49.62
18.1 | 39.3333 | 44.02 | 47.6 | 49.16 | 49.34
18.2 | 39.04 | 43.85 | 47.36 | 48.925 | 49.1
18.3 | 38.825 | 43.6333 | 47.14 | 48.62 | 48.875
18.4 | 38.5833 | 43.34 | 46.925 | 48.45 | 48.6714
18.5 | 38.375 | 43.1333 | 46.66 | 48.2667 | 48.475
18.6 | 38.125 | 42.86 | 46.425 | 47.95 | 48.14
18.7 | 37.88 | 42.6667 | 46.2 | 47.8 | 47.98
18.8 | 37.7 | 42.38 | 46 | 47.56 | 47.74
18.9 | 37.375 | 42.02 | 45.75 | 47.3 | 47.525
19 | 37.225 | 41.875 | 45.6 | 47.075 | 47.3
19.1 | 37 | 41.65 | 45.36 | 46.8 | 47.08
19.2 | 36.8 | 41.5 | 45.0667 | 46.6 | 46.9
19.3 | 36.64 | 41.275 | 44.9 | 46.425 | 46.725
19.4 | 36.425 | 41.0667 | 44.725 | 46.2 | 46.44
19.5 | 36.2 | 40.825 | 44.475 | 45.95 | 46.25
19.6 | 35.975 | 40.525 | 44.22 | 45.74 | 46.1
19.7 | 35.7667 | 40.35 | 44.02 | 45.6 | 45.9
19.8 | 35.58 | 40.15 | 43.82 | 45.38 | 45.6
19.9 | 35.4 | 39.88 | 43.6333 | 45.2 | 45.36
20 | 35.16 | 39.7 | 43.42 | 44.9333 | 45.16
20.1 | 34.94 | 39.46 | 43.2 | 44.68 | 44.9
20.2 | 34.7333 | 39.24 | 42.9 | 44.44 | 44.76
20.3 | 34.5 | 39.04 | 42.68 | 44.225 | 44.54
20.4 | 34.38 | 38.8 | 42.4667 | 44.1 | 44.3
20.5 | 34.2 | 38.6 | 42.3 | 43.85 | 44.15
20.6 | 34 | 38.4 | 42.12 | 43.7 | 43.9
20.7 | 33.82 | 38.175 | 41.92 | 43.475 | 43.78
20.8 | 33.575 | 38.04 | 41.8 | 43.24 | 43.55
20.9 | 33.35 | 37.82 | 41.5167 | 43.1 | 43.3
21 | 33.16 | 37.675 | 41.35 | 42.8 | 43.14
21.1 | 33 | 37.45 | 41.125 | 42.7 | 43
21.2 | 32.82 | 37.3 | 40.86 | 42.5 | 42.775
21.3 | 32.6 | 37 | 40.625 | 42.3 | 42.6
21.4 | 32.4 | 36.8 | 40.475 | 42.1 | 42.3
21.5 | 32.2 | 36.7 | 40.24 | 41.8833 | 42.18
21.6 | 32.025 | 36.42 | 40.1 | 41.7 | 42
21.7 | 31.84 | 36.3 | 39.88 | 41.5667 | 41.88
21.8 | 31.62 | 36.08 | 39.7 | 41.35 | 41.6
21.9 | 31.44 | 35.86 | 39.5 | 41.12 | 41.4
22 | 31.32 | 35.725 | 39.275 | 40.94 | 41.2
22.1 | 31.08 | 35.48 | 39.16 | 40.725 | 41.04
22.2 | 30.96 | 35.2667 | 38.925 | 40.5 | 40.8714
22.3 | 30.7667 | 35.1 | 38.74 | 40.3333 | 40.68
22.4 | 30.575 | 34.9714 | 38.54 | 40.225 | 40.48
22.5 | 30.45 | 34.7833 | 38.4 | 40.04 | 40.35
22.6 | 30.2 | 34.6 | 38.175 | 39.825 | 40.2
22.7 | 30.1 | 34.4 | 38 | 39.6333 | 39.95
22.8 | 30 | 34.25 | 37.8333 | 39.4 | 39.72
22.9 | 29.8 | 34.025 | 37.6333 | 39.2667 | 39.52
23 | 29.6 | 33.9 | 37.4 | 39.18 | 39.4
23.1 | 29.4 | 33.6667 | 37.3 | 38.9 | 39.22
23.2 | 29.2333 | 33.4333 | 37.125 | 38.7167 | 39
23.3 | 29.1 | 33.325 | 36.925 | 38.54 | 38.88
23.4 | 28.9167 | 33.15 | 36.8 | 38.32 | 38.68
23.5 | 28.72 | 32.925 | 36.5833 | 38.1 | 38.5
23.6 | 28.6 | 32.8 | 36.4167 | 37.98 | 38.3
23.7 | 28.48 | 32.575 | 36.3 | 37.86 | 38.1
23.8 | 28.25 | 32.4 | 36.0667 | 37.68 | 37.9
23.9 | 28.1 | 32.25 | 35.85 | 37.5 | 37.75
24 | 28 | 32.04 | 35.62 | 37.3 | 37.6167
24.1 | 27.84 | 31.9 | 35.45 | 37.16 | 37.38
24.2 | 27.64 | 31.68 | 35.3 | 36.9667 | 37.26
24.3 | 27.5 | 31.5 | 35.1 | 36.76 | 37.05
24.4 | 27.3 | 31.32 | 35 | 36.5833 | 36.84
24.5 | 27.1167 | 31.2 | 34.75 | 36.5 | 36.7
24.6 | 27 | 31.0333 | 34.6 | 36.3 | 36.52
24.7 | 26.875 | 30.88 | 34.425 | 36.15 | 36.42
24.8 | 26.65 | 30.7 | 34.3 | 35.92 | 36.2
24.9 | 26.45 | 30.62 | 34.15 | 35.78 | 36.08
25 | 26.32 | 30.38 | 33.925 | 35.66 | 35.88
25.1 | 26.2 | 30.2 | 33.6833 | 35.375 | 35.7
25.2 | 26 | 30.025 | 33.575 | 35.26 | 35.5
25.3 | 25.9 | 29.9 | 33.46 | 35.06 | 35.38
25.4 | 25.7714 | 29.725 | 33.24 | 34.8833 | 35.18
25.5 | 25.6 | 29.6 | 33.1 | 34.7 | 35
25.6 | 25.5333 | 29.4 | 32.9333 | 34.6 | 34.9
25.7 | 25.4 | 29.3 | 32.8 | 34.46 | 34.75
25.8 | 25.2 | 29.2 | 32.6 | 34.2667 | 34.6
25.9 | 25.05 | 29 | 32.475 | 34.1 | 34.42
26 | 24.9667 | 28.85 | 32.3333 | 33.95 | 34.3333
26.1 | 24.8 | 28.7 | 32.2 | 33.8 | 34.14
26.2 | 24.6 | 28.5 | 31.98 | 33.6 | 33.96
26.3 | 24.5 | 28.34 | 31.8 | 33.5 | 33.75
26.4 | 24.38 | 28.2 | 31.7 | 33.3333 | 33.64
26.5 | 24.3 | 28.1 | 31.6 | 33.2 | 33.5
26.6 | 24.14 | 27.925 | 31.475 | 33.075 | 33.3
26.7 | 24 | 27.84 | 31.2333 | 32.86 | 33.16
26.8 | 23.8 | 27.7 | 31.0333 | 32.75 | 33.1
26.9 | 23.7 | 27.48 | 30.95 | 32.55 | 32.85
27 | 23.6 | 27.375 | 30.775 | 32.38 | 32.7
27.1 | 23.42 | 27.15 | 30.64 | 32.2667 | 32.6
27.2 | 23.2 | 27 | 30.525 | 32.15 | 32.45
27.3 | 23.175 | 26.9 | 30.4 | 32 | 32.34
27.4 | 23 | 26.78 | 30.16 | 31.8 | 32.125
27.5 | 22.9 | 26.6 | 30.0667 | 31.6833 | 32
27.6 | 22.78 | 26.5 | 29.88 | 31.5167 | 31.8
27.7 | 22.6 | 26.4 | 29.7 | 31.34 | 31.7
27.8 | 22.52 | 26.2 | 29.6143 | 31.225 | 31.5
27.9 | 22.42 | 26.1 | 29.4 | 31.1 | 31.4
28 | 22.2833 | 25.9 | 29.3 | 30.98 | 31.3
28.1 | 22.2 | 25.8 | 29.1 | 30.8 | 31.1
28.2 | 22.075 | 25.7 | 29 | 30.7 | 30.9
28.3 | 21.88 | 25.6 | 28.9 | 30.5167 | 30.8
28.4 | 21.775 | 25.4167 | 28.7 | 30.36 | 30.68
28.5 | 21.625 | 25.3 | 28.575 | 30.24 | 30.55
28.6 | 21.55 | 25.2 | 28.42 | 30.1 | 30.35
28.7 | 21.425 | 25 | 28.26 | 29.9 | 30.2
28.8 | 21.3 | 24.925 | 28.12 | 29.7833 | 30.1
28.9 | 21.2 | 24.78 | 28 | 29.7 | 30
29 | 21.0167 | 24.68 | 27.9 | 29.5333 | 29.9
29.1 | 20.975 | 24.475 | 27.8 | 29.4 | 29.8
29.2 | 20.8 | 24.36 | 27.7 | 29.3 | 29.525
29.3 | 20.66 | 24.25 | 27.5 | 29.1 | 29.44
29.4 | 20.54 | 24.12 | 27.3 | 29 | 29.3
29.5 | 20.4 | 23.98 | 27.1667 | 28.825 | 29.2
29.6 | 20.4 | 23.8 | 27.075 | 28.7167 | 29.0333
29.7 | 20.2833 | 23.7 | 26.925 | 28.58 | 28.88
29.8 | 20.2 | 23.6 | 26.8 | 28.46 | 28.7667
29.9 | 20.075 | 23.5 | 26.7 | 28.36 | 28.6
30 | 19.9 | 23.4 | 26.6 | 28.14 | 28.5

## SSD80 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 43.1 | 49.8 | 63.5 | 74 | 77.5
0.1 | 48.16 | 55.575 | 68.3 | 78.15 | 81.15
0.2 | 56.1 | 63.3375 | 74.45 | 83.5 | 86
0.3 | 67.15 | 72.875 | 82.6 | 89.06 | 90.95
0.4 | 75.92 | 80.1429 | 88.1 | 92.9 | 94.5
0.5 | 82.45 | 86.04 | 92.15 | 95.75 | 96.8
0.6 | 87.5 | 90.4 | 94.8667 | 97.6 | 98.3
0.7 | 91.35 | 93.5 | 96.875 | 98.8 | 99.25
0.8 | 94.1 | 95.82 | 98.25 | 99.5 | 99.7
0.9 | 96.24 | 97.45 | 99.175 | 99.85 | 99.925
1 | 97.8 | 98.575 | 99.75 | 100 | 100
1.1 | 98.8333 | 99.325 | 100 | 100 | 100
1.2 | 99.475 | 99.825 | 100 | 100 | 99.925
1.3 | 99.825 | 100 | 100 | 99.8 | 99.65
1.4 | 100 | 100 | 99.9 | 99.5 | 99.425
1.5 | 100 | 100 | 99.6 | 99.14 | 99.05
1.6 | 100 | 99.8333 | 99.4 | 98.82 | 98.75
1.7 | 99.775 | 99.55 | 99 | 98.45 | 98.25
1.8 | 99.44 | 99.2 | 98.5 | 98.05 | 97.8
1.9 | 98.975 | 98.8 | 98.15 | 97.55 | 97.38
2 | 98.575 | 98.45 | 97.74 | 97.15 | 97
2.1 | 98.15 | 97.95 | 97.4 | 96.775 | 96.6
2.2 | 97.56 | 97.5125 | 96.8 | 96.35 | 96.1
2.3 | 97.1 | 96.975 | 96.3667 | 95.96 | 95.7
2.4 | 96.5 | 96.55 | 95.9 | 95.48 | 95.3
2.5 | 96 | 96.05 | 95.525 | 95 | 94.9
2.6 | 95.5 | 95.6 | 95.05 | 94.6 | 94.45
2.7 | 94.85 | 95.06 | 94.55 | 94.2 | 94.05
2.8 | 94.4 | 94.6 | 94.15 | 93.7 | 93.55
2.9 | 93.8 | 94.2 | 93.65 | 93.3 | 93.15
3 | 93.15 | 93.725 | 93.25 | 92.85 | 92.75
3.1 | 92.6 | 93.15 | 92.9 | 92.4 | 92.35
3.2 | 92.15 | 92.75 | 92.35 | 91.95 | 91.95
3.3 | 91.65 | 92.3 | 91.9 | 91.5 | 91.475
3.4 | 91.05 | 91.8 | 91.5 | 91.15 | 91.0667
3.5 | 90.45 | 91.3333 | 91.08 | 90.64 | 90.6333
3.6 | 89.95 | 90.9 | 90.6 | 90.25 | 90.2
3.7 | 89.3667 | 90.4 | 90.1 | 89.85 | 89.8
3.8 | 88.9 | 89.9 | 89.7 | 89.35 | 89.3
3.9 | 88.3333 | 89.4 | 89.175 | 88.925 | 88.95
4 | 87.8333 | 88.94 | 88.825 | 88.55 | 88.54
4.1 | 87.2 | 88.45 | 88.35 | 88.15 | 88
4.2 | 86.7 | 88.0375 | 87.95 | 87.6667 | 87.6
4.3 | 86.26 | 87.55 | 87.525 | 87.35 | 87.2
4.4 | 85.65 | 87.05 | 87 | 86.9 | 86.8333
4.5 | 85.16 | 86.55 | 86.6333 | 86.5 | 86.4
4.6 | 84.62 | 86.125 | 86.2 | 86 | 86
4.7 | 84.125 | 85.6333 | 85.75 | 85.5333 | 85.6
4.8 | 83.55 | 85.2 | 85.325 | 85.15 | 85.175
4.9 | 83.025 | 84.7 | 84.95 | 84.8 | 84.75
5 | 82.55 | 84.25 | 84.45 | 84.4 | 84.325
5.1 | 82 | 83.825 | 84 | 83.9 | 83.975
5.2 | 81.55 | 83.35 | 83.6 | 83.575 | 83.48
5.3 | 81 | 82.9 | 83.2 | 83.15 | 83.2
5.4 | 80.3 | 82.425 | 82.75 | 82.75 | 82.68
5.5 | 79.85 | 82 | 82.3 | 82.35 | 82.3
5.6 | 79.48 | 81.6 | 81.86 | 81.76 | 81.9
5.7 | 78.96 | 81.1 | 81.55 | 81.4 | 81.4
5.8 | 78.35 | 80.6429 | 81.0714 | 81 | 81.1667
5.9 | 77.925 | 80.25 | 80.65 | 80.7 | 80.75
6 | 77.525 | 79.72 | 80.2 | 80.3 | 80.3375
6.1 | 76.875 | 79.25 | 79.8 | 79.96 | 79.9667
6.2 | 76.36 | 78.9 | 79.5 | 79.625 | 79.7
6.3 | 76 | 78.44 | 79.02 | 79.175 | 79.22
6.4 | 75.4 | 78.0333 | 78.65 | 78.75 | 78.775
6.5 | 75.1 | 77.4667 | 78.25 | 78.35 | 78.4333
6.6 | 74.5 | 77.1 | 77.875 | 78.025 | 78.1
6.7 | 74.1 | 76.7 | 77.55 | 77.65 | 77.65
6.8 | 73.64 | 76.2667 | 77.0625 | 77.15 | 77.2
6.9 | 73.15 | 75.8 | 76.6714 | 76.8 | 76.85
7 | 72.6875 | 75.35 | 76.26 | 76.4 | 76.6
7.1 | 72.25 | 74.95 | 75.875 | 76 | 76.1429
7.2 | 71.75 | 74.525 | 75.45 | 75.675 | 75.8
7.3 | 71.35 | 74.25 | 75.1 | 75.35 | 75.45
7.4 | 70.8667 | 73.75 | 74.775 | 74.925 | 75.05
7.5 | 70.45 | 73.3 | 74.35 | 74.65 | 74.65
7.6 | 70 | 72.9 | 73.975 | 74.225 | 74.325
7.7 | 69.5 | 72.46 | 73.625 | 73.9 | 73.9667
7.8 | 69.15 | 72.1 | 73.125 | 73.55 | 73.575
7.9 | 68.6714 | 71.7857 | 72.725 | 73.1 | 73.25
8 | 68.2333 | 71.275 | 72.4 | 72.76 | 72.8
8.1 | 67.72 | 70.95 | 72 | 72.375 | 72.5
8.2 | 67.35 | 70.4875 | 71.7 | 72 | 72.2
8.3 | 66.96 | 70.1125 | 71.3 | 71.6625 | 71.8
8.4 | 66.45 | 69.6889 | 70.9 | 71.2875 | 71.3667
8.5 | 66.1167 | 69.35 | 70.56 | 70.85 | 71.025
8.6 | 65.7 | 68.8667 | 70.25 | 70.5875 | 70.7
8.7 | 65.3143 | 68.5 | 69.75 | 70.2125 | 70.4
8.8 | 64.85 | 68 | 69.425 | 69.95 | 70
8.9 | 64.375 | 67.8 | 69.08 | 69.6 | 69.72
9 | 63.975 | 67.25 | 68.7333 | 69.16 | 69.325
9.1 | 63.55 | 66.9 | 68.3667 | 68.85 | 68.975
9.2 | 63.175 | 66.625 | 68 | 68.62 | 68.725
9.3 | 62.8 | 66.125 | 67.6 | 68.25 | 68.275
9.4 | 62.48 | 65.775 | 67.2 | 67.75 | 67.975
9.5 | 62.1 | 65.4 | 66.9 | 67.375 | 67.6
9.6 | 61.5 | 64.975 | 66.5333 | 67.125 | 67.25
9.7 | 61.225 | 64.65 | 66.225 | 66.75 | 66.94
9.8 | 60.75 | 64.225 | 65.95 | 66.4 | 66.62
9.9 | 60.4 | 63.8667 | 65.45 | 66.1 | 66.225
10 | 59.95 | 63.4667 | 65.1625 | 65.8 | 65.8125
10.1 | 59.525 | 63.1125 | 64.7875 | 65.4 | 65.4833
10.2 | 59.3 | 62.7375 | 64.4125 | 65 | 65.2125
10.3 | 58.78 | 62.4 | 64.1444 | 64.7 | 64.9
10.4 | 58.4 | 62 | 63.9 | 64.45 | 64.6
10.5 | 58.15 | 61.7 | 63.425 | 64.125 | 64.2
10.6 | 57.7 | 61.25 | 63.12 | 63.7625 | 63.92
10.7 | 57.4 | 60.9 | 62.775 | 63.4 | 63.62
10.8 | 56.9143 | 60.575 | 62.425 | 63.1333 | 63.35
10.9 | 56.6222 | 60.325 | 62.06 | 62.82 | 62.975
11 | 56.2625 | 59.975 | 61.8 | 62.5 | 62.6333
11.1 | 55.8875 | 59.5 | 61.45 | 62.1143 | 62.325
11.2 | 55.5125 | 59.175 | 61.1 | 61.8 | 62
11.3 | 55.225 | 58.85 | 60.85 | 61.525 | 61.8
11.4 | 54.775 | 58.5333 | 60.5 | 61.2 | 61.5
11.5 | 54.5143 | 58.175 | 60.15 | 60.95 | 61.15
11.6 | 54.2 | 57.8 | 59.85 | 60.65 | 60.8
11.7 | 53.9 | 57.4 | 59.5 | 60.275 | 60.5
11.8 | 53.4 | 57.1 | 59.22 | 60.025 | 60.2
11.9 | 53.15 | 56.75 | 58.925 | 59.68 | 59.9
12 | 52.84 | 56.5333 | 58.575 | 59.38 | 59.58
12.1 | 52.46 | 56.1 | 58.1333 | 59.08 | 59.3
12.2 | 52.125 | 55.8 | 57.9 | 58.8 | 59
12.3 | 51.775 | 55.5 | 57.6 | 58.475 | 58.6
12.4 | 51.525 | 55.3 | 57.3 | 58.225 | 58.375
12.5 | 51.175 | 54.925 | 57 | 57.775 | 58.025
12.6 | 50.9 | 54.625 | 56.7 | 57.625 | 57.8
12.7 | 50.5333 | 54.2333 | 56.3333 | 57.3 | 57.525
12.8 | 50.175 | 53.9 | 56.05 | 57.05 | 57.3
12.9 | 49.9 | 53.6 | 55.8 | 56.7 | 56.95
13 | 49.6 | 53.3333 | 55.5333 | 56.45 | 56.6
13.1 | 49.3 | 52.95 | 55.16 | 56 | 56.375
13.2 | 49 | 52.625 | 54.925 | 55.86 | 56.125
13.3 | 48.65 | 52.3667 | 54.6 | 55.525 | 55.775
13.4 | 48.3 | 52.15 | 54.3667 | 55.25 | 55.525
13.5 | 48.05 | 51.8 | 54 | 55 | 55.25
13.6 | 47.8 | 51.5 | 53.76 | 54.75 | 54.9143
13.7 | 47.4 | 51.2 | 53.45 | 54.45 | 54.6667
13.8 | 47.1 | 50.95 | 53.175 | 54.1 | 54.4
13.9 | 46.8 | 50.6 | 52.85 | 53.82 | 54.075
14 | 46.52 | 50.2714 | 52.675 | 53.6 | 53.9125
14.1 | 46.2 | 50 | 52.425 | 53.3429 | 53.575
14.2 | 45.9429 | 49.675 | 52.075 | 53 | 53.325
14.3 | 45.7 | 49.3 | 51.825 | 52.78 | 53.0125
14.4 | 45.35 | 49.05 | 51.475 | 52.525 | 52.725
14.5 | 45.0857 | 48.82 | 51.24 | 52.2625 | 52.56
14.6 | 44.825 | 48.45 | 50.9 | 51.925 | 52.25
14.7 | 44.55 | 48.2 | 50.65 | 51.7625 | 51.9
14.8 | 44.175 | 47.9 | 50.3 | 51.38 | 51.65
14.9 | 43.9 | 47.6 | 50.05 | 51.2 | 51.3
15 | 43.6333 | 47.32 | 49.8 | 50.84 | 51.075
15.1 | 43.4 | 47.1 | 49.55 | 50.6 | 50.9
15.2 | 43.1 | 46.875 | 49.3 | 50.34 | 50.575
15.3 | 42.85 | 46.525 | 48.95 | 50.05 | 50.325
15.4 | 42.525 | 46.275 | 48.74 | 49.825 | 50.15
15.5 | 42.275 | 46.025 | 48.475 | 49.6 | 49.8
15.6 | 41.95 | 45.775 | 48.225 | 49.35 | 49.6
15.7 | 41.775 | 45.425 | 47.975 | 49.1 | 49.35
15.8 | 41.5 | 45.2333 | 47.725 | 48.8 | 49.1
15.9 | 41.175 | 44.95 | 47.45 | 48.6125 | 48.85
16 | 40.925 | 44.7 | 47.225 | 48.375 | 48.6
16.1 | 40.7 | 44.475 | 46.875 | 48.125 | 48.36
16.2 | 40.5 | 44.2 | 46.6333 | 47.875 | 48.175
16.3 | 40.25 | 43.925 | 46.5 | 47.625 | 47.9
16.4 | 40.0333 | 43.725 | 46.2 | 47.375 | 47.65
16.5 | 39.7 | 43.375 | 46 | 47.05 | 47.375
16.6 | 39.56 | 43.1333 | 45.65 | 46.9 | 47.1
16.7 | 39.325 | 42.9 | 45.4 | 46.6 | 46.9
16.8 | 38.975 | 42.65 | 45.15 | 46.3667 | 46.65
16.9 | 38.8 | 42.3 | 44.925 | 46.125 | 46.4333
17 | 38.475 | 42.15 | 44.66 | 45.875 | 46.2
17.1 | 38.225 | 41.8 | 44.5 | 45.625 | 45.95
17.2 | 38.0667 | 41.65 | 44.275 | 45.375 | 45.7429
17.3 | 37.85 | 41.4 | 44.025 | 45.225 | 45.525
17.4 | 37.6 | 41.15 | 43.75 | 45 | 45.225
17.5 | 37.325 | 41 | 43.525 | 44.65 | 45.075
17.6 | 37.1 | 40.7 | 43.275 | 44.5 | 44.825
17.7 | 36.8333 | 40.525 | 43.05 | 44.25 | 44.6375
17.8 | 36.7 | 40.175 | 42.88 | 44 | 44.45
17.9 | 36.45 | 40.025 | 42.6 | 43.775 | 44.1
18 | 36.2 | 39.775 | 42.4 | 43.6 | 43.95
18.1 | 36.05 | 39.6 | 42.15 | 43.4 | 43.7
18.2 | 35.8 | 39.325 | 42 | 43.2 | 43.475
18.3 | 35.46 | 39.075 | 41.75 | 42.95 | 43.225
18.4 | 35.325 | 38.9 | 41.5 | 42.7 | 43.075
18.5 | 35.1 | 38.7 | 41.26 | 42.55 | 42.825
18.6 | 34.9 | 38.35 | 41 | 42.3 | 42.575
18.7 | 34.675 | 38.2 | 40.8 | 42.05 | 42.425
18.8 | 34.45 | 37.96 | 40.6 | 41.825 | 42.2
18.9 | 34.2 | 37.7667 | 40.35 | 41.575 | 42
19 | 34 | 37.5 | 40.1 | 41.425 | 41.7
19.1 | 33.8 | 37.3333 | 39.92 | 41.1 | 41.45
19.2 | 33.5667 | 37.1 | 39.7714 | 40.925 | 41.3
19.3 | 33.35 | 36.925 | 39.52 | 40.8 | 41.1
19.4 | 33.1556 | 36.6 | 39.275 | 40.55 | 40.9
19.5 | 33.0375 | 36.45 | 39.0625 | 40.3 | 40.6667
19.6 | 32.825 | 36.2 | 38.875 | 40.15 | 40.525
19.7 | 32.575 | 36.05 | 38.74 | 39.9 | 40.275
19.8 | 32.35 | 35.8 | 38.4 | 39.7 | 40.025
19.9 | 32.1875 | 35.6 | 38.25 | 39.5 | 39.875
20 | 32.025 | 35.36 | 38 | 39.25 | 39.625
20.1 | 31.8375 | 35.2125 | 37.85 | 39.125 | 39.4333
20.2 | 31.6444 | 34.975 | 37.625 | 38.875 | 39.25
20.3 | 31.4 | 34.7625 | 37.4 | 38.625 | 39
20.4 | 31.25 | 34.575 | 37.225 | 38.4 | 38.8
20.5 | 30.96 | 34.325 | 36.975 | 38.225 | 38.6
20.6 | 30.8 | 34.175 | 36.825 | 38.1 | 38.475
20.7 | 30.625 | 33.9571 | 36.6667 | 37.85 | 38.225
20.8 | 30.4 | 33.75 | 36.4 | 37.7 | 38.075
20.9 | 30.24 | 33.6 | 36.25 | 37.45 | 37.825
21 | 30.1 | 33.35 | 36.05 | 37.24 | 37.675
21.1 | 29.85 | 33.22 | 35.85 | 37.2 | 37.5
21.2 | 29.7 | 33 | 35.6222 | 37 | 37.3143
21.3 | 29.5667 | 32.825 | 35.4375 | 36.8 | 37.1
21.4 | 29.3667 | 32.6 | 35.3125 | 36.5 | 36.95
21.5 | 29.3 | 32.5 | 35.0875 | 36.4 | 36.7111
21.6 | 29 | 32.3 | 34.925 | 36.15 | 36.5375
21.7 | 28.875 | 32.04 | 34.675 | 36 | 36.3625
21.8 | 28.625 | 31.9 | 34.5125 | 35.8286 | 36.175
21.9 | 28.5 | 31.75 | 34.3333 | 35.7 | 36.0125
22 | 28.325 | 31.6 | 34.2 | 35.4556 | 35.775
22.1 | 28.1 | 31.3 | 33.95 | 35.2625 | 35.6
22.2 | 27.925 | 31.2 | 33.7 | 35.1375 | 35.5
22.3 | 27.775 | 31 | 33.6 | 34.925 | 35.25
22.4 | 27.6 | 30.825 | 33.42 | 34.6875 | 35.1
22.5 | 27.48 | 30.675 | 33.275 | 34.6 | 34.9
22.6 | 27.325 | 30.425 | 33.025 | 34.4 | 34.72
22.7 | 27.2 | 30.3 | 32.88 | 34.14 | 34.575
22.8 | 26.95 | 30.12 | 32.75 | 34 | 34.325
22.9 | 26.8 | 29.975 | 32.5 | 33.85 | 34.175
23 | 26.65 | 29.725 | 32.4 | 33.65 | 34.05
23.1 | 26.5 | 29.575 | 32.2 | 33.5 | 33.8
23.2 | 26.3 | 29.425 | 32 | 33.3 | 33.65
23.3 | 26.1 | 29.26 | 31.82 | 33.16 | 33.5
23.4 | 26 | 29.05 | 31.675 | 33 | 33.35
23.5 | 25.8 | 28.9 | 31.45 | 32.8 | 33.2
23.6 | 25.7 | 28.75 | 31.3 | 32.7 | 33
23.7 | 25.5 | 28.6 | 31.175 | 32.5 | 32.75
23.8 | 25.375 | 28.4 | 30.95 | 32.35 | 32.6
23.9 | 25.225 | 28.3 | 30.84 | 32.1 | 32.475
24 | 25.075 | 28.1333 | 30.68 | 32 | 32.4
24.1 | 24.925 | 28 | 30.525 | 31.8 | 32.2
24.2 | 24.8 | 27.8 | 30.4 | 31.6 | 31.95
24.3 | 24.6 | 27.6 | 30.2 | 31.5 | 31.8
24.4 | 24.5 | 27.45 | 30 | 31.375 | 31.675
24.5 | 24.4 | 27.325 | 29.8714 | 31.125 | 31.5
24.6 | 24.2 | 27.175 | 29.6556 | 31 | 31.35
24.7 | 24.1 | 27.025 | 29.5375 | 30.825 | 31.2
24.8 | 23.9 | 26.8 | 29.4125 | 30.72 | 31.0667
24.9 | 23.7286 | 26.6667 | 29.1875 | 30.5 | 30.875
25 | 23.62 | 26.5333 | 29.05 | 30.375 | 30.725
25.1 | 23.5 | 26.4 | 29 | 30.1857 | 30.6
25.2 | 23.34 | 26.225 | 28.7167 | 30.05 | 30.4333
25.3 | 23.2 | 26.075 | 28.55 | 29.9333 | 30.3
25.4 | 23.05 | 25.925 | 28.4 | 29.7833 | 30.1333
25.5 | 22.9 | 25.775 | 28.3 | 29.6167 | 29.9
25.6 | 22.75 | 25.625 | 28.1 | 29.475 | 29.8
25.7 | 22.6333 | 25.5 | 28 | 29.2625 | 29.62
25.8 | 22.52 | 25.34 | 27.8 | 29.1875 | 29.5
25.9 | 22.4 | 25.2 | 27.7 | 29 | 29.4
26 | 22.25 | 25.1 | 27.525 | 28.8667 | 29.1875
26.1 | 22.1667 | 24.96 | 27.375 | 28.7 | 29.04
26.2 | 22 | 24.75 | 27.225 | 28.6 | 28.9
26.3 | 21.8333 | 24.6 | 27.1 | 28.475 | 28.76
26.4 | 21.7 | 24.5 | 26.95 | 28.225 | 28.6
26.5 | 21.6 | 24.325 | 26.8 | 28.1 | 28.5
26.6 | 21.475 | 24.275 | 26.6667 | 28 | 28.3
26.7 | 21.325 | 24.1 | 26.52 | 27.875 | 28.2
26.8 | 21.2 | 23.975 | 26.4 | 27.725 | 28.075
26.9 | 21.025 | 23.84 | 26.3 | 27.575 | 27.9
27 | 20.975 | 23.8 | 26.1 | 27.425 | 27.775
27.1 | 20.84 | 23.6 | 25.9333 | 27.3 | 27.625
27.2 | 20.7 | 23.4 | 25.7875 | 27.2 | 27.5
27.3 | 20.65 | 23.3333 | 25.6625 | 27 | 27.35
27.4 | 20.5 | 23.2 | 25.5375 | 26.9 | 27.2
27.5 | 20.3333 | 23.05 | 25.475 | 26.75 | 27.05
27.6 | 20.2 | 22.925 | 25.3333 | 26.6 | 26.9
27.7 | 20.125 | 22.8111 | 25.2 | 26.5 | 26.8
27.8 | 20 | 22.675 | 25.025 | 26.3 | 26.6
27.9 | 19.9625 | 22.5 | 24.8875 | 26.2 | 26.525
28 | 19.8375 | 22.4 | 24.7625 | 26.075 | 26.4
28.1 | 19.7222 | 22.2 | 24.6375 | 25.925 | 26.3
28.2 | 19.6 | 22.1 | 24.5125 | 25.775 | 26.1
28.3 | 19.475 | 22 | 24.3875 | 25.62 | 26
28.4 | 19.35 | 21.9 | 24.2667 | 25.5667 | 25.875
28.5 | 19.225 | 21.75 | 24.1429 | 25.4 | 25.725
28.6 | 19.1 | 21.6 | 24 | 25.3 | 25.6167
28.7 | 19 | 21.575 | 23.875 | 25.1 | 25.5
28.8 | 18.86 | 21.4 | 23.7 | 25 | 25.4
28.9 | 18.8 | 21.275 | 23.6 | 24.9 | 25.2
29 | 18.675 | 21.2 | 23.5 | 24.8 | 25.0444
29.1 | 18.525 | 21.08 | 23.4 | 24.65 | 25
29.2 | 18.475 | 20.95 | 23.2 | 24.5 | 24.9
29.3 | 18.3 | 20.8 | 23.1 | 24.3833 | 24.775
29.4 | 18.2 | 20.7 | 23 | 24.2375 | 24.65
29.5 | 18.1 | 20.6 | 22.9 | 24.1125 | 24.5143
29.6 | 18 | 20.5 | 22.8 | 24.0429 | 24.4
29.7 | 17.9 | 20.325 | 22.6 | 23.9 | 24.25
29.8 | 17.8 | 20.275 | 22.5 | 23.775 | 24.1
29.9 | 17.7 | 20.1 | 22.4 | 23.7 | 24
30 | 17.6 | 20 | 22.275 | 23.575 | 23.9

## SSD120 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 42.3 | 48 | 56 | 62.9 | 65.7
0.1 | 47.8 | 53.05 | 61.05 | 67.34 | 69.85
0.2 | 56.075 | 60.82 | 68 | 73 | 75.7
0.3 | 66.95 | 70.775 | 75.875 | 80.7 | 82.8
0.4 | 75.075 | 78.36 | 82.45 | 86.3 | 88
0.5 | 81.7 | 83.8 | 87.475 | 90.3 | 91.6
0.6 | 86.525 | 88.15 | 91.075 | 93.25 | 94.4
0.7 | 90.3 | 91.6 | 93.8 | 95.5 | 96.225
0.8 | 93.15 | 94.225 | 95.8 | 97.15 | 97.5
0.9 | 95.325 | 95.9 | 97.325 | 98.3 | 98.775
1 | 97 | 97.6 | 98.475 | 99.1 | 99.35
1.1 | 98.3 | 98.5 | 99.2333 | 99.6 | 99.725
1.2 | 99.275 | 99.25 | 99.65 | 99.85 | 99.9
1.3 | 99.75 | 99.75 | 99.9 | 100 | 100
1.4 | 100 | 100 | 100 | 100 | 100
1.5 | 100 | 100 | 100 | 100 | 99.98
1.6 | 100 | 100 | 100 | 99.8 | 99.85
1.7 | 99.9 | 99.7 | 99.75 | 99.6 | 99.7
1.8 | 99.72 | 99.5 | 99.5 | 99.42 | 99.45
1.9 | 99.5667 | 99.3 | 99.35 | 99.2 | 99.2
2 | 99.24 | 98.95 | 99.1 | 98.875 | 98.9
2.1 | 98.85 | 98.675 | 98.75 | 98.625 | 98.5
2.2 | 98.45 | 98.525 | 98.4 | 98.25 | 98.25
2.3 | 98.05 | 98.1833 | 98.1667 | 97.925 | 98
2.4 | 97.625 | 97.9125 | 97.7667 | 97.66 | 97.65
2.5 | 97.1667 | 97.525 | 97.4 | 97.35 | 97.3
2.6 | 96.75 | 97.1 | 97.075 | 97 | 97.05
2.7 | 96.2 | 96.65 | 96.75 | 96.65 | 96.65
2.8 | 95.75 | 96.34 | 96.475 | 96.4 | 96.3
2.9 | 95.3667 | 95.9667 | 96.225 | 96.05 | 96
3 | 94.9 | 95.5857 | 95.7667 | 95.8 | 95.72
3.1 | 94.5 | 95.2333 | 95.475 | 95.45 | 95.4333
3.2 | 93.95 | 94.9333 | 95.05 | 95.1 | 95
3.3 | 93.6 | 94.4 | 94.75 | 94.75 | 94.675
3.4 | 93.05 | 94.1 | 94.45 | 94.52 | 94.425
3.5 | 92.7 | 93.64 | 94.1 | 94.1333 | 94.1
3.6 | 92.225 | 93.36 | 93.7 | 93.8667 | 93.75
3.7 | 91.825 | 93 | 93.4 | 93.475 | 93.5
3.8 | 91.4 | 92.54 | 93.05 | 93.225 | 93.15
3.9 | 90.8 | 92.1125 | 92.7 | 92.825 | 92.9
4 | 90.48 | 91.775 | 92.4 | 92.55 | 92.4667
4.1 | 90.0333 | 91.45 | 92.15 | 92.225 | 92.2
4.2 | 89.45 | 90.9625 | 91.7125 | 91.675 | 91.75
4.3 | 89.025 | 90.5778 | 91.3375 | 91.425 | 91.5
4.4 | 88.625 | 90.3429 | 91 | 91.15 | 91.15
4.5 | 88.3 | 89.9 | 90.68 | 90.84 | 90.8
4.6 | 87.8 | 89.64 | 90.2857 | 90.5667 | 90.6
4.7 | 87.35 | 89.25 | 90.025 | 90.24 | 90.32
4.8 | 86.85 | 88.75 | 89.75 | 90 | 89.95
4.9 | 86.25 | 88.425 | 89.25 | 89.5 | 89.55
5 | 85.925 | 87.95 | 88.98 | 89.1 | 89.275
5.1 | 85.5 | 87.48 | 88.65 | 88.9 | 88.95
5.2 | 85 | 87.1 | 88.3 | 88.65 | 88.7
5.3 | 84.56 | 86.7 | 87.88 | 88.15 | 88.3
5.4 | 84.24 | 86.5 | 87.625 | 87.825 | 87.9
5.5 | 83.7667 | 86.15 | 87.32 | 87.6 | 87.65
5.6 | 83.25 | 85.825 | 87.025 | 87.35 | 87.3
5.7 | 82.8333 | 85.475 | 86.65 | 86.9 | 86.9
5.8 | 82.4 | 85.075 | 86.25 | 86.6 | 86.6
5.9 | 82 | 84.75 | 85.9 | 86.35 | 86.375
6 | 81.7 | 84.425 | 85.55 | 86 | 86.05
6.1 | 81.3 | 84 | 85.3 | 85.65 | 85.775
6.2 | 80.7 | 83.65 | 84.92 | 85.425 | 85.35
6.3 | 80.1667 | 83.2 | 84.625 | 85.05 | 85.0667
6.4 | 79.9 | 82.8 | 84.2714 | 84.725 | 84.85
6.5 | 79.4667 | 82.425 | 83.9667 | 84.475 | 84.525
6.6 | 79.175 | 82.15 | 83.55 | 84.15 | 84.175
6.7 | 78.575 | 81.8 | 83.225 | 83.7667 | 83.8
6.8 | 78.3 | 81.4 | 83 | 83.45 | 83.5
6.9 | 77.9 | 81.05 | 82.55 | 83.1 | 83.2
7 | 77.5 | 80.8 | 82.24 | 82.8 | 83
7.1 | 77.175 | 80.45 | 82.06 | 82.5 | 82.56
7.2 | 76.6 | 80.05 | 81.6 | 82.1667 | 82.25
7.3 | 76.2667 | 79.8 | 81.2556 | 81.9 | 81.975
7.4 | 76.025 | 79.3 | 80.95 | 81.6333 | 81.725
7.5 | 75.45 | 78.9 | 80.5 | 81.35 | 81.275
7.6 | 74.9875 | 78.55 | 80.4 | 81.05 | 81.025
7.7 | 74.7444 | 78.25 | 80.0667 | 80.725 | 80.675
7.8 | 74.325 | 77.95 | 79.7 | 80.375 | 80.425
7.9 | 73.7429 | 77.45 | 79.36 | 79.95 | 80.05
8 | 73.3 | 77.1 | 79.0667 | 79.675 | 79.74
8.1 | 72.95 | 76.85 | 78.82 | 79.4 | 79.5
8.2 | 72.6 | 76.6 | 78.45 | 79.0333 | 79.1
8.3 | 72.25 | 76.25 | 78.2 | 78.6667 | 78.8
8.4 | 71.8 | 75.75 | 77.875 | 78.4 | 78.55
8.5 | 71.36 | 75.275 | 77.45 | 78.15 | 78.225
8.6 | 71.05 | 75 | 77.08 | 77.9 | 77.9
8.7 | 70.6 | 74.725 | 76.8 | 77.65 | 77.725
8.8 | 70.3 | 74.4 | 76.5 | 77.325 | 77.4
8.9 | 69.9 | 74.1667 | 76.2333 | 76.975 | 77.05
9 | 69.5333 | 73.7 | 75.9 | 76.625 | 76.8
9.1 | 69.1 | 73.325 | 75.5 | 76.28 | 76.45
9.2 | 68.65 | 72.9 | 75.3 | 75.975 | 76.1
9.3 | 68.3 | 72.5667 | 75 | 75.65 | 75.8
9.4 | 67.94 | 72.2375 | 74.575 | 75.375 | 75.425
9.5 | 67.66 | 71.8625 | 74.375 | 75.05 | 75.275
9.6 | 67.3 | 71.525 | 74.125 | 74.9 | 75.025
9.7 | 66.92 | 71.2125 | 73.78 | 74.45 | 74.675
9.8 | 66.6 | 70.8375 | 73.5 | 74.2 | 74.35
9.9 | 66.1 | 70.5667 | 73 | 73.9 | 74.2
10 | 65.7 | 70.25 | 72.75 | 73.65 | 73.75
10.1 | 65.3 | 69.9 | 72.5 | 73.3 | 73.5
10.2 | 65.0333 | 69.62 | 72.225 | 73.05 | 73.25
10.3 | 64.6333 | 69.425 | 71.85 | 72.7 | 73
10.4 | 64.25 | 68.975 | 71.38 | 72.45 | 72.6333
10.5 | 63.925 | 68.65 | 71.1 | 72.3 | 72.35
10.6 | 63.6 | 68.38 | 70.95 | 71.9 | 72
10.7 | 63.25 | 68 | 70.6 | 71.7333 | 71.7571
10.8 | 62.85 | 67.5 | 70.275 | 71.3 | 71.35
10.9 | 62.5 | 67.25 | 70.025 | 71.075 | 71.075
11 | 62.2333 | 67 | 69.76 | 70.825 | 70.85
11.1 | 61.9 | 66.66 | 69.425 | 70.5 | 70.625
11.2 | 61.6 | 66.425 | 69.075 | 70.2333 | 70.4167
11.3 | 61.1 | 66.175 | 68.825 | 69.8667 | 70.1125
11.4 | 60.7 | 65.9 | 68.55 | 69.6 | 69.725
11.5 | 60.375 | 65.475 | 68.225 | 69.3667 | 69.48
11.6 | 60.125 | 65.14 | 67.975 | 69.15 | 69.225
11.7 | 59.775 | 64.74 | 67.7 | 68.8 | 68.98
11.8 | 59.55 | 64.5 | 67.45 | 68.6 | 68.65
11.9 | 59.175 | 64.25 | 67.1 | 68.25 | 68.475
12 | 58.88 | 63.95 | 66.78 | 67.925 | 68.15
12.1 | 58.5 | 63.65 | 66.5 | 67.675 | 67.875
12.2 | 58.35 | 63.3 | 66.25 | 67.35 | 67.625
12.3 | 57.9 | 62.8556 | 65.9 | 67.075 | 67.3
12.4 | 57.5 | 62.5875 | 65.65 | 66.75 | 67.05
12.5 | 57.2 | 62.2571 | 65.4 | 66.475 | 66.72
12.6 | 56.95 | 62.025 | 64.975 | 66.34 | 66.55
12.7 | 56.6667 | 61.86 | 64.725 | 65.975 | 66.325
12.8 | 56.3 | 61.45 | 64.45 | 65.8 | 65.95
12.9 | 55.95 | 61.2 | 64.15 | 65.5 | 65.7
13 | 55.7 | 60.7 | 63.9 | 65.25 | 65.425
13.1 | 55.25 | 60.5 | 63.65 | 64.9 | 65.15
13.2 | 55.1 | 60.2 | 63.5 | 64.6 | 64.975
13.3 | 54.95 | 60.05 | 63.15 | 64.35 | 64.7
13.4 | 54.525 | 59.7333 | 62.9 | 64.2 | 64.4
13.5 | 54.075 | 59.3 | 62.5857 | 63.975 | 64.1
13.6 | 53.825 | 59.075 | 62.3222 | 63.7 | 63.8
13.7 | 53.625 | 58.8 | 62.15 | 63.3 | 63.55
13.8 | 53.3 | 58.575 | 61.825 | 63.05 | 63.2
13.9 | 52.9 | 58.14 | 61.6 | 62.875 | 63.05
14 | 52.725 | 57.9 | 61.3 | 62.625 | 62.8
14.1 | 52.35 | 57.65 | 60.96 | 62.28 | 62.5
14.2 | 52.0333 | 57.45 | 60.75 | 62.04 | 62.2
14.3 | 51.9 | 57.1 | 60.5 | 61.775 | 62.1
14.4 | 51.55 | 56.85 | 60.075 | 61.525 | 61.75
14.5 | 51.325 | 56.6 | 59.94 | 61.2667 | 61.5
14.6 | 51.05 | 56.35 | 59.7 | 60.98 | 61.2
14.7 | 50.7 | 55.9 | 59.4 | 60.7333 | 61.05
14.8 | 50.35 | 55.56 | 59.2 | 60.5 | 60.7
14.9 | 50.1 | 55.42 | 58.95 | 60.3 | 60.5
15 | 49.75 | 55.06 | 58.5 | 60 | 60.2
15.1 | 49.4 | 54.825 | 58.25 | 59.65 | 60
15.2 | 49.25 | 54.6 | 58 | 59.4 | 59.725
15.3 | 49 | 54.36 | 57.85 | 59.06 | 59.4333
15.4 | 48.775 | 54.025 | 57.6 | 58.925 | 59.1
15.5 | 48.525 | 53.775 | 57.25 | 58.575 | 58.875
15.6 | 48.2 | 53.5 | 57.1 | 58.425 | 58.7
15.7 | 48.05 | 53.2333 | 56.875 | 58.275 | 58.4667
15.8 | 47.7333 | 52.975 | 56.55 | 57.95 | 58.2
15.9 | 47.5 | 52.725 | 56.275 | 57.7 | 57.95
16 | 47.15 | 52.575 | 56.025 | 57.525 | 57.7
16.1 | 46.9667 | 52.34 | 55.7333 | 57.25 | 57.5
16.2 | 46.6333 | 52 | 55.58 | 57 | 57.3
16.3 | 46.3 | 51.75 | 55.35 | 56.825 | 57.15
16.4 | 46.1 | 51.4 | 55.05 | 56.575 | 56.85
16.5 | 45.9 | 51.15 | 54.8 | 56.225 | 56.55
16.6 | 45.65 | 50.9 | 54.6 | 56 | 56.3
16.7 | 45.4 | 50.65 | 54.3 | 55.84 | 56.05
16.8 | 45.15 | 50.42 | 54.1 | 55.5667 | 55.8
16.9 | 44.9 | 50.175 | 53.85 | 55.3625 | 55.625
17 | 44.65 | 49.925 | 53.6 | 55.1333 | 55.4
17.1 | 44.5 | 49.575 | 53.35 | 54.8286 | 55.1667
17.2 | 44.125 | 49.35 | 53.2 | 54.5778 | 54.9
17.3 | 43.8 | 49.1 | 52.95 | 54.4 | 54.725
17.4 | 43.55 | 48.95 | 52.72 | 54.2 | 54.5
17.5 | 43.4 | 48.7 | 52.475 | 54.05 | 54.35
17.6 | 43.075 | 48.55 | 52.2 | 53.8 | 54.1
17.7 | 42.925 | 48.3 | 52 | 53.575 | 53.85
17.8 | 42.6 | 47.96 | 51.6667 | 53.3 | 53.6
17.9 | 42.3 | 47.725 | 51.475 | 53.05 | 53.45
18 | 42.175 | 47.475 | 51.225 | 52.8 | 53.1
18.1 | 41.975 | 47.225 | 51 | 52.7 | 52.9
18.2 | 41.65 | 47.075 | 50.725 | 52.45 | 52.7
18.3 | 41.5 | 46.925 | 50.5 | 52.1 | 52.45
18.4 | 41.2 | 46.6667 | 50.3 | 51.9 | 52.275
18.5 | 41 | 46.42 | 50.15 | 51.7 | 52.025
18.6 | 40.775 | 46.25 | 49.9 | 51.46 | 51.775
18.7 | 40.6 | 45.925 | 49.6667 | 51.225 | 51.525
18.8 | 40.35 | 45.675 | 49.4 | 51.075 | 51.4
18.9 | 40.0667 | 45.525 | 49.2 | 50.825 | 51.1333
19 | 39.8 | 45.275 | 49 | 50.5833 | 50.8
19.1 | 39.7 | 44.9667 | 48.85 | 50.375 | 50.7
19.2 | 39.425 | 44.8 | 48.5429 | 50.2 | 50.45
19.3 | 39.25 | 44.5333 | 48.3 | 49.95 | 50.3
19.4 | 39.025 | 44.375 | 48.06 | 49.6857 | 50
19.5 | 38.775 | 44.1 | 47.9 | 49.5 | 49.82
19.6 | 38.6 | 43.9 | 47.675 | 49.2857 | 49.6
19.7 | 38.4 | 43.6 | 47.425 | 49.1 | 49.425
19.8 | 38.18 | 43.4 | 47.2 | 48.85 | 49.1333
19.9 | 37.98 | 43.22 | 46.95 | 48.6571 | 49.0333
20 | 37.7333 | 43 | 46.7 | 48.375 | 48.8
20.1 | 37.575 | 42.72 | 46.65 | 48.125 | 48.5
20.2 | 37.35 | 42.575 | 46.4 | 47.975 | 48.3
20.3 | 37.075 | 42.325 | 46.175 | 47.85 | 48.1
20.4 | 36.98 | 42.175 | 46 | 47.6 | 47.9
20.5 | 36.65 | 41.925 | 45.82 | 47.4 | 47.7
20.6 | 36.5 | 41.625 | 45.5667 | 47.2 | 47.56
20.7 | 36.35 | 41.48 | 45.3333 | 47.075 | 47.3
20.8 | 36.2 | 41.25 | 45.05 | 46.825 | 47.2
20.9 | 35.85 | 41.1 | 44.9 | 46.6 | 47.025
21 | 35.7 | 40.95 | 44.725 | 46.425 | 46.7
21.1 | 35.56 | 40.7 | 44.475 | 46.18 | 46.6
21.2 | 35.375 | 40.55 | 44.225 | 46.05 | 46.35
21.3 | 35.225 | 40.3 | 44.1 | 45.8667 | 46.2
21.4 | 34.975 | 40.06 | 43.925 | 45.7 | 45.975
21.5 | 34.825 | 39.9 | 43.7 | 45.45 | 45.725
21.6 | 34.575 | 39.675 | 43.55 | 45.2 | 45.575
21.7 | 34.425 | 39.5333 | 43.3 | 45.0143 | 45.325
21.8 | 34.2 | 39.3 | 43.05 | 44.775 | 45.175
21.9 | 34.025 | 39.15 | 42.9 | 44.6 | 45
22 | 33.88 | 38.88 | 42.75 | 44.425 | 44.8
22.1 | 33.7333 | 38.7 | 42.5 | 44.275 | 44.55
22.2 | 33.6 | 38.4 | 42.3 | 44.025 | 44.4
22.3 | 33.35 | 38.25 | 42.1 | 43.8375 | 44.2
22.4 | 33.1 | 38.1 | 41.9333 | 43.6444 | 44.1
22.5 | 32.975 | 37.95 | 41.8 | 43.4889 | 43.9
22.6 | 32.8 | 37.7 | 41.625 | 43.3429 | 43.6
22.7 | 32.575 | 37.56 | 41.475 | 43.0333 | 43.45
22.8 | 32.425 | 37.4 | 41.225 | 42.8143 | 43.3
22.9 | 32.3 | 37.3 | 41 | 42.6714 | 43.1
23 | 32.075 | 37.1 | 40.8667 | 42.5 | 42.925
23.1 | 31.9 | 36.8 | 40.65 | 42.35 | 42.7
23.2 | 31.75 | 36.7 | 40.5 | 42.2 | 42.5
23.3 | 31.5 | 36.5 | 40.35 | 42.05 | 42.3
23.4 | 31.325 | 36.225 | 40.1 | 41.74 | 42.125
23.5 | 31.2 | 36.075 | 40 | 41.56 | 42
23.6 | 31.025 | 35.9 | 39.8 | 41.4 | 41.75
23.7 | 30.8625 | 35.6 | 39.56 | 41.275 | 41.6
23.8 | 30.6875 | 35.44 | 39.4 | 41.1 | 41.475
23.9 | 30.5625 | 35.3 | 39.16 | 40.9 | 41.225
24 | 30.38 | 35.15 | 39.025 | 40.75 | 41.075
24.1 | 30.25 | 35.05 | 38.875 | 40.5 | 40.9
24.2 | 30.1 | 34.92 | 38.7 | 40.3 | 40.8
24.3 | 29.9 | 34.7 | 38.6 | 40.125 | 40.55
24.4 | 29.7 | 34.425 | 38.4 | 39.975 | 40.4
24.5 | 29.6 | 34.3 | 38.175 | 39.825 | 40.1667
24.6 | 29.5 | 34.175 | 38 | 39.7167 | 40.05
24.7 | 29.3 | 34 | 37.775 | 39.5333 | 39.925
24.8 | 29.1 | 33.8167 | 37.6 | 39.3 | 39.7
24.9 | 28.9 | 33.6167 | 37.5 | 39.2 | 39.5
25 | 28.8 | 33.46 | 37.3 | 39 | 39.3125
25.1 | 28.6 | 33.3222 | 37.1 | 38.775 | 39.175
25.2 | 28.5 | 33.1 | 36.95 | 38.64 | 39.025
25.3 | 28.4 | 32.925 | 36.8 | 38.48 | 38.7875
25.4 | 28.2 | 32.8 | 36.65 | 38.35 | 38.625
25.5 | 28 | 32.65 | 36.5 | 38.2 | 38.4875
25.6 | 27.9 | 32.5 | 36.275 | 37.95 | 38.34
25.7 | 27.775 | 32.35 | 36.125 | 37.7429 | 38.2
25.8 | 27.6 | 32.2 | 35.94 | 37.65 | 38
25.9 | 27.3875 | 32 | 35.8 | 37.45 | 37.8
26 | 27.2625 | 31.9 | 35.65 | 37.325 | 37.65
26.1 | 27.1375 | 31.7 | 35.45 | 37.1 | 37.52
26.2 | 27.0222 | 31.6 | 35.325 | 36.92 | 37.3
26.3 | 26.8857 | 31.35 | 35.175 | 36.8 | 37.225
26.4 | 26.775 | 31.2 | 34.925 | 36.6 | 37.075
26.5 | 26.64 | 31.075 | 34.8 | 36.4333 | 36.825
26.6 | 26.4 | 30.925 | 34.625 | 36.3 | 36.675
26.7 | 26.3 | 30.8 | 34.4875 | 36.2 | 36.5
26.8 | 26.2 | 30.7 | 34.325 | 36.0333 | 36.4
26.9 | 26.05 | 30.55 | 34.1375 | 35.9 | 36.25
27 | 25.8 | 30.3 | 34 | 35.75 | 36.1
27.1 | 25.7 | 30.2 | 33.9444 | 35.6 | 35.95
27.2 | 25.6 | 29.975 | 33.7286 | 35.4 | 35.8333
27.3 | 25.4 | 29.85 | 33.6 | 35.3 | 35.6
27.4 | 25.275 | 29.725 | 33.4 | 35.14 | 35.5
27.5 | 25.1 | 29.6 | 33.3 | 34.925 | 35.3429
27.6 | 25 | 29.475 | 33.15 | 34.7375 | 35.225
27.7 | 24.9 | 29.2625 | 32.9 | 34.7 | 35
27.8 | 24.8 | 29.1375 | 32.8 | 34.4333 | 34.825
27.9 | 24.7 | 28.925 | 32.6 | 34.3 | 34.675
28 | 24.4875 | 28.78 | 32.4 | 34.1889 | 34.525
28.1 | 24.3667 | 28.6 | 32.3 | 34.05 | 34.4
28.2 | 24.25 | 28.475 | 32.2 | 33.9 | 34.225
28.3 | 24.125 | 28.325 | 32 | 33.8 | 34.1
28.4 | 24 | 28.25 | 31.9 | 33.55 | 34
28.5 | 23.8 | 28.1 | 31.76 | 33.425 | 33.7667
28.6 | 23.7 | 27.95 | 31.525 | 33.3 | 33.6
28.7 | 23.6 | 27.8 | 31.4 | 33.175 | 33.5
28.8 | 23.5 | 27.7 | 31.325 | 33.025 | 33.3
28.9 | 23.425 | 27.55 | 31.1 | 32.875 | 33.15
29 | 23.3 | 27.4 | 31 | 32.7 | 33.05
29.1 | 23.1 | 27.35 | 30.8 | 32.6 | 32.9
29.2 | 23 | 27.1 | 30.7 | 32.5 | 32.75
29.3 | 22.9 | 27 | 30.6 | 32.35 | 32.7
29.4 | 22.8 | 26.8 | 30.5 | 32.15 | 32.45
29.5 | 22.65 | 26.75 | 30.36 | 31.95 | 32.3
29.6 | 22.5 | 26.6 | 30.1 | 31.9 | 32.2
29.7 | 22.38 | 26.45 | 30 | 31.78 | 32.025
29.8 | 22.2143 | 26.3125 | 29.9 | 31.625 | 31.9
29.9 | 22.1 | 26.1833 | 29.775 | 31.5 | 31.85
30 | 22.025 | 26.075 | 29.6 | 31.325 | 31.625

## Wedge 30 Lateral Profiles
Lateral Profiles
Wedge 30
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.5 cm FieldSize:4x4 cm2 | Depth:1.5 cm FieldSize:10x10 cm2 | Depth:1.5 cm FieldSize:20x20 cm2 | Depth:1.5 cm FieldSize:30x30 cm2 | Depth:1.5 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [18 empty] | 17.98 | 19.48 |  |  |  | 25.3 | 26.9333
-22.8 | [18 empty] | 18.2 | 19.72 |  |  |  | 25.625 | 27.28
-22.7 | [18 empty] | 18.36 | 19.9333 |  |  |  | 26.0167 | 27.5
-22.6 | [18 empty] | 18.6333 | 20.1 |  |  |  | 26.35 | 27.925
-22.5 | [18 empty] | 18.9429 | 20.35 |  |  |  | 26.62 | 28.275
-22.4 | [18 empty] | 19.1571 | 20.6333 |  |  |  | 26.96 | 28.52
-22.3 | [18 empty] | 19.4 | 20.8833 |  |  |  | 27.38 | 28.9
-22.2 | [18 empty] | 19.66 | 21.1333 |  |  |  | 27.7333 | 29.3143
-22.1 | [18 empty] | 19.92 | 21.5167 |  |  |  | 28 | 29.68
-22 | [18 empty] | 20.1857 | 21.7429 |  |  |  | 28.4 | 30.02
-21.9 | [18 empty] | 20.5 | 22.0333 |  |  |  | 28.75 | 30.3667
-21.8 | [18 empty] | 20.8 | 22.2 |  |  |  | 29.16 | 30.72
-21.7 | [18 empty] | 21.05 | 22.5 |  |  |  | 29.55 | 31.16
-21.6 | [18 empty] | 21.3833 | 22.8 |  |  |  | 30.02 | 31.6
-21.5 | [18 empty] | 21.6714 | 23.2333 |  |  |  | 30.3 | 31.9
-21.4 | [13 empty] | 13.9 | 15.2429 |  |  |  | 22 | 23.5167 |  |  |  | 30.66 | 32.28
-21.3 | [13 empty] | 14.04 | 15.4 |  |  |  | 22.3333 | 23.7667 |  |  |  | 31.2 | 32.66
-21.2 | [13 empty] | 14.2 | 15.6143 |  |  |  | 22.6333 | 24.0857 |  |  |  | 31.4667 | 33.15
-21.1 | [13 empty] | 14.375 | 15.8571 |  |  |  | 22.9429 | 24.425 |  |  |  | 31.9 | 33.6
-21 | [13 empty] | 14.6333 | 16.0833 |  |  |  | 23.35 | 24.8333 |  |  |  | 32.4 | 34.05
-20.9 | [13 empty] | 14.8 | 16.3 |  |  |  | 23.6333 | 25.1 |  |  |  | 32.85 | 34.6
-20.8 | [13 empty] | 15.025 | 16.4714 |  |  |  | 24.0571 | 25.4667 |  |  |  | 33.48 | 35.0333
-20.7 | [8 empty] | 11.5029 | 12.86 |  |  |  | 15.25 | 16.6667 |  |  |  | 24.3667 | 25.8333 |  |  |  | 34.1 | 35.5667
-20.6 | [8 empty] | 11.7029 | 13 |  |  |  | 15.475 | 16.9 |  |  |  | 24.7 | 26.2143 |  |  |  | 34.6 | 36.1
-20.5 | [8 empty] | 11.853 | 13.2 |  |  |  | 15.74 | 17.15 |  |  |  | 25.05 | 26.6 |  |  |  | 35.25 | 36.7667
-20.4 | [8 empty] | 11.978 | 13.375 |  |  |  | 15.96 | 17.42 |  |  |  | 25.4667 | 26.94 |  |  |  | 36.1 | 37.6
-20.3 | [8 empty] | 12.2031 | 13.6 |  |  |  | 16.22 | 17.5857 |  |  |  | 25.82 | 27.4 |  |  |  | 37.05 | 38.6
-20.2 |  |  |  | 10.44 | 11.825 |  |  |  | 12.3031 | 13.7 |  |  |  | 16.475 | 17.8667 |  |  |  | 26.2143 | 27.7333 |  |  |  | 38.4333 | 39.7667
-20.1 |  |  |  | 10.56 | 11.96 |  |  |  | 12.5031 | 13.82 |  |  |  | 16.7 | 18.22 |  |  |  | 26.7143 | 28.05 |  |  |  | 40.1 | 41.4
-20 |  |  |  | 10.7 | 12.18 |  |  |  | 12.7032 | 14.025 |  |  |  | 16.925 | 18.42 |  |  |  | 27.05 | 28.5 |  |  |  | 42.7 | 44.06
-19.9 |  |  |  | 10.9 | 12.3 |  |  |  | 12.8282 | 14.2 |  |  |  | 17.25 | 18.7 |  |  |  | 27.5125 | 29 |  |  |  | 46.02 | 47.42
-19.8 |  |  |  | 11 | 12.4 |  |  |  | 12.9633 | 14.4 |  |  |  | 17.52 | 18.9143 |  |  |  | 27.95 | 29.4667 |  |  |  | 51.6 | 52.38
-19.7 |  |  |  | 11.1 | 12.575 |  |  |  | 13.1033 | 14.6 |  |  |  | 17.7667 | 19.2667 |  |  |  | 28.45 | 29.7571 |  |  |  | 58.76 | 59.4
-19.6 |  |  |  | 11.3 | 12.7 |  |  |  | 13.3784 | 14.8 |  |  |  | 18.04 | 19.4857 |  |  |  | 28.8571 | 30.34 |  |  |  | 67.3 | 67.2
-19.5 |  |  |  | 11.4 | 12.8333 |  |  |  | 13.5284 | 15 |  |  |  | 18.45 | 19.85 |  |  |  | 29.32 | 30.8286 |  |  |  | 76 | 75.6
-19.4 |  |  |  | 11.625 | 13 |  |  |  | 13.7284 | 15.175 |  |  |  | 18.6667 | 20.1571 |  |  |  | 29.8 | 31.2667 |  |  |  | 84.8333 | 84.82
-19.3 |  |  |  | 11.775 | 13.2 |  |  |  | 13.9785 | 15.4 |  |  |  | 19.05 | 20.4 |  |  |  | 30.34 | 31.7875 |  |  |  | 93.74 | 93
-19.2 |  |  |  | 11.95 | 13.35 |  |  |  | 14.2036 | 15.7 |  |  |  | 19.375 | 20.8 |  |  |  | 30.9286 | 32.3857 |  |  |  | 101 | 100.14
-19.1 |  |  |  | 12.025 | 13.48 |  |  |  | 14.4286 | 15.85 |  |  |  | 19.725 | 21.08 |  |  |  | 31.5 | 32.9333 |  |  |  | 106.46 | 105.4
-19 |  |  |  | 12.175 | 13.625 |  |  |  | 14.6537 | 16.075 |  |  |  | 20.05 | 21.4667 |  |  |  | 32.3 | 33.7143 |  |  |  | 110.04 | 108.9
-18.9 |  |  |  | 12.3333 | 13.85 |  |  |  | 14.9037 | 16.225 |  |  |  | 20.36 | 21.8 |  |  |  | 33.1429 | 34.5 |  |  |  | 112.3667 | 111.1667
-18.8 |  |  |  | 12.6 | 14 |  |  |  | 15.1638 | 16.5 |  |  |  | 20.75 | 22.1333 |  |  |  | 34.2 | 35.6 |  |  |  | 113.9 | 112.6667
-18.7 |  |  |  | 12.72 | 14.2 |  |  |  | 15.3288 | 16.8 |  |  |  | 21.12 | 22.52 |  |  |  | 35.55 | 36.8286 |  |  |  | 115.06 | 113.6
-18.6 |  |  |  | 12.9 | 14.375 |  |  |  | 15.5789 | 17.02 |  |  |  | 21.5 | 22.92 |  |  |  | 37.35 | 38.6125 |  |  |  | 115.6 | 114.34
-18.5 |  |  |  | 13.2 | 14.6 |  |  |  | 15.904 | 17.32 |  |  |  | 21.92 | 23.2667 |  |  |  | 40.0167 | 41.4143 |  |  |  | 116.22 | 115.02
-18.4 |  |  |  | 13.25 | 14.78 |  |  |  | 16.154 | 17.55 |  |  |  | 22.3 | 23.7 |  |  |  | 44.3857 | 45.8667 |  |  |  | 116.88 | 115.375
-18.3 |  |  |  | 13.5 | 14.9 |  |  |  | 16.4641 | 17.86 |  |  |  | 22.7 | 24.0714 |  |  |  | 51.0714 | 52.0571 |  |  |  | 117.3 | 115.84
-18.2 |  |  |  | 13.7 | 15.16 |  |  |  | 16.7042 | 18.125 |  |  |  | 23.15 | 24.58 |  |  |  | 59.5625 | 60.45 |  |  |  | 117.5 | 116.2
-18.1 |  |  |  | 13.9 | 15.3 |  |  |  | 17.0443 | 18.4 |  |  |  | 23.6 | 24.95 |  |  |  | 70.4 | 70.8714 |  |  |  | 117.84 | 116.625
-18 |  |  |  | 14.125 | 15.54 |  |  |  | 17.3544 | 18.65 |  |  |  | 24.1 | 25.45 |  |  |  | 81.5 | 81.7833 |  |  |  | 118.15 | 116.8333
-17.9 |  |  |  | 14.36 | 15.775 |  |  |  | 17.6294 | 19 |  |  |  | 24.62 | 25.96 |  |  |  | 92.38 | 92.6167 |  |  | 15.6 | 118.5 | 117
-17.8 |  |  |  | 14.5 | 16.025 |  |  |  | 17.9545 | 19.35 |  |  |  | 25.125 | 26.48 |  |  |  | 102.8143 | 102.55 |  |  | 15.85 | 118.5333 | 117.3
-17.7 |  |  |  | 14.75 | 16.26 |  |  |  | 18.3046 | 19.7 |  |  |  | 25.7 | 27.1 |  |  |  | 111 | 110.4 |  |  | 16.1167 | 118.7 | 117.45
-17.6 |  |  |  | 15 | 16.48 |  |  |  | 18.6046 | 20 |  |  |  | 26.36 | 27.6714 |  |  |  | 117.1 | 116.1 |  |  | 16.4 | 118.98 | 117.6
-17.5 |  |  |  | 15.25 | 16.725 |  |  |  | 18.9047 | 20.425 |  |  |  | 27.12 | 28.42 |  |  |  | 120.9 | 119.5 |  |  | 16.5833 | 119.2 | 117.8833
-17.4 |  |  |  | 15.52 | 16.85 |  |  |  | 19.3548 | 20.8 |  |  |  | 28.05 | 29.3 |  |  |  | 122.9333 | 121.5857 |  |  | 16.8 | 119.3333 | 118
-17.3 |  |  |  | 15.76 | 17.1 |  |  |  | 19.7383 | 21.15 |  |  |  | 29.1 | 30.4143 |  |  |  | 124.4 | 122.9167 |  |  | 16.9667 | 119.5 | 118.2
-17.2 |  |  |  | 15.96 | 17.375 |  |  |  | 20.105 | 21.56 |  |  |  | 30.45 | 31.7667 |  |  |  | 125.3143 | 124.1 |  |  | 17.14 | 119.6667 | 118.3
-17.1 |  |  |  | 16.22 | 17.68 |  |  |  | 20.5851 | 22.02 |  |  |  | 32.56 | 33.6833 |  |  |  | 126.02 | 124.7 |  |  | 17.4333 | 119.8 | 118.4
-17 |  |  |  | 16.5 | 17.95 |  |  |  | 21.1553 | 22.4 |  |  |  | 35.6 | 37 |  |  |  | 126.5857 | 125.1857 |  |  | 17.7167 | 119.86 | 118.525
-16.9 |  |  |  | 16.78 | 18.2333 |  |  |  | 21.6054 | 23.06 |  |  |  | 40.84 | 41.84 |  |  | 12.9 | 127.05 | 125.5667 |  |  | 17.8833 | 119.975 | 118.65
-16.8 |  |  |  | 17.075 | 18.475 |  |  |  | 22.2306 | 23.7 |  |  |  | 49.1333 | 49.62 |  |  | 13.02 | 127.46 | 125.96 |  |  | 18.075 | 120.1 | 118.7667
-16.7 |  |  |  | 17.35 | 18.7667 |  |  |  | 23.0558 | 24.4 |  |  |  | 61 | 61.45 |  |  | 13.2 | 127.6286 | 126.25 |  |  | 18.36 | 120.2 | 119
-16.6 |  |  |  | 17.65 | 19.1 |  |  |  | 23.931 | 25.325 |  |  |  | 74.3 | 74.3 |  |  | 13.44 | 127.7714 | 126.45 |  |  | 18.65 | 120.3167 | 119.1
-16.5 |  |  |  | 18 | 19.425 |  |  |  | 25.3064 | 26.625 |  |  |  | 88.3 | 88.36 |  |  | 13.7 | 127.98 | 126.5857 |  |  | 18.84 | 120.4 | 119.1
-16.4 |  |  |  | 18.4 | 19.8 |  |  |  | 27.1068 | 28.4 |  |  |  | 103.475 | 102.4333 |  |  | 13.85 | 128.1286 | 126.7333 |  |  | 19.0333 | 120.5 | 119.2
-16.3 |  |  |  | 18.8 | 20.25 |  |  |  | 29.8575 | 30.95 |  |  |  | 116.3 | 114.28 |  |  | 14.0167 | 128.2833 | 126.9 |  |  | 19.35 | 120.54 | 119.3
-16.2 |  |  |  | 19.45 | 20.8 |  |  |  | 34.1085 | 35.38 |  |  |  | 126.0333 | 124.4667 |  |  | 14.1833 | 128.48 | 127 |  |  | 19.66 | 120.6 | 119.3
-16.1 |  |  |  | 20.125 | 21.56 |  |  |  | 42.3106 | 43.55 |  |  |  | 132.55 | 131.0857 |  |  | 14.4 | 128.6 | 127.075 |  |  | 20.05 | 120.6667 | 119.38
-16 |  |  |  | 21.1 | 22.5 |  |  |  | 54.8337 | 55.4 |  |  |  | 136.7 | 134.9333 |  |  | 14.64 | 128.7 | 127.2 |  |  | 20.4 | 120.7 | 119.4
-15.9 |  |  |  | 22.65 | 24 |  |  |  | 69.7174 | 70.3 |  |  | 9.45 | 138.7 | 136.96 |  |  | 14.8667 | 128.7 | 127.3 |  |  | 20.7 | 120.8 | 119.5
-15.8 |  |  |  | 24.825 | 26.15 |  |  |  | 85.6814 | 86.2 |  |  | 9.5 | 140.1 | 138.3 |  |  | 15.08 | 128.7857 | 127.3429 |  |  | 21.04 | 120.8333 | 119.5
-15.7 |  |  |  | 29.025 | 30.25 |  |  |  | 102.2056 | 102.9 |  |  | 9.64 | 140.7 | 139.14 |  |  | 15.3 | 128.8 | 127.4 |  |  | 21.275 | 120.9 | 119.6
-15.6 |  |  |  | 36.475 | 37.175 |  |  |  | 117.829 | 116.8 |  |  | 9.875 | 141.4 | 139.6 |  |  | 15.55 | 128.9 | 127.4667 |  |  | 21.6 | 120.9 | 119.6667
-15.5 |  |  |  | 49.275 | 50.05 |  |  |  | 131.3082 | 130.175 |  |  | 9.98 | 141.65 | 139.9 |  |  | 15.9 | 128.9 | 127.5 |  |  | 21.9 | 120.9 | 119.7
-15.4 |  |  |  | 64.9 | 65.975 |  |  | 7.8 | 140.685 | 139.06 |  |  | 10.18 | 141.9 | 140.14 |  |  | 16.1667 | 128.9 | 127.5 |  |  | 22.2714 | 120.9 | 119.7
-15.3 |  |  |  | 83.4 | 82.1 |  |  | 7.9 | 145.7865 | 144.16 |  |  | 10.3 | 142.025 | 140.3 |  |  | 16.38 | 128.9 | 127.5 |  |  | 22.5 | 120.9 | 119.75
-15.2 |  |  |  | 103 | 101.025 |  |  | 8 | 148.462 | 146.6 |  |  | 10.45 | 142.1 | 140.4 |  |  | 16.62 | 128.9 | 127.5 |  |  | 22.78 | 120.98 | 119.7667
-15.1 |  |  | 7.1486 | 120.4 | 118.54 |  |  | 8.14 | 149.837 | 148.1 |  |  | 10.6 | 142.1 | 140.4 |  |  | 16.9333 | 128.9 | 127.4 |  |  | 23.25 | 121 | 119.7
-15 |  |  | 7.2585 | 136 | 134 |  |  | 8.24 | 150.5713 | 148.75 |  |  | 10.7833 | 142.1 | 140.4 |  |  | 17.1 | 128.8 | 127.4 |  |  | 23.56 | 121 | 119.7
-14.9 |  |  | 7.2985 | 146.66 | 144.6 |  |  | 8.36 | 150.838 | 149.2 |  |  | 10.9333 | 142.1 | 140.4 |  |  | 17.375 | 128.8 | 127.4 |  |  | 23.9 | 120.9833 | 119.7
-14.8 |  |  | 7.3985 | 153 | 150.8333 |  |  | 8.45 | 151.038 | 149.225 |  |  | 11.18 | 142 | 140.4 |  |  | 17.7 | 128.75 | 127.3429 |  |  | 24.35 | 120.8 | 119.7
-14.7 |  |  | 7.4985 | 155.75 | 153.6 |  |  | 8.58 | 151.138 | 149.3 |  |  | 11.3 | 141.875 | 140.24 |  |  | 18.025 | 128.7 | 127.3 |  |  | 24.7 | 120.8 | 119.7
-14.6 |  |  | 7.5985 | 157.1 | 154.8667 |  |  | 8.7 | 151.138 | 149.3 |  |  | 11.5 | 141.7 | 140.1833 |  |  | 18.4 | 128.65 | 127.3 |  |  | 24.96 | 120.8 | 119.7
-14.5 |  |  | 7.6985 | 157.5 | 155.4 |  |  | 8.85 | 151.138 | 149.3 |  |  | 11.725 | 141.6 | 140.0167 |  |  | 18.65 | 128.5 | 127.2571 |  |  | 25.48 | 120.8 | 119.65
-14.4 |  |  | 7.7734 | 157.6 | 155.4 |  |  | 8.96 | 150.998 | 149.175 |  |  | 11.96 | 141.4667 | 139.85 |  |  | 18.9333 | 128.5 | 127.2 |  |  | 25.95 | 120.8 | 119.6
-14.3 |  |  | 7.9234 | 157.6 | 155.5 |  |  | 9.08 | 150.698 | 149.05 |  |  | 12.175 | 141.3333 | 139.6667 |  |  | 19.3 | 128.3714 | 127.2 |  |  | 26.4 | 120.8 | 119.6
-14.2 |  |  | 8.0384 | 157.375 | 155.4 |  |  | 9.22 | 150.463 | 148.74 |  |  | 12.38 | 141.2 | 139.5 |  |  | 19.5667 | 128.3 | 127.1857 |  |  | 26.9 | 120.8 | 119.6
-14.1 |  |  | 8.1984 | 157.225 | 155.125 |  |  | 9.35 | 150.2047 | 148.4 |  |  | 12.58 | 141.1 | 139.4 |  |  | 19.8667 | 128.2429 | 127.05 |  |  | 27.55 | 120.7 | 119.675
-14 |  |  | 8.2983 | 156.9 | 154.7 |  |  | 9.5 | 149.862 | 148.15 |  |  | 12.82 | 140.78 | 139.2333 |  |  | 20.3 | 128.2 | 127 |  |  | 28.1 | 120.7 | 119.625
-13.9 |  |  | 8.365 | 156.35 | 154.3 |  |  | 9.7 | 149.5703 | 147.825 |  |  | 13.05 | 140.625 | 139.0667 |  |  | 20.58 | 128.0571 | 126.9 |  |  | 28.8 | 120.7 | 119.6
-13.8 |  |  | 8.4983 | 155.92 | 153.8 |  |  | 9.825 | 149.137 | 147.575 |  |  | 13.275 | 140.4 | 138.8 |  |  | 20.95 | 128 | 126.8 |  |  | 29.7 | 120.675 | 119.6
-13.7 |  |  | 8.5983 | 155.2667 | 153.26 |  |  | 10.05 | 148.8703 | 147.25 |  |  | 13.52 | 140.24 | 138.64 |  |  | 21.4 | 127.8833 | 126.7 |  |  | 30.9 | 120.6 | 119.52
-13.6 |  |  | 8.6983 | 154.875 | 152.85 |  |  | 10.16 | 148.537 | 146.9 |  |  | 13.84 | 140 | 138.46 |  |  | 21.84 | 127.8 | 126.6167 |  |  | 32.5 | 120.46 | 119.4
-13.5 |  |  | 8.8982 | 154.45 | 152.4 |  |  | 10.38 | 148.1703 | 146.55 |  |  | 14.06 | 139.7833 | 138.1667 |  |  | 22.3 | 127.6571 | 126.4571 |  |  | 34.95 | 120.35 | 119.3
-13.4 |  |  | 8.9982 | 153.9 | 151.85 |  |  | 10.6 | 147.837 | 146.1 |  |  | 14.4 | 139.58 | 137.9667 |  |  | 22.725 | 127.5833 | 126.3143 |  |  | 38.45 | 120.2 | 119.2333
-13.3 |  |  | 9.0982 | 153.25 | 151.4 |  |  | 10.75 | 147.397 | 145.75 |  |  | 14.65 | 139.26 | 137.78 |  |  | 23.14 | 127.4375 | 126.2 |  |  | 43.28 | 120.1167 | 119.2
-13.2 |  |  | 9.2481 | 152.75 | 150.85 |  |  | 10.88 | 146.937 | 145.35 |  |  | 14.96 | 138.9667 | 137.6 |  |  | 23.65 | 127.2714 | 126.2 |  |  | 49.9 | 120.1 | 119.1833
-13.1 |  |  | 9.4181 | 152.2667 | 150.3 |  |  | 11.2 | 146.587 | 144.95 |  |  | 15.18 | 138.7 | 137.2286 |  |  | 24.3 | 127.1375 | 126.0667 |  |  | 58 | 120.1 | 119.1
-13 |  |  | 9.5731 | 151.85 | 149.875 |  |  | 11.375 | 146.237 | 144.56 |  |  | 15.62 | 138.42 | 137.0333 |  |  | 24.9667 | 127.0833 | 125.925 |  |  | 67.4857 | 120 | 119
-12.9 |  |  | 9.7647 | 151.3 | 149.25 |  |  | 11.58 | 145.836 | 144.15 |  |  | 15.875 | 138.225 | 136.7 |  |  | 25.6 | 127 | 125.7667 |  |  | 75.8667 | 120 | 118.9
-12.8 |  |  | 9.898 | 150.7 | 148.75 |  |  | 11.8 | 145.276 | 143.8 |  |  | 16.225 | 137.84 | 136.4333 |  |  | 26.65 | 126.8333 | 125.5833 |  |  | 84.7 | 119.9 | 118.82
-12.7 |  |  | 10.058 | 150.05 | 148.24 |  |  | 12.05 | 144.996 | 143.475 |  |  | 16.65 | 137.575 | 136.1714 |  |  | 27.86 | 126.68 | 125.4167 |  |  | 91.74 | 119.8 | 118.7
-12.6 |  |  | 10.198 | 149.6 | 147.8 |  |  | 12.26 | 144.436 | 142.95 |  |  | 16.95 | 137.24 | 135.92 |  |  | 29.6667 | 126.5143 | 125.2667 |  |  | 97.2 | 119.7 | 118.66
-12.5 |  |  | 10.4479 | 149.075 | 147.34 |  |  | 12.5333 | 143.936 | 142.5 |  |  | 17.32 | 137 | 135.625 |  |  | 32.1333 | 126.3667 | 125.1833 |  |  | 100.7 | 119.6 | 118.6
-12.4 |  |  | 10.5979 | 148.52 | 146.8 |  |  | 12.86 | 143.696 | 142.22 |  |  | 17.7667 | 136.725 | 135.4167 |  |  | 36.05 | 126.1833 | 125.0167 |  |  | 103.1 | 119.3833 | 118.5
-12.3 |  |  | 10.7978 | 148 | 146.225 |  |  | 13.1 | 143.261 | 141.7 |  |  | 18.1 | 136.5 | 135.1 |  |  | 42 | 126.0167 | 124.8333 |  |  | 104.4 | 119.3 | 118.38
-12.2 |  |  | 10.9978 | 147.4333 | 145.825 |  |  | 13.34 | 142.786 | 141.2667 |  |  | 18.7 | 136.175 | 134.86 |  |  | 49.8167 | 125.8667 | 124.7 |  |  | 105.375 | 119.2 | 118.2833
-12.1 |  |  | 11.1645 | 147.05 | 145.225 |  |  | 13.7 | 142.386 | 141 |  |  | 19.3 | 135.925 | 134.5286 |  |  | 59.2 | 125.6286 | 124.5333 |  |  | 106.25 | 119.1 | 118.175
-12 |  |  | 11.3977 | 146.26 | 144.65 |  |  | 14.075 | 141.8853 | 140.58 |  |  | 19.9 | 135.575 | 134.2143 |  |  | 70.45 | 125.4333 | 124.3833 |  |  | 106.78 | 118.9667 | 118.1
-11.9 |  |  | 11.6227 | 145.75 | 144.15 |  |  | 14.44 | 141.535 | 140.18 |  |  | 20.7 | 135.3 | 133.9286 |  |  | 80.3 | 125.3 | 124.2167 |  |  | 107.18 | 118.9 | 118
-11.8 |  |  | 11.7976 | 145.275 | 143.6 |  |  | 14.84 | 141.035 | 139.8 |  |  | 21.6 | 135.05 | 133.68 |  |  | 89.625 | 125.1857 | 124.02 |  |  | 107.58 | 118.88 | 117.98
-11.7 |  |  | 12.0976 | 144.75 | 143.25 |  |  | 15.18 | 140.685 | 139.2667 |  |  | 22.775 | 134.5 | 133.2667 |  |  | 97.75 | 124.94 | 123.94 |  |  | 107.95 | 118.8 | 117.9
-11.6 |  |  | 12.3642 | 144.18 | 142.65 |  |  | 15.7 | 140.255 | 138.8 |  |  | 24.66 | 134.375 | 133.0286 |  |  | 103.15 | 124.8 | 123.7833 |  |  | 108.2 | 118.7 | 117.8
-11.5 |  |  | 12.5975 | 143.8 | 142.3 |  |  | 16.15 | 139.775 | 138.375 |  |  | 27.25 | 134.0667 | 132.6857 |  |  | 107 | 124.7333 | 123.575 |  |  | 108.34 | 118.6 | 117.6833
-11.4 |  |  | 12.8474 | 143.25 | 141.56 |  |  | 17 | 139.485 | 138.04 |  |  | 31.6833 | 133.7 | 132.45 |  |  | 108.9 | 124.46 | 123.3 |  | 8.56 | 108.65 | 118.44 | 117.5
-11.3 |  |  | 13.1374 | 142.675 | 141.2 |  |  | 18 | 138.935 | 137.5667 |  |  | 38.35 | 133.275 | 132.0571 |  |  | 110.3 | 124.3 | 123.1167 |  | 8.6667 | 108.875 | 118.3 | 117.4
-11.2 |  |  | 13.5473 | 142.1 | 140.85 |  |  | 19.25 | 138.575 | 137.2 |  |  | 48.5 | 133.025 | 131.8333 |  |  | 111.22 | 124.1167 | 123 |  | 8.95 | 109.08 | 118.16 | 117.35
-11.1 |  |  | 14.0972 | 141.7 | 140.32 |  |  | 21.1 | 138.035 | 136.8 |  |  | 59.5 | 132.6 | 131.3833 |  |  | 111.84 | 123.8333 | 122.8857 |  | 9.1 | 109.3 | 118.06 | 117.3
-11 |  |  | 14.7171 | 141.175 | 139.8 |  |  | 24.06 | 137.584 | 136.4 |  |  | 72.7 | 132.3 | 131.1333 |  |  | 112.36 | 123.6429 | 122.7 |  | 9.3 | 109.42 | 117.9667 | 117.16
-10.9 |  |  | 15.6568 | 140.85 | 139.2 |  |  | 28.9 | 137.214 | 136.025 |  |  | 84.86 | 131.98 | 130.8 |  | 6.675 | 112.75 | 123.4286 | 122.4833 |  | 9.4 | 109.6 | 117.8 | 116.95
-10.8 |  |  | 16.8966 | 140.2 | 138.7 |  |  | 36.8 | 136.834 | 135.525 |  |  | 96.8 | 131.6 | 130.3714 |  | 6.8 | 112.925 | 123.28 | 122.3 |  | 9.5 | 109.7 | 117.72 | 116.8333
-10.7 |  |  | 18.9463 | 139.7 | 138.26 |  |  | 48.725 | 136.384 | 135.1 |  |  | 105.76 | 131.32 | 130.1333 |  | 6.9 | 113.25 | 123.0833 | 122.1167 |  | 9.65 | 109.7833 | 117.52 | 116.7
-10.6 |  |  | 22.7621 | 139.3 | 137.75 |  |  | 62.3 | 135.884 | 134.75 |  |  | 112.1 | 131 | 129.7857 |  | 7 | 113.4333 | 122.8571 | 121.8857 |  | 9.96 | 109.85 | 117.4 | 116.5
-10.5 |  |  | 29.2341 | 138.74 | 137.22 |  |  | 76 | 135.334 | 134.3 |  |  | 115.8 | 130.6 | 129.5857 |  | 7.1 | 113.6 | 122.6429 | 121.72 |  | 10.1 | 110 | 117.2333 | 116.5
-10.4 |  |  | 39.767 | 138.2 | 136.72 |  |  | 89.925 | 134.984 | 133.82 |  | 4.5 | 117.925 | 130.22 | 129.24 |  | 7.2 | 113.7833 | 122.4286 | 121.5143 |  | 10.275 | 110 | 117.1 | 116.4
-10.3 |  |  | 54.0642 | 137.65 | 136.3 |  |  | 102.5 | 134.514 | 133.45 |  | 4.5 | 119.1 | 129.92 | 128.88 |  | 7.3 | 113.9 | 122.3 | 121.3143 |  | 10.4 | 110.1 | 116.9333 | 116.2833
-10.2 |  |  | 68.7462 | 137.3 | 135.85 |  | 3.4 | 112.85 | 134.084 | 133 |  | 4.6 | 119.9 | 129.6 | 128.5 |  | 7.5 | 114.04 | 122.0857 | 121.1667 |  | 10.6 | 110.2 | 116.7667 | 116.1167
-10.1 |  |  | 84.8497 | 136.7 | 135.25 |  | 3.4 | 119.24 | 133.613 | 132.55 |  | 4.7 | 120.42 | 129.34 | 128.2167 |  | 7.6 | 114.14 | 121.8667 | 120.9667 |  | 10.8 | 110.3 | 116.6 | 116.0333
-10 |  | 3.1 | 99.2053 | 136.1 | 134.8 |  | 3.5 | 122.8 | 133.233 | 132.2 |  | 4.8 | 120.775 | 129 | 127.92 |  | 7.7667 | 114.2 | 121.6375 | 120.7167 |  | 11.1 | 110.3333 | 116.5 | 115.86
-9.9 |  | 3.2 | 112.1277 | 135.65 | 134.38 |  | 3.6 | 124.54 | 132.733 | 131.75 |  | 4.9 | 121 | 128.75 | 127.6 |  | 7.8667 | 114.3 | 121.38 | 120.5429 |  | 11.3 | 110.38 | 116.3833 | 115.625
-9.8 |  | 3.225 | 121.136 | 135.28 | 133.9667 |  | 3.65 | 125.7333 | 132.393 | 131.45 |  | 5 | 121.1 | 128.275 | 127.2375 |  | 8.1 | 114.3 | 121.2 | 120.3 |  | 11.4 | 110.4 | 116.2167 | 115.5
-9.7 |  | 3.3 | 126.15 | 134.8 | 133.5 |  | 3.7 | 126.325 | 131.853 | 130.96 |  | 5.1 | 121.3 | 127.88 | 126.92 |  | 8.2333 | 114.3 | 120.94 | 120.14 |  | 11.7 | 110.4 | 116.08 | 115.45
-9.6 |  | 3.3 | 128.694 | 134.32 | 133.0333 |  | 3.8 | 126.55 | 131.433 | 130.65 |  | 5.2 | 121.4 | 127.55 | 126.62 |  | 8.4 | 114.3 | 120.72 | 119.86 |  | 11.9 | 110.4 | 116 | 115.32
-9.5 |  | 3.4 | 129.674 | 133.85 | 132.475 |  | 3.8 | 126.7 | 131.033 | 130.1 |  | 5.3 | 121.4 | 127.26 | 126.2667 |  | 8.5667 | 114.4 | 120.5714 | 119.6833 |  | 12 | 110.4 | 115.8333 | 115.2
-9.4 |  | 3.4 | 130.349 | 133.4 | 132.1 |  | 3.9 | 126.8 | 130.683 | 129.825 |  | 5.4 | 121.4 | 126.96 | 126 |  | 8.6333 | 114.4 | 120.3333 | 119.5 |  | 12.2 | 110.5 | 115.5667 | 115
-9.3 |  | 3.5 | 130.499 | 132.85 | 131.62 |  | 3.9 | 126.9 | 130.283 | 129.35 |  | 5.5 | 121.4 | 126.625 | 125.6 |  | 8.9 | 114.4 | 120.075 | 119.3667 |  | 12.4667 | 110.5 | 115.5 | 114.8
-9.2 |  | 3.5 | 130.574 | 132.3 | 131.1 |  | 4 | 126.8 | 129.882 | 129.02 |  | 5.675 | 121.3 | 126.25 | 125.2667 |  | 9.1333 | 114.4 | 119.84 | 119.125 |  | 12.65 | 110.5 | 115.34 | 114.7
-9.1 |  | 3.6 | 130.474 | 131.86 | 130.75 |  | 4.1 | 126.7 | 129.432 | 128.65 |  | 5.7667 | 121.2 | 125.85 | 125 |  | 9.3 | 114.4 | 119.5429 | 118.9167 |  | 13 | 110.5 | 115.225 | 114.52
-9 |  | 3.6 | 130.324 | 131.4 | 130.38 |  | 4.2 | 126.6 | 129.132 | 128.1 |  | 5.875 | 121.14 | 125.475 | 124.7 |  | 9.5 | 114.3667 | 119.3167 | 118.72 |  | 13.1667 | 110.5 | 115.0833 | 114.4167
-8.9 |  | 3.7 | 130.074 | 130.92 | 129.85 |  | 4.2333 | 126.425 | 128.657 | 127.8 |  | 5.9 | 121 | 125.25 | 124.3 |  | 9.6667 | 114.3 | 119.0667 | 118.35 |  | 13.4 | 110.5 | 114.8714 | 114.2
-8.8 |  | 3.8 | 129.854 | 130.375 | 129.35 |  | 4.3 | 126.2 | 128.082 | 127.375 |  | 6.1 | 120.9 | 124.9 | 123.94 |  | 10 | 114.3 | 118.8667 | 118.1714 |  | 13.75 | 110.5 | 114.7 | 114.2
-8.7 |  | 3.8 | 129.474 | 129.975 | 128.9 |  | 4.4 | 125.975 | 127.772 | 126.95 |  | 6.2 | 120.8 | 124.5667 | 123.6 |  | 10.1667 | 114.2 | 118.6286 | 117.8857 |  | 14 | 110.4 | 114.5667 | 113.9
-8.6 |  | 3.8667 | 129.199 | 129.5 | 128.5 |  | 4.525 | 125.725 | 127.432 | 126.5 |  | 6.4 | 120.7 | 124.1667 | 123.3 |  | 10.325 | 114.15 | 118.44 | 117.7 |  | 14.35 | 110.4 | 114.3 | 113.8333
-8.5 |  | 3.9 | 128.849 | 129.14 | 128 |  | 4.6 | 125.5 | 126.982 | 126.14 |  | 6.5 | 120.5 | 123.9 | 123.08 |  | 10.55 | 114 | 118.1667 | 117.54 |  | 14.6 | 110.3 | 114.175 | 113.6667
-8.4 |  | 4 | 128.524 | 128.64 | 127.55 |  | 4.7 | 125.2667 | 126.632 | 125.76 |  | 6.6333 | 120.35 | 123.55 | 122.7 |  | 10.8333 | 113.925 | 118.0143 | 117.3429 |  | 14.9 | 110.2 | 113.9833 | 113.4667
-8.3 |  | 4.1 | 128.174 | 128.3 | 127.1 |  | 4.8333 | 125 | 126.232 | 125.45 |  | 6.8667 | 120.22 | 123.22 | 122.4 |  | 11.0667 | 113.8 | 117.8333 | 117.0571 |  | 15.175 | 110.2 | 113.9 | 113.26
-8.2 |  | 4.1667 | 127.824 | 127.75 | 126.725 |  | 4.9 | 124.575 | 125.791 | 125.05 |  | 7 | 120.1 | 122.9333 | 122.1571 |  | 11.4 | 113.725 | 117.5667 | 116.8667 |  | 15.5333 | 110.1 | 113.75 | 113.1
-8.1 |  | 4.2333 | 127.455 | 127.3 | 126.3 |  | 5 | 124.3 | 125.381 | 124.6 |  | 7.1667 | 119.775 | 122.6 | 121.7714 |  | 11.625 | 113.5833 | 117.35 | 116.6833 |  | 15.95 | 110.1 | 113.7 | 113.0167
-8 |  | 4.3 | 127.055 | 126.8 | 125.7 |  | 5.1 | 124.05 | 124.881 | 124.1 |  | 7.4 | 119.55 | 122.175 | 121.375 |  | 11.95 | 113.5 | 117.1833 | 116.4333 |  | 16.3 | 110.1 | 113.48 | 112.85
-7.9 |  | 4.4 | 126.575 | 126.25 | 125.45 |  | 5.25 | 123.82 | 124.481 | 123.78 |  | 7.55 | 119.3 | 121.8333 | 120.98 |  | 12.3 | 113.375 | 117 | 116.14 |  | 16.6 | 110.1 | 113.3 | 112.6833
-7.8 |  | 4.5 | 126.255 | 125.9333 | 125 |  | 5.4 | 123.55 | 124.081 | 123.38 |  | 7.8 | 119.05 | 121.42 | 120.6 |  | 12.6 | 113.3 | 116.6857 | 115.9 |  | 16.95 | 110 | 113.1167 | 112.6
-7.7 |  | 4.6 | 125.755 | 125.5 | 124.6 |  | 5.5667 | 123.2 | 123.731 | 123.04 |  | 8 | 118.86 | 121.15 | 120.36 |  | 12.9667 | 113.2 | 116.3667 | 115.6333 |  | 17.5 | 109.9 | 112.9 | 112.44
-7.6 |  | 4.675 | 125.375 | 125.06 | 124.06 |  | 5.7 | 122.8 | 123.211 | 122.6 |  | 8.1667 | 118.64 | 120.82 | 119.9333 |  | 13.1 | 113.0667 | 116.0429 | 115.45 |  | 17.9667 | 109.9 | 112.7 | 112.2
-7.5 |  | 4.8 | 124.995 | 124.58 | 123.65 |  | 5.8333 | 122.5 | 122.8977 | 122.2 |  | 8.5 | 118.44 | 120.36 | 119.7 |  | 13.5 | 112.9 | 115.8286 | 115.2714 | 3.2 | 18.5 | 109.8 | 112.6 | 112.0667
-7.4 |  | 4.9 | 124.525 | 124.2 | 123.3 |  | 6 | 122.125 | 122.481 | 121.85 |  | 8.7 | 118.12 | 120.05 | 119.44 | 2.4 | 13.8667 | 112.8 | 115.6 | 115.0571 | 3.2 | 19.1 | 109.8 | 112.4167 | 111.8833
-7.3 |  | 5 | 123.995 | 123.76 | 122.9 |  | 6.2 | 121.82 | 122.131 | 121.52 |  | 8.9333 | 117.8 | 119.7 | 119.1333 | 2.3 | 14.275 | 112.7 | 115.2857 | 114.825 | 3.2 | 20.0667 | 109.8 | 112.3 | 111.7
-7.2 |  | 5.1 | 123.6417 | 123.2667 | 122.4667 |  | 6.3 | 121.5 | 121.68 | 121.15 |  | 9.3333 | 117.675 | 119.4 | 118.8 | 2.3 | 14.7667 | 112.52 | 115.075 | 114.5333 | 3.2667 | 21.075 | 109.7 | 112.0667 | 111.6
-7.1 |  | 5.2667 | 123.25 | 122.85 | 122.125 | 1.1 | 6.5 | 121.06 | 121.27 | 120.58 | 1.4 | 9.6 | 117.35 | 119 | 118.3667 | 2.35 | 15.2 | 112.4 | 114.8143 | 114.35 | 3.3 | 22.8 | 109.6 | 111.8 | 111.5
-7 | 1 | 5.325 | 122.875 | 122.4 | 121.55 | 1.1 | 6.8 | 120.65 | 120.88 | 120.22 | 1.5 | 9.8667 | 117.1 | 118.8 | 118.1 | 2.4 | 15.8333 | 112.2 | 114.6 | 114.2 | 3.3 | 24.9 | 109.5 | 111.7 | 111.4
-6.9 | 1 | 5.5 | 122.426 | 121.95 | 121.1 | 1.1 | 7 | 120.34 | 120.45 | 119.8667 | 1.5 | 10.2667 | 116.9 | 118.45 | 117.66 | 2.5 | 16.6 | 112.05 | 114.3667 | 113.92 | 3.2 | 28.2 | 109.5 | 111.6 | 111.3
-6.8 | 1 | 5.6333 | 121.926 | 121.7 | 120.7333 | 1.1 | 7.2 | 119.96 | 120.05 | 119.45 | 1.5 | 10.5667 | 116.55 | 118.125 | 117.34 | 2.6 | 17.45 | 111.9333 | 114.175 | 113.65 | 3.35 | 33.1333 | 109.3167 | 111.45 | 111.1333
-6.7 | 1 | 5.8 | 121.676 | 121.1 | 120.35 | 1.2 | 7.5 | 119.64 | 119.755 | 119.05 | 1.6 | 10.9333 | 116.25 | 117.8 | 117.15 | 2.6 | 18.4333 | 111.8 | 114.0429 | 113.3667 | 3.4 | 40.5 | 109.2 | 111.4 | 110.9
-6.6 | 1 | 6 | 121.126 | 120.7667 | 119.925 | 1.2 | 7.7667 | 119.26 | 119.305 | 118.8 | 1.6 | 11.4 | 115.95 | 117.3 | 116.86 | 2.7 | 20.1 | 111.7 | 113.8667 | 113.2 | 3.55 | 48.875 | 109.0833 | 111.3 | 110.7167
-6.5 | 1 | 6.1 | 120.696 | 120.34 | 119.45 | 1.2 | 8 | 118.9 | 118.805 | 118.26 | 1.7 | 12 | 115.875 | 117.02 | 116.5 | 2.7 | 22.1333 | 111.6 | 113.575 | 113.0333 | 3.65 | 58.2 | 109 | 111.2 | 110.62
-6.4 | 1 | 6.4 | 120.2093 | 119.88 | 119.125 | 1.2 | 8.4 | 118.5 | 118.48 | 117.96 | 1.7 | 12.6 | 115.6333 | 116.7167 | 116.1667 | 2.8 | 25.8 | 111.4 | 113.4125 | 112.8571 | 3.8 | 67.95 | 108.92 | 110.9667 | 110.5333
-6.3 | 1 | 6.6 | 119.796 | 119.275 | 118.7 | 1.2 | 8.8 | 118.2 | 118.105 | 117.52 | 1.7 | 13.4 | 115.3 | 116.5 | 115.82 | 2.9 | 31 | 111.3 | 113.2 | 112.6833 | 4 | 76.7 | 108.9 | 110.7667 | 110.5
-6.2 | 1 | 6.85 | 119.376 | 118.85 | 118.2333 | 1.3 | 9.4333 | 117.8 | 117.754 | 117.125 | 1.7 | 14.5 | 114.95 | 116.08 | 115.5667 | 2.9 | 38.6 | 111.2 | 113 | 112.5167 | 4.05 | 84.5333 | 108.84 | 110.6 | 110.25
-6.1 | 1 | 7.1333 | 118.951 | 118.4667 | 117.825 | 1.3 | 10 | 117.4 | 117.279 | 116.8 | 1.8 | 16.1667 | 114.8 | 115.72 | 115.2571 | 3 | 48.2 | 111.1 | 112.7143 | 112.3167 | 4.1 | 90.6 | 108.7667 | 110.4333 | 110.1
-6 | 1 | 7.6333 | 118.676 | 118.15 | 117.6 | 1.3 | 10.9 | 117.15 | 117.129 | 116.58 | 1.9 | 18.4 | 114.56 | 115.425 | 114.8667 | 3 | 57.85 | 110.96 | 112.38 | 112.06 | 4.2 | 94.35 | 108.7 | 110.3 | 110
-5.9 | 1.05 | 8.2 | 118.176 | 117.75 | 117.05 | 1.3 | 12.15 | 116.78 | 116.729 | 116.25 | 1.9 | 22.5 | 114.35 | 115.22 | 114.6571 | 3.2 | 69.1 | 110.76 | 112.2 | 111.8833 | 4.2 | 97 | 108.7 | 110.1833 | 109.8
-5.8 | 1.1 | 9 | 117.896 | 117.32 | 116.7 | 1.35 | 13.9 | 116.4 | 116.329 | 115.925 | 1.95 | 28.65 | 114.1 | 114.9667 | 114.4 | 3.3 | 78 | 110.6667 | 112.1 | 111.7167 | 4.3 | 98.4 | 108.6167 | 110 | 109.6333
-5.7 | 1.1 | 10.3 | 117.477 | 116.95 | 116.3 | 1.4 | 17.1 | 116.175 | 116.069 | 115.575 | 2 | 37.1667 | 113.85 | 114.675 | 114.2167 | 3.35 | 86.5667 | 110.6 | 111.9429 | 111.36 | 4.45 | 99.7 | 108.35 | 109.8286 | 109.54
-5.6 | 1.1 | 12.125 | 117.127 | 116.6333 | 115.92 | 1.4 | 22.15 | 115.96 | 115.704 | 115.26 | 2.1 | 47.55 | 113.6 | 114.38 | 113.95 | 3.4 | 92.6 | 110.42 | 111.775 | 111.18 | 4.6 | 100.375 | 108.2833 | 109.56 | 109.3667
-5.5 | 1.2 | 15.6 | 116.7437 | 116.32 | 115.5833 | 1.4 | 30.25 | 115.55 | 115.379 | 115.025 | 2.2 | 59.6333 | 113.475 | 114.2 | 113.72 | 3.5 | 96.65 | 110.3 | 111.525 | 111.0143 | 4.7 | 100.925 | 108.1167 | 109.325 | 109.14
-5.4 | 1.2 | 21.7 | 116.327 | 116 | 115.4 | 1.5 | 41.2 | 115.3 | 115.0957 | 114.675 | 2.2 | 71.4 | 113.26 | 113.9 | 113.42 | 3.6 | 99.0667 | 110.2 | 111.3857 | 110.85 | 4.8 | 101.4 | 108.1 | 109.2 | 109.04
-5.3 | 1.2 | 31.6 | 116.077 | 115.5667 | 115.06 | 1.5 | 53.5667 | 115.05 | 114.879 | 114.5 | 2.3 | 82.1667 | 113 | 113.58 | 113.06 | 3.7 | 100.7 | 110.0333 | 111.1667 | 110.7 | 4.9 | 101.7333 | 108 | 109.1 | 108.9
-5.2 | 1.3 | 43.9 | 115.877 | 115.175 | 114.68 | 1.55 | 66.2667 | 114.7 | 114.554 | 114.175 | 2.4 | 90.8667 | 112.78 | 113.36 | 112.9 | 3.8 | 101.5333 | 109.8333 | 110.96 | 110.5167 | 5.1 | 102 | 107.9 | 108.9667 | 108.7
-5.1 | 1.3 | 57.8 | 115.502 | 114.925 | 114.475 | 1.6 | 77.8333 | 114.46 | 114.269 | 113.84 | 2.5 | 97.1 | 112.6 | 113.16 | 112.6333 | 3.9 | 102.05 | 109.62 | 110.6286 | 110.275 | 5.3 | 102.25 | 107.86 | 108.7 | 108.5
-5 | 1.3 | 70.6 | 115.177 | 114.6667 | 114.2 | 1.7 | 88.35 | 114.1 | 114.029 | 113.4 | 2.5 | 100.8333 | 112.325 | 112.95 | 112.45 | 3.95 | 102.6333 | 109.4333 | 110.3333 | 110.0143 | 5.4 | 102.3 | 107.78 | 108.52 | 108.3833
-4.9 | 1.3 | 83.65 | 114.857 | 114.325 | 113.94 | 1.7 | 96.8 | 113.8167 | 113.678 | 113.275 | 2.6 | 102.8 | 112 | 112.64 | 112.1667 | 4.1 | 102.8667 | 109.2333 | 110.1667 | 109.7714 | 5.6 | 102.6 | 107.5667 | 108.38 | 108.275
-4.8 | 1.4 | 94.4 | 114.602 | 114.075 | 113.6 | 1.8 | 102.15 | 113.6 | 113.348 | 112.95 | 2.7 | 104.3667 | 111.7667 | 112.28 | 111.8 | 4.3 | 103.2 | 109.1 | 110 | 109.5429 | 5.7 | 102.7333 | 107.34 | 108.3 | 108.1667
-4.7 | 1.4 | 102 | 114.377 | 113.78 | 113.2667 | 1.9 | 104.9 | 113.35 | 112.9947 | 112.7 | 2.7 | 105.1 | 111.56 | 112 | 111.55 | 4.4 | 103.425 | 108.9 | 109.7667 | 109.3714 | 5.8 | 102.8 | 107.3 | 108.125 | 107.9667
-4.6 | 1.5 | 106.4 | 113.927 | 113.44 | 112.975 | 1.9 | 106.6 | 113.04 | 112.678 | 112.38 | 2.9 | 105.7 | 111.375 | 111.76 | 111.2857 | 4.55 | 103.65 | 108.78 | 109.5714 | 109.2167 | 6 | 102.9 | 107.16 | 107.95 | 107.68
-4.5 | 1.5 | 108.5 | 113.577 | 113.16 | 112.6 | 2 | 107.5 | 112.675 | 112.328 | 112.15 | 2.9 | 106 | 111.04 | 111.44 | 111.06 | 4.8 | 103.8 | 108.56 | 109.4 | 109.05 | 6.1 | 103 | 107.025 | 107.75 | 107.5
-4.4 | 1.5 | 109.6 | 113.252 | 112.825 | 112.32 | 2.05 | 107.9667 | 112.42 | 112.103 | 111.82 | 3.1 | 106.2 | 110.8333 | 111.1 | 110.76 | 4.9 | 103.8 | 108.475 | 109.2 | 108.9 | 6.3 | 103.1 | 106.8167 | 107.6 | 107.3667
-4.3 | 1.6 | 110.2667 | 112.902 | 112.45 | 112.04 | 2.1 | 108.2 | 112.1 | 111.853 | 111.55 | 3.2 | 106.3 | 110.5 | 110.86 | 110.6 | 5 | 104 | 108.325 | 108.8571 | 108.62 | 6.4 | 103.2 | 106.7 | 107.5 | 107.2167
-4.2 | 1.6 | 110.4333 | 112.478 | 112.125 | 111.625 | 2.2 | 108.5 | 111.7 | 111.508 | 111.225 | 3.3 | 106.5 | 110.325 | 110.575 | 110.3 | 5.2 | 104 | 108.1 | 108.6667 | 108.4429 | 6.6 | 103.275 | 106.6 | 107.25 | 107.1
-4.1 | 1.7 | 110.5 | 112.278 | 111.875 | 111.275 | 2.3 | 108.6 | 111.5 | 111.078 | 110.85 | 3.5 | 106.5 | 110.0333 | 110.325 | 109.98 | 5.4 | 104.1 | 107.95 | 108.4833 | 108.2286 | 7 | 103.4 | 106.5 | 107.075 | 106.925
-4 | 1.7 | 110.5 | 111.828 | 111.55 | 111.075 | 2.4 | 108.6 | 111.2 | 110.868 | 110.5 | 3.6 | 106.5 | 109.75 | 110.08 | 109.6714 | 5.7 | 104.1 | 107.8 | 108.3167 | 107.9667 | 7.2 | 103.4 | 106.3167 | 106.85 | 106.7
-3.9 | 1.8 | 110.5 | 111.478 | 111.15 | 110.74 | 2.5 | 108.575 | 110.9 | 110.528 | 110.26 | 3.7 | 106.5 | 109.48 | 109.86 | 109.5 | 5.9 | 104.2 | 107.625 | 108.0625 | 107.8 | 7.4 | 103.4 | 106.2 | 106.7833 | 106.6
-3.8 | 1.8333 | 110.4 | 111.228 | 110.8333 | 110.4 | 2.6 | 108.5 | 110.5333 | 110.228 | 110 | 4 | 106.5 | 109.125 | 109.5667 | 109.22 | 6.1 | 104.2 | 107.4 | 107.9 | 107.65 | 7.75 | 103.425 | 106.16 | 106.7 | 106.5
-3.7 | 1.9 | 110.125 | 110.803 | 110.4 | 110.0667 | 2.7 | 108.4 | 110.18 | 109.887 | 109.775 | 4.1 | 106.45 | 108.9 | 109.3429 | 109.02 | 6.4 | 104.3 | 107.2 | 107.6857 | 107.5 | 8.05 | 103.5 | 106.0667 | 106.54 | 106.3429
-3.6 | 2 | 109.9 | 110.378 | 110.15 | 109.7 | 2.9 | 108.3 | 109.95 | 109.587 | 109.54 | 4.3 | 106.5 | 108.7 | 109 | 108.74 | 6.6 | 104.3 | 107.05 | 107.5667 | 107.3286 | 8.4 | 103.5 | 105.9667 | 106.4 | 106.28
-3.5 | 2.1 | 109.7 | 110.153 | 109.76 | 109.36 | 3.1 | 108.1333 | 109.6 | 109.387 | 109.25 | 4.6 | 106.4 | 108.525 | 108.76 | 108.5 | 7 | 104.3 | 106.8 | 107.4 | 107.1 | 9 | 103.5 | 105.7667 | 106.3 | 106.06
-3.4 | 2.2 | 109.5 | 109.858 | 109.54 | 109.18 | 3.2 | 108.0333 | 109.3333 | 108.977 | 108.825 | 4.85 | 106.4 | 108.3 | 108.48 | 108.1833 | 7.4 | 104.3 | 106.7 | 107.2167 | 106.9143 | 9.9333 | 103.5 | 105.6 | 106.1 | 105.96
-3.3 | 2.3 | 109.35 | 109.603 | 109.3 | 108.925 | 3.5 | 107.9 | 109.1 | 108.827 | 108.62 | 5.2 | 106.3 | 108.1 | 108.2 | 108.0167 | 7.8 | 104.3 | 106.54 | 106.9857 | 106.78 | 10.7 | 103.5 | 105.5 | 105.8333 | 105.75
-3.2 | 2.4 | 109.1 | 109.278 | 108.96 | 108.68 | 3.7 | 107.8 | 108.78 | 108.477 | 108.34 | 5.5 | 106.4 | 107.8333 | 108.05 | 107.7333 | 8.6 | 104.4 | 106.42 | 106.8 | 106.6 | 11.85 | 103.5 | 105.3 | 105.6333 | 105.6
-3.1 | 2.6 | 108.9 | 108.953 | 108.825 | 108.475 | 4.1 | 107.6333 | 108.5333 | 108.327 | 108.075 | 6 | 106.3 | 107.5 | 107.7 | 107.56 | 9.45 | 104.3 | 106.3333 | 106.575 | 106.425 | 14.1 | 103.5 | 105.2 | 105.5 | 105.44
-3 | 2.85 | 108.7 | 108.603 | 108.45 | 108.1 | 4.5 | 107.5 | 108.4 | 108.027 | 107.825 | 6.65 | 106.2 | 107.3 | 107.52 | 107.2667 | 10.7 | 104.3 | 106.14 | 106.32 | 106.1429 | 17.35 | 103.45 | 105.0667 | 105.4 | 105.34
-2.9 | 3.2 | 108.45 | 108.398 | 108.2667 | 107.85 | 5 | 107.3 | 108.075 | 107.827 | 107.62 | 7.4 | 106.1 | 107.05 | 107.3 | 107.1143 | 12.8 | 104.2 | 106 | 106.1 | 106 | 22.1 | 103.4 | 104.96 | 105.3 | 105.12
-2.8 | 3.75 | 108.2 | 108.078 | 107.94 | 107.6 | 5.8 | 107.1333 | 107.8 | 107.467 | 107.45 | 8.8 | 106 | 106.925 | 107.04 | 106.95 | 16.2 | 104.1 | 105.82 | 105.9 | 105.8 | 28.7 | 103.4 | 104.8 | 105.1333 | 105
-2.7 | 4.45 | 108 | 107.828 | 107.675 | 107.36 | 7.2 | 107 | 107.6 | 107.307 | 107.2 | 10.8 | 105.9 | 106.64 | 106.7833 | 106.58 | 21 | 104.1 | 105.7 | 105.7333 | 105.575 | 37.85 | 103.4 | 104.7 | 104.9667 | 104.84
-2.6 | 5.65 | 107.7333 | 107.578 | 107.425 | 107.05 | 9 | 106.7 | 107.325 | 107.087 | 106.875 | 13.7 | 105.8 | 106.46 | 106.575 | 106.3286 | 28.45 | 104 | 105.475 | 105.54 | 105.325 | 47.6 | 103.4 | 104.56 | 104.8 | 104.6333
-2.5 | 7.6 | 107.5667 | 107.329 | 107.1667 | 106.8 | 12 | 106.6 | 107.05 | 106.787 | 106.58 | 18.5 | 105.7 | 106.18 | 106.34 | 106.05 | 38.1 | 103.9 | 105.26 | 105.3143 | 105.1286 | 59 | 103.4 | 104.3 | 104.62 | 104.48
-2.4 | 11.1 | 107.4 | 107.079 | 106.9 | 106.575 | 17.55 | 106.4333 | 106.875 | 106.527 | 106.325 | 27.2 | 105.5 | 105.95 | 106.04 | 105.8714 | 49.4 | 103.8 | 105.06 | 105.1 | 104.9667 | 69.85 | 103.4 | 104.18 | 104.35 | 104.3
-2.3 | 17.7 | 107.1667 | 106.779 | 106.55 | 106.2 | 25.75 | 106.3 | 106.525 | 106.307 | 106.0667 | 37.6 | 105.3 | 105.72 | 105.82 | 105.7 | 60.2 | 103.75 | 104.8 | 104.9429 | 104.8 | 78.6 | 103.4 | 104.1 | 104.18 | 104.1333
-2.2 | 28.1 | 106.9 | 106.529 | 106.4 | 105.98 | 36.95 | 106.0333 | 106.3 | 105.9664 | 105.875 | 48.3 | 105.1 | 105.55 | 105.675 | 105.4333 | 71.4 | 103.6 | 104.7 | 104.725 | 104.6 | 87.2 | 103.3 | 103.9333 | 104 | 104.0333
-2.1 | 38.95 | 106.5667 | 106.179 | 106 | 105.78 | 49.3 | 105.8333 | 106 | 105.746 | 105.525 | 60.9 | 105 | 105.2 | 105.35 | 105.25 | 81.05 | 103.5 | 104.5 | 104.4667 | 104.4 | 92.65 | 103.1333 | 103.75 | 103.825 | 103.8
-2 | 53.3 | 106.3 | 105.929 | 105.775 | 105.6 | 61.2 | 105.5667 | 105.7 | 105.486 | 105.3167 | 72.9 | 104.8333 | 104.98 | 105.1 | 104.9667 | 89 | 103.3667 | 104.34 | 104.2 | 104.1429 | 96.1 | 103 | 103.5833 | 103.7 | 103.625
-1.9 | 65.9 | 105.9 | 105.619 | 105.44 | 105.25 | 73.1 | 105.3667 | 105.36 | 105.166 | 105.1 | 82.1 | 104.6 | 104.75 | 104.875 | 104.7167 | 93.8 | 103.3 | 104 | 104.0333 | 103.9333 | 97.8 | 102.9 | 103.5 | 103.6 | 103.5
-1.8 | 77.8 | 105.4333 | 105.329 | 105.225 | 105 | 84.1 | 105.1 | 105.06 | 104.906 | 104.925 | 90.4 | 104.35 | 104.5 | 104.64 | 104.5 | 97.2 | 103.1 | 103.775 | 103.8 | 103.7 | 99.1 | 102.78 | 103.3143 | 103.4333 | 103.32
-1.7 | 88.1 | 105.2667 | 105.0123 | 104.9 | 104.84 | 92.1 | 104.8 | 104.92 | 104.676 | 104.5 | 95.2 | 104.2 | 104.2833 | 104.4 | 104.26 | 98.9 | 103 | 103.6 | 103.65 | 103.5167 | 99.7 | 102.6 | 103.25 | 103.2 | 103.1333
-1.6 | 95.4 | 104.9667 | 104.679 | 104.6333 | 104.45 | 96.65 | 104.5 | 104.6 | 104.401 | 104.1667 | 98.15 | 104 | 104.1 | 104.18 | 103.96 | 99.9 | 102.8 | 103.45 | 103.4333 | 103.3429 | 100.5 | 102.4667 | 103.06 | 102.9 | 102.95
-1.5 | 99.15 | 104.8 | 104.429 | 104.4 | 104.2 | 98.9 | 104.2 | 104.26 | 104.076 | 104 | 99.6 | 103.8 | 103.775 | 103.86 | 103.7429 | 100.75 | 102.7 | 103.225 | 103.2625 | 103.1667 | 100.8 | 102.4 | 102.8571 | 102.75 | 102.7667
-1.4 | 100.8 | 104.5 | 104.204 | 104.1 | 103.875 | 100.4 | 103.95 | 104.04 | 103.776 | 103.775 | 100.6 | 103.5667 | 103.45 | 103.56 | 103.5 | 101.1667 | 102.5 | 102.9667 | 102.9571 | 102.9167 | 101 | 102.2 | 102.58 | 102.5 | 102.56
-1.3 | 101.9 | 104.1 | 103.854 | 103.6667 | 103.54 | 101 | 103.7333 | 103.74 | 103.526 | 103.6 | 101.15 | 103.2 | 103.2 | 103.25 | 103.2333 | 101.3 | 102.4 | 102.74 | 102.8714 | 102.64 | 101.3 | 102.075 | 102.475 | 102.3333 | 102.2667
-1.2 | 102.3 | 103.8 | 103.529 | 103.425 | 103.32 | 101.4 | 103.4 | 103.375 | 103.086 | 103.2 | 101.3 | 103 | 102.98 | 103.025 | 102.9667 | 101.4 | 102.225 | 102.54 | 102.6333 | 102.4 | 101.35 | 101.875 | 102.3 | 102.2 | 102.075
-1.1 | 102.4 | 103.45 | 103.2123 | 103.125 | 103.075 | 101.6 | 103.125 | 103.14 | 102.876 | 102.9 | 101.5 | 102.8 | 102.72 | 102.7 | 102.7 | 101.5 | 102.1 | 102.34 | 102.4333 | 102.24 | 101.35 | 101.775 | 102.0429 | 102 | 101.85
-1 | 102.4 | 103.05 | 102.879 | 102.875 | 102.725 | 101.7 | 102.7333 | 102.8667 | 102.626 | 102.65 | 101.5 | 102.5667 | 102.44 | 102.48 | 102.4333 | 101.5 | 101.775 | 102.1333 | 102.1833 | 102.0714 | 101.3 | 101.6 | 101.8 | 101.9 | 101.7
-0.9 | 102.3 | 102.7333 | 102.629 | 102.6 | 102.48 | 101.6 | 102.6 | 102.525 | 102.3927 | 102.325 | 101.5 | 102.3 | 102.08 | 102.25 | 102.1 | 101.4 | 101.6 | 101.925 | 101.9 | 101.8333 | 101.3 | 101.5 | 101.7 | 101.64 | 101.5
-0.8 | 102.1 | 102.5333 | 102.305 | 102.275 | 102.16 | 101.6 | 102.3 | 102.1667 | 102.0757 | 101.98 | 101.5 | 102.075 | 101.8333 | 101.9833 | 101.8833 | 101.35 | 101.4 | 101.7 | 101.7429 | 101.65 | 101.2 | 101.4 | 101.5 | 101.4167 | 101.375
-0.7 | 101.85 | 102.2333 | 102.04 | 101.925 | 101.8 | 101.5 | 102 | 101.9 | 101.75 | 101.775 | 101.4 | 101.8 | 101.55 | 101.76 | 101.6333 | 101.25 | 101.2 | 101.45 | 101.5 | 101.4333 | 101 | 101.275 | 101.3 | 101.25 | 101.125
-0.6 | 101.7 | 101.8333 | 101.68 | 101.675 | 101.55 | 101.3 | 101.7 | 101.65 | 101.5 | 101.6 | 101.2 | 101.45 | 101.3 | 101.45 | 101.4833 | 101.05 | 101 | 101.22 | 101.22 | 101.1 | 100.8 | 101 | 101.1571 | 101.1 | 101
-0.5 | 101.3 | 101.5 | 101.255 | 101.35 | 101.3 | 101.1 | 101.4667 | 101.3 | 101.25 | 101.3 | 101.1 | 101.2 | 101.08 | 101.2 | 101.3 | 100.9 | 100.8667 | 101 | 100.9667 | 100.85 | 100.6 | 100.875 | 101 | 100.98 | 100.8
-0.4 | 101 | 101.1667 | 101.04 | 101.0667 | 101.04 | 100.8 | 101.2 | 100.95 | 101.045 | 101.1 | 100.95 | 101 | 100.78 | 101 | 100.94 | 100.8 | 100.7 | 100.8 | 100.7 | 100.775 | 100.5 | 100.7 | 100.82 | 100.76 | 100.68
-0.3 | 100.8 | 100.8 | 100.78 | 100.82 | 100.74 | 100.6 | 100.7333 | 100.725 | 100.725 | 100.7 | 100.7 | 100.7667 | 100.58 | 100.8 | 100.7333 | 100.6 | 100.4667 | 100.54 | 100.6 | 100.5286 | 100.35 | 100.55 | 100.6 | 100.4833 | 100.4
-0.2 | 100.6 | 100.6 | 100.555 | 100.5 | 100.575 | 100.4 | 100.5 | 100.48 | 100.475 | 100.525 | 100.45 | 100.5 | 100.35 | 100.55 | 100.42 | 100.4 | 100.2 | 100.4 | 100.3667 | 100.2714 | 100.3 | 100.4 | 100.4833 | 100.3 | 100.26
-0.1 | 100.25 | 100.2333 | 100.2467 | 100.175 | 100.275 | 100.3 | 100.2333 | 100.25 | 100.15 | 100.28 | 100.2 | 100.25 | 100.2 | 100.2 | 100.1667 | 100.1 | 100.175 | 100.2 | 100.1875 | 100.1 | 100.2 | 100.2 | 100.3167 | 100.1 | 100.15
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.7 | 99.675 | 99.5801 | 99.8 | 99.8 | 99.6 | 99.725 | 99.825 | 99.7499 | 99.74 | 99.8 | 99.7333 | 99.75 | 99.8 | 99.8 | 99.9 | 99.7667 | 99.9 | 99.8 | 99.8167 | 99.7 | 99.9 | 99.8 | 99.75 | 99.8333
0.2 | 99.4 | 99.5 | 99.3201 | 99.525 | 99.56 | 99.4 | 99.5 | 99.6 | 99.3999 | 99.6 | 99.45 | 99.525 | 99.42 | 99.45 | 99.62 | 99.7 | 99.6 | 99.6667 | 99.5571 | 99.65 | 99.5 | 99.6333 | 99.58 | 99.52 | 99.6333
0.3 | 99.1 | 99.3 | 99.0802 | 99.3 | 99.425 | 99 | 99.2 | 99.25 | 99.2748 | 99.3 | 99.2 | 99.3 | 99.2 | 99.3 | 99.4333 | 99.4 | 99.175 | 99.4333 | 99.3857 | 99.4667 | 99.3 | 99.5 | 99.45 | 99.32 | 99.4
0.4 | 98.8 | 98.975 | 98.8802 | 99.02 | 99.06 | 98.8 | 98.9667 | 99.1 | 99.1248 | 99.125 | 98.9 | 99.1 | 99.025 | 99.1 | 99.1 | 99.1 | 99 | 99.2 | 99.2429 | 99.2167 | 99.15 | 99.3 | 99.325 | 99.2 | 99.25
0.5 | 98.5 | 98.6 | 98.6803 | 98.8 | 98.825 | 98.6 | 98.7667 | 98.82 | 98.8847 | 98.975 | 98.6 | 98.9 | 98.8 | 98.96 | 98.92 | 98.9 | 98.725 | 99.075 | 98.9833 | 99.0333 | 99 | 99.1 | 99.2 | 99.075 | 99.1
0.6 | 98.2 | 98.3667 | 98.4053 | 98.6 | 98.7 | 98.4 | 98.5 | 98.64 | 98.7247 | 98.8 | 98.25 | 98.775 | 98.54 | 98.725 | 98.6667 | 98.6 | 98.5 | 98.85 | 98.8167 | 98.8556 | 98.8 | 98.9333 | 99.1 | 98.95 | 99
0.7 | 98 | 98.1333 | 98.247 | 98.4 | 98.44 | 98 | 98.3333 | 98.5 | 98.4996 | 98.575 | 98 | 98.55 | 98.3 | 98.6 | 98.5 | 98.4 | 98.3 | 98.6167 | 98.5667 | 98.6571 | 98.55 | 98.725 | 99 | 98.74 | 98.8
0.8 | 97.7 | 97.9667 | 97.9804 | 98.25 | 98.3 | 97.6 | 98.1 | 98.2 | 98.2746 | 98.3 | 97.8 | 98.275 | 98.05 | 98.35 | 98.36 | 98.1 | 98.1 | 98.425 | 98.4143 | 98.4 | 98.3 | 98.6 | 98.8833 | 98.64 | 98.68
0.9 | 97.2 | 97.725 | 97.8054 | 98.02 | 98.1 | 97.1 | 97.8 | 97.96 | 98.0245 | 98.14 | 97.3 | 98 | 97.85 | 98.12 | 98.18 | 97.7 | 98 | 98.1667 | 98.275 | 98.2167 | 97.9 | 98.4 | 98.7167 | 98.5 | 98.4
1 | 96.9 | 97.4667 | 97.6472 | 97.8 | 97.95 | 96.6 | 97.625 | 97.775 | 97.7995 | 97.92 | 97 | 97.7333 | 97.7 | 97.975 | 97.9833 | 97.5 | 97.8 | 98 | 98.12 | 98.04 | 97.6 | 98.1667 | 98.6 | 98.28 | 98.2
1.1 | 96.2 | 97.2 | 97.2805 | 97.6 | 97.675 | 96.15 | 97.4333 | 97.625 | 97.5644 | 97.75 | 96.4 | 97.5 | 97.575 | 97.76 | 97.7571 | 97.2 | 97.6 | 97.9 | 97.8429 | 97.96 | 97.3 | 97.925 | 98.375 | 98.08 | 98.1
1.2 | 95.7 | 96.95 | 97.1306 | 97.4 | 97.5 | 95.5 | 97.1 | 97.3 | 97.4244 | 97.54 | 96 | 97.2333 | 97.225 | 97.55 | 97.56 | 96.8 | 97.2667 | 97.66 | 97.58 | 97.7833 | 96.9 | 97.775 | 98.1 | 97.95 | 97.98
1.3 | 94.9 | 96.7 | 96.8606 | 97.25 | 97.26 | 94.55 | 96.85 | 97.16 | 97.2243 | 97.375 | 95.4 | 97.0667 | 97 | 97.325 | 97.4 | 96.1 | 97.025 | 97.5667 | 97.4143 | 97.6286 | 96.6 | 97.5 | 97.925 | 97.7 | 97.7833
1.4 | 93.6 | 96.4333 | 96.6557 | 97.1 | 97.04 | 93.6 | 96.5667 | 96.875 | 97.0443 | 97.2 | 94.35 | 96.8 | 96.775 | 97.2 | 97.2286 | 95.7 | 96.8 | 97.35 | 97.25 | 97.3778 | 96.3 | 97.4 | 97.75 | 97.6 | 97.7
1.5 | 91.75 | 96.3 | 96.4474 | 96.7 | 96.92 | 91.8 | 96.4 | 96.625 | 96.8642 | 96.98 | 93.1 | 96.6 | 96.6 | 96.95 | 97.0833 | 94.9 | 96.6 | 97.1 | 97.0833 | 97.175 | 95.8 | 97.25 | 97.56 | 97.45 | 97.5
1.6 | 88.35 | 96 | 96.2807 | 96.54 | 96.675 | 89.7 | 96.1667 | 96.4 | 96.6242 | 96.75 | 91.6 | 96.4 | 96.325 | 96.7 | 96.9 | 94 | 96.4 | 96.98 | 96.9 | 97.02 | 94.95 | 97 | 97.4833 | 97.275 | 97.375
1.7 | 82.75 | 95.7667 | 96.0808 | 96.4 | 96.425 | 85.9 | 95.8333 | 96.25 | 96.4408 | 96.5 | 89.15 | 96.0333 | 96.175 | 96.5 | 96.6333 | 92.6 | 96.2 | 96.8 | 96.82 | 96.8571 | 94.4 | 96.775 | 97.3167 | 97.06 | 97.26
1.8 | 74.3 | 95.5667 | 95.8558 | 96.15 | 96.275 | 79.1 | 95.6333 | 96 | 96.2408 | 96.375 | 84.9 | 95.8 | 95.9667 | 96.3 | 96.46 | 91 | 95.9 | 96.7 | 96.625 | 96.7 | 93.4 | 96.45 | 97.05 | 96.975 | 97.0667
1.9 | 63.75 | 95.3 | 95.5809 | 95.95 | 96.025 | 69.8 | 95.4 | 95.78 | 95.944 | 96.125 | 79.1 | 95.5 | 95.7 | 96.1 | 96.3125 | 88.5 | 95.65 | 96.44 | 96.4375 | 96.5167 | 92.1 | 96.125 | 96.95 | 96.85 | 96.9
2 | 51.3 | 94.9667 | 95.2609 | 95.775 | 95.85 | 59.3 | 95.1 | 95.58 | 95.7239 | 95.9 | 70.6 | 95.0333 | 95.525 | 95.9167 | 96.04 | 84.25 | 95.4 | 96.16 | 96.1667 | 96.275 | 90.3 | 95.9 | 96.7667 | 96.6333 | 96.7
2.1 | 39.35 | 94.7333 | 95.0143 | 95.525 | 95.6 | 48.4 | 94.8 | 95.36 | 95.6489 | 95.76 | 60 | 94.85 | 95.32 | 95.72 | 95.84 | 78.3 | 95.1 | 95.9333 | 96.0143 | 96.04 | 87.8 | 95.725 | 96.6 | 96.4333 | 96.5
2.2 | 27.95 | 94.5 | 94.7211 | 95.26 | 95.3667 | 36.55 | 94.55 | 94.975 | 95.3989 | 95.4667 | 49.1 | 94.6 | 95.08 | 95.5 | 95.64 | 70.2 | 94.75 | 95.7 | 95.85 | 95.8714 | 83.15 | 95.575 | 96.35 | 96.2667 | 96.35
2.3 | 18.5 | 94.15 | 94.4561 | 95.1 | 95.1 | 26.4 | 94.3 | 94.8 | 95.1488 | 95.3 | 37.9 | 94.4 | 94.85 | 95.25 | 95.44 | 60.7 | 94.55 | 95.525 | 95.68 | 95.7 | 77.8 | 95.25 | 96.16 | 96.1 | 96.18
2.4 | 11.8 | 93.8667 | 94.2062 | 94.9167 | 94.8 | 17.8 | 93.9 | 94.575 | 94.7987 | 94.925 | 27.2 | 94.0667 | 94.62 | 95.025 | 95.1667 | 50 | 94.3 | 95.2667 | 95.5 | 95.5143 | 69.5 | 94.9 | 95.9833 | 95.9 | 96
2.5 | 7.9 | 93.525 | 93.9412 | 94.6 | 94.6 | 12.2 | 93.65 | 94.275 | 94.5837 | 94.8 | 19.7 | 93.7 | 94.34 | 94.775 | 95 | 39.4 | 94 | 94.95 | 95.2 | 95.325 | 60.9 | 94.7333 | 95.8 | 95.6833 | 95.8
2.6 | 5.6 | 93.2 | 93.7312 | 94.4 | 94.5 | 8.8 | 93.3 | 94.025 | 94.3636 | 94.6333 | 14.3 | 93.4 | 94.1 | 94.625 | 94.8 | 29.7 | 93.6667 | 94.7 | 95 | 95.15 | 50.45 | 94.3667 | 95.6 | 95.5167 | 95.6
2.7 | 4.4 | 93 | 93.3813 | 94.12 | 94.22 | 6.9 | 93 | 93.9 | 94.0235 | 94.45 | 10.8 | 93.1 | 93.96 | 94.375 | 94.44 | 22.7 | 93.4 | 94.54 | 94.8286 | 94.98 | 40.8 | 94.175 | 95.48 | 95.42 | 95.4
2.8 | 3.7 | 92.7333 | 93.2313 | 93.85 | 93.9333 | 5.6 | 92.7 | 93.5 | 93.9235 | 94.1 | 8.85 | 92.7333 | 93.64 | 94.2 | 94.2667 | 16.7 | 93.125 | 94.3 | 94.6 | 94.8 | 31.7 | 93.8667 | 95.2 | 95.2 | 95.18
2.9 | 3.1 | 92.2 | 92.8814 | 93.6 | 93.7 | 4.7 | 92.4 | 93.3 | 93.7234 | 93.9 | 7.5 | 92.35 | 93.4 | 93.88 | 94.1 | 13.3 | 92.9333 | 94.125 | 94.3833 | 94.4167 | 24.4667 | 93.625 | 94.9333 | 95.08 | 94.975
3 | 2.7 | 92 | 92.6565 | 93.35 | 93.5 | 4.2 | 91.8333 | 93.025 | 93.5234 | 93.6 | 6.5 | 92 | 93.16 | 93.68 | 93.8571 | 10.85 | 92.5333 | 93.8667 | 94.2286 | 94.325 | 18.35 | 93.4 | 94.78 | 94.85 | 94.8
3.1 | 2.5 | 91.65 | 92.3565 | 93.1 | 93.24 | 3.8 | 91.5667 | 92.8 | 93.2233 | 93.4 | 5.9 | 91.6333 | 92.86 | 93.5 | 93.7 | 9.5 | 92.1 | 93.6 | 94.0333 | 94.1286 | 14.9 | 93 | 94.6143 | 94.6 | 94.58
3.2 | 2.2 | 91.4 | 92.0316 | 92.825 | 93 | 3.6 | 91.2333 | 92.56 | 93.0483 | 93.2 | 5.4 | 91.3333 | 92.475 | 93.3 | 93.4667 | 8.6 | 91.6667 | 93.38 | 93.8143 | 93.86 | 12.35 | 92.6 | 94.4 | 94.3667 | 94.48
3.3 | 2.1 | 90.9667 | 91.8816 | 92.575 | 92.7333 | 3.2 | 90.8 | 92.25 | 92.8232 | 93 | 4.9 | 90.7 | 92.35 | 93 | 93.12 | 7.9 | 91.4 | 93.16 | 93.5571 | 93.6 | 11.1 | 92.4 | 94.0833 | 94.24 | 94.225
3.4 | 2 | 90.6 | 91.5817 | 92.4 | 92.52 | 3 | 90.5 | 92.02 | 92.5231 | 92.8667 | 4.6 | 90.3667 | 92.15 | 92.85 | 92.9 | 7.25 | 91.2 | 93.025 | 93.3833 | 93.5143 | 10.1667 | 92.0333 | 93.8833 | 94.025 | 94.125
3.5 | 1.95 | 90.35 | 91.2817 | 92.2 | 92.4 | 2.85 | 90.2 | 91.775 | 92.3481 | 92.575 | 4.4 | 90.1 | 91.9 | 92.64 | 92.76 | 6.9 | 90.725 | 92.66 | 93.175 | 93.3 | 9.3 | 91.725 | 93.7 | 93.88 | 94.0333
3.6 | 1.9 | 89.9333 | 91.1568 | 91.95 | 92.16 | 2.7 | 89.7 | 91.55 | 92.123 | 92.3 | 4.2 | 89.7 | 91.65 | 92.475 | 92.46 | 6.6 | 90.2333 | 92.425 | 92.925 | 93.0667 | 8.4 | 91.3 | 93.35 | 93.68 | 93.825
3.7 | 1.8 | 89.5667 | 91.0068 | 91.7 | 91.925 | 2.6 | 89.4 | 91.225 | 91.898 | 92.125 | 4 | 89.2 | 91.34 | 92.225 | 92.2667 | 6.2 | 89.8 | 92.18 | 92.6833 | 92.8 | 8.2 | 91.0667 | 93.1667 | 93.4 | 93.52
3.8 | 1.7 | 89.2667 | 90.7019 | 91.55 | 91.7 | 2.4 | 89 | 91 | 91.6629 | 91.9667 | 3.8 | 88.925 | 91.15 | 91.95 | 92.1 | 5.8 | 89.4 | 92 | 92.4333 | 92.65 | 7.6 | 90.8 | 92.9 | 93.2667 | 93.4
3.9 | 1.7 | 88.85 | 90.5819 | 91.3 | 91.4833 | 2.3 | 88.4667 | 90.8 | 91.4229 | 91.72 | 3.6 | 88.5 | 90.9 | 91.6667 | 91.8333 | 5.7 | 89.0667 | 91.7 | 92.25 | 92.4333 | 7.3 | 90.3333 | 92.72 | 92.9833 | 93.2167
4 | 1.6 | 88.5 | 90.2819 | 91.15 | 91.275 | 2.2 | 88.0667 | 90.675 | 91.1728 | 91.4 | 3.5 | 87.9333 | 90.7 | 91.5 | 91.6571 | 5.4 | 88.7 | 91.6 | 92.1 | 92.3 | 7.1 | 89.9 | 92.55 | 92.9 | 93.0167
4.1 | 1.5 | 88.1 | 90.082 | 90.92 | 91.1 | 2.1 | 87.6667 | 90.45 | 90.9427 | 91.3 | 3.3 | 87.4667 | 90.54 | 91.28 | 91.4833 | 5.2333 | 88.2333 | 91.32 | 91.9667 | 92.0429 | 6.8 | 89.5 | 92.325 | 92.7 | 92.84
4.2 | 1.5 | 87.4333 | 89.832 | 90.8 | 90.975 | 2 | 87.1667 | 90.25 | 90.8227 | 91.15 | 3.1667 | 87.0667 | 90.3 | 91.14 | 91.3286 | 5.1 | 87.85 | 90.95 | 91.7 | 91.825 | 6.55 | 89.175 | 92.12 | 92.4667 | 92.6333
4.3 | 1.4 | 86.7 | 89.5821 | 90.6 | 90.7333 | 1.9 | 86.5 | 90 | 90.6727 | 91 | 3 | 86.5667 | 89.9833 | 90.92 | 91.04 | 5 | 87.375 | 90.8 | 91.5286 | 91.5833 | 6.2 | 88.7 | 91.78 | 92.2 | 92.5
4.4 | 1.4 | 85.8 | 89.4321 | 90.425 | 90.58 | 1.9 | 85.85 | 89.8 | 90.5226 | 90.675 | 2.9 | 86.2 | 89.9 | 90.75 | 90.8333 | 4.75 | 86.9667 | 90.6667 | 91.3 | 91.4 | 6.1 | 88.3 | 91.6 | 92.04 | 92.3286
4.5 | 1.3 | 84.5 | 89.1822 | 90.1667 | 90.45 | 1.8 | 84.9333 | 89.52 | 90.2559 | 90.55 | 2.8 | 85.5667 | 89.65 | 90.52 | 90.6 | 4.6 | 86.5 | 90.46 | 91.12 | 91.3 | 5.9 | 88 | 91.4 | 91.9 | 92.1429
4.6 | 1.3 | 82.1333 | 88.9322 | 90.025 | 90.1 | 1.7 | 83.8 | 89.38 | 90.0975 | 90.3 | 2.7 | 84.9 | 89.5 | 90.38 | 90.5 | 4.5 | 86.1333 | 90.2333 | 90.88 | 91.1714 | 5.8 | 87.725 | 91.14 | 91.72 | 91.95
4.7 | 1.2 | 78.7 | 88.7822 | 89.76 | 90 | 1.6 | 82.1 | 89.15 | 89.8725 | 90.1333 | 2.6 | 84.0333 | 89.25 | 90.26 | 90.4 | 4.3 | 85.5667 | 90 | 90.6714 | 91.0167 | 5.6 | 87.2333 | 90.9333 | 91.52 | 91.75
4.8 | 1.2 | 72.6 | 88.5423 | 89.525 | 89.76 | 1.6 | 79.7 | 88.85 | 89.7224 | 89.95 | 2.6 | 83 | 89 | 90.04 | 90.18 | 4.2 | 85.05 | 89.82 | 90.5167 | 90.7571 | 5.4 | 86.8 | 90.75 | 91.3 | 91.525
4.9 | 1.2 | 64.1 | 88.3823 | 89.35 | 89.56 | 1.5 | 75.5 | 88.68 | 89.4824 | 89.7 | 2.5 | 81.6 | 88.78 | 89.7 | 89.9833 | 4.1 | 84.7 | 89.5333 | 90.36 | 90.5333 | 5.3 | 86.375 | 90.5 | 91.0667 | 91.3333
5 | 1.1 | 54.1667 | 88.0824 | 89.2 | 89.325 | 1.5 | 69.4 | 88.44 | 89.2556 | 89.5 | 2.3 | 79.475 | 88.5667 | 89.475 | 89.7333 | 4 | 83.9333 | 89.3667 | 90.1833 | 90.2833 | 5.2 | 86.0333 | 90.3 | 91 | 91.15
5.1 | 1.1 | 43.9 | 87.8574 | 88.9 | 89.075 | 1.5 | 60.7667 | 88.225 | 89.0223 | 89.325 | 2.3 | 76.4 | 88.4 | 89.25 | 89.5571 | 3.85 | 83.2 | 89.08 | 89.9571 | 90 | 5.1 | 85.5 | 90.12 | 90.88 | 90.92
5.2 | 1.1 | 33.5 | 87.6075 | 88.6 | 88.82 | 1.4 | 51.6 | 88 | 88.8822 | 89.175 | 2.2 | 71.4 | 88.1 | 89.1 | 89.28 | 3.7 | 82.4 | 88.84 | 89.775 | 89.9429 | 4.9 | 85.0667 | 89.94 | 90.62 | 90.7
5.3 | 1.1 | 25.1 | 87.3425 | 88.375 | 88.65 | 1.4 | 42.3 | 87.74 | 88.5721 | 88.925 | 2.1 | 64.2667 | 87.8667 | 89 | 89.1 | 3.6 | 81.3 | 88.62 | 89.5625 | 89.6857 | 4.8 | 84.6667 | 89.58 | 90.36 | 90.55
5.4 | 1.1 | 17.4 | 87.1076 | 88.25 | 88.4 | 1.3 | 32.95 | 87.475 | 88.3888 | 88.7 | 2.1 | 56.4333 | 87.6 | 88.6667 | 88.9 | 3.55 | 80 | 88.4 | 89.4167 | 89.5429 | 4.65 | 84 | 89.42 | 90.14 | 90.45
5.5 | 1 | 12.5333 | 86.7576 | 88 | 88.2 | 1.3 | 24.9 | 87.1667 | 88.1971 | 88.475 | 2 | 47.5667 | 87.375 | 88.5 | 88.7 | 3.4 | 77.7 | 88.1 | 89.2429 | 89.2857 | 4.5 | 83.35 | 89.1 | 90.0333 | 90.26
5.6 | 1 | 10.0333 | 86.5327 | 87.7 | 87.92 | 1.3 | 18.1 | 86.9 | 87.972 | 88.25 | 1.9 | 38.2667 | 87.1 | 88.2 | 88.4667 | 3.3 | 74.8667 | 87.875 | 89.0167 | 89.1333 | 4.35 | 82.5 | 89 | 89.8 | 90.05
5.7 | 1 | 8.4 | 86.1828 | 87.5 | 87.7 | 1.3 | 14.2 | 86.6 | 87.7219 | 88.06 | 1.85 | 30.45 | 86.825 | 87.9667 | 88.24 | 3.2 | 70.325 | 87.6 | 88.85 | 88.975 | 4.3 | 81.7667 | 88.85 | 89.6 | 89.82
5.8 | 1 | 7.5 | 86.0328 | 87.2 | 87.45 | 1.2 | 11.6 | 86.3667 | 87.4619 | 87.78 | 1.8 | 24 | 86.675 | 87.76 | 88.0333 | 3.1 | 64.3667 | 87.44 | 88.5714 | 88.82 | 4.1 | 81 | 88.7 | 89.4 | 89.675
5.9 | 0.9 | 6.8333 | 85.6829 | 86.9 | 87.225 | 1.2 | 10.3 | 86.2 | 87.3018 | 87.5 | 1.8 | 18.5 | 86.35 | 87.6 | 87.8143 | 3 | 57.9 | 87.14 | 88.3333 | 88.64 | 4.1 | 79.3 | 88.45 | 89.28 | 89.3
6 | 0.9 | 6.4 | 85.4029 | 86.7 | 86.98 | 1.1 | 9.1 | 85.85 | 87.0418 | 87.225 | 1.7 | 15.675 | 86.075 | 87.4 | 87.5857 | 2.95 | 49.5333 | 86.925 | 88.1143 | 88.44 | 4 | 77.4 | 88.2333 | 89 | 89.1667
6.1 | 0.9 | 6 | 85.108 | 86.525 | 86.72 | 1.1 | 8.4667 | 85.6 | 86.8467 | 87.1 | 1.7 | 13.65 | 85.72 | 87.1 | 87.3667 | 2.85 | 41.6 | 86.575 | 87.8833 | 88.1333 | 3.9 | 74.64 | 87.9333 | 88.9 | 89.08
6.2 | 0.9 | 5.7333 | 84.9497 | 86.275 | 86.44 | 1.1 | 7.95 | 85.3 | 86.6217 | 86.8 | 1.6 | 12.3333 | 85.45 | 86.8333 | 87.2 | 2.8 | 34.4 | 86.35 | 87.7167 | 87.9 | 3.8 | 70.35 | 87.6 | 88.76 | 88.9
6.3 | 0.9 | 5.5333 | 84.6581 | 86.05 | 86.15 | 1.1 | 7.4667 | 85.1 | 86.3216 | 86.575 | 1.6 | 11.4333 | 85.2 | 86.68 | 86.9667 | 2.7 | 27.9333 | 85.98 | 87.55 | 87.7667 | 3.7 | 64.8 | 87.3 | 88.625 | 88.7167
6.4 | 0.9 | 5.4 | 84.3831 | 85.775 | 85.975 | 1.1 | 7.1 | 84.8 | 86.0615 | 86.25 | 1.5 | 10.7 | 84.925 | 86.55 | 86.6714 | 2.7 | 23.3333 | 85.8 | 87.2333 | 87.42 | 3.7 | 58.175 | 87.1 | 88.375 | 88.6167
6.5 | 0.9 | 5.2 | 84.1332 | 85.525 | 85.725 | 1.1 | 6.8333 | 84.44 | 85.9215 | 86.1 | 1.5 | 10.2333 | 84.56 | 86.3 | 86.4333 | 2.6 | 20 | 85.52 | 86.9286 | 87.1833 | 3.6 | 51.3667 | 86.86 | 88.125 | 88.4
6.6 | 0.8 | 5 | 83.8832 | 85.2333 | 85.5667 | 1.1 | 6.5333 | 84.2 | 85.5214 | 85.9 | 1.5 | 9.7667 | 84.25 | 86.02 | 86.2429 | 2.6 | 17.7333 | 85.3 | 86.7667 | 87.0167 | 3.5 | 43.4 | 86.66 | 87.92 | 88.2
6.7 | 0.8 | 4.9 | 83.6433 | 84.975 | 85.325 | 1 | 6.4 | 84.025 | 85.3013 | 85.65 | 1.5 | 9.4333 | 84.025 | 85.7667 | 85.9333 | 2.5 | 16.25 | 85.025 | 86.6 | 86.7667 | 3.4 | 36.4667 | 86.4 | 87.7429 | 87.88
6.8 | 0.8 | 4.8 | 83.2334 | 84.8 | 85.06 | 1 | 6.2 | 83.7 | 85.0713 | 85.4 | 1.4 | 9.05 | 83.725 | 85.46 | 85.725 | 2.5 | 15.4333 | 84.8333 | 86.3857 | 86.5857 | 3.3 | 30.5 | 86.1 | 87.54 | 87.7
6.9 | 0.8 | 4.6 | 82.9584 | 84.6 | 84.84 | 1 | 5.975 | 83.46 | 84.7012 | 85.14 | 1.4 | 8.65 | 83.5 | 85.3 | 85.5167 | 2.5 | 14.6 | 84.625 | 86.18 | 86.3667 | 3.2 | 26.0667 | 85.82 | 87.3 | 87.5667
7 | 0.8 | 4.5333 | 82.7085 | 84.26 | 84.5 | 0.9 | 5.8 | 83.14 | 84.5211 | 84.9 | 1.3 | 8.5 | 83.16 | 85.125 | 85.3143 | 2.4 | 14.0333 | 84.2667 | 86 | 86.1667 | 3.2 | 22.8667 | 85.5 | 87.1667 | 87.2833
7.1 |  | 4.4 | 82.4585 | 84.025 | 84.375 | 1 | 5.55 | 82.8 | 84.3011 | 84.675 | 1.3 | 8.15 | 82.925 | 84.82 | 85.1 | 2.4 | 13.4667 | 83.9333 | 85.74 | 85.9286 | 3.1 | 20.5 | 85.2667 | 86.9333 | 87.075
7.2 |  | 4.375 | 82.1836 | 83.7 | 84.0333 |  | 5.4 | 82.5 | 84.071 | 84.425 |  | 7.925 | 82.62 | 84.55 | 84.84 | 2.35 | 12.9667 | 83.55 | 85.5333 | 85.8 | 3.1 | 19.1333 | 85.1 | 86.7 | 86.88
7.3 |  | 4.2 | 81.8336 | 83.58 | 83.82 |  | 5.3 | 82.22 | 83.921 | 84.22 |  | 7.7 | 82.375 | 84.2 | 84.62 | 2.3 | 12.5333 | 83.2 | 85.2286 | 85.5 | 3 | 18 | 84.7667 | 86.46 | 86.6
7.4 |  | 4.1 | 81.5637 | 83.25 | 83.6 |  | 5.1 | 81.95 | 83.5959 | 83.9 |  | 7.4 | 81.98 | 84.06 | 84.3667 | 2 | 12.2333 | 82.96 | 85.0667 | 85.0833 | 2.8 | 17.3 | 84.575 | 86.28 | 86.425
7.5 |  | 4 | 81.3087 | 83.02 | 83.325 |  | 5 | 81.6 | 83.3809 | 83.65 |  | 7.2 | 81.7333 | 83.875 | 84.1 |  | 12 | 82.62 | 84.85 | 84.9375 | 2.65 | 16.7667 | 84.25 | 86.06 | 86.2
7.6 |  | 4 | 80.9588 | 82.86 | 83.05 |  | 4.9 | 81.3667 | 83.1458 | 83.45 |  | 7.1 | 81.4333 | 83.64 | 83.8333 |  | 11.7 | 82.38 | 84.5667 | 84.7667 | 2.4 | 16.3 | 83.88 | 85.75 | 85.9667
7.7 |  | 3.9 | 80.7338 | 82.52 | 82.8 |  | 4.7 | 80.98 | 82.954 | 83.24 |  | 6.9 | 81.18 | 83.4 | 83.52 |  | 11.3333 | 82.2 | 84.3833 | 84.6143 |  | 15.6333 | 83.65 | 85.525 | 85.7333
7.8 |  | 3.8 | 80.3839 | 82.275 | 82.525 |  | 4.6 | 80.75 | 82.6957 | 83 |  | 6.7 | 80.95 | 83.2 | 83.34 |  | 11.0667 | 81.9 | 84.2167 | 84.42 |  | 15.3 | 83.38 | 85.4 | 85.54
7.9 |  | 3.7 | 80.134 | 82 | 82.3 |  | 4.5 | 80.4 | 82.4206 | 82.65 |  | 6.5 | 80.5 | 82.9 | 83.2 |  | 10.875 | 81.55 | 83.9667 | 84.24 |  | 14.85 | 83.175 | 85.22 | 85.3
8 |  | 3.7 | 79.809 | 81.7333 | 82.1 |  | 4.4 | 80.18 | 82.1872 | 82.4 |  | 6.4 | 80.225 | 82.7167 | 82.8857 |  | 10.6 | 81.26 | 83.7 | 84.02 |  | 14.5 | 82.8667 | 84.925 | 85.175
8.1 |  | 3.6 | 79.5174 | 81.575 | 81.82 |  | 4.3 | 79.8 | 81.9205 | 82.16 |  | 6.2 | 79.775 | 82.4 | 82.7333 |  | 10.4 | 80.92 | 83.5 | 83.7 |  | 14.2667 | 82.7 | 84.8 | 84.96
8.2 |  | 3.5 | 79.2842 | 81.3 | 81.625 |  | 4.22 | 79.4667 | 81.6704 | 81.94 |  | 6.1 | 79.56 | 82.25 | 82.54 |  | 10.1 | 80.525 | 83.2 | 83.3857 |  | 13.9 | 82.35 | 84.525 | 84.75
8.3 |  | 3.5 | 78.8842 | 81.025 | 81.3 |  | 4.2 | 79.125 | 81.4404 | 81.825 |  | 5.9667 | 79.25 | 81.92 | 82.1667 |  | 9.9 | 80.2 | 82.9667 | 83.2125 |  | 13.7 | 81.95 | 84.3 | 84.6
8.4 |  | 3.4 | 78.551 | 80.8 | 81.075 |  | 4.1 | 78.9 | 81.187 | 81.5 |  | 5.8 | 79 | 81.7 | 81.9667 |  | 9.725 | 79.8 | 82.7429 | 83.0667 |  | 13.4 | 81.65 | 84.04 | 84.38
8.5 |  | 3.4 | 78.2843 | 80.46 | 80.86 |  | 4 | 78.375 | 80.9202 | 81.25 |  | 5.7 | 78.52 | 81.4 | 81.6833 |  | 9.3667 | 79.6 | 82.4571 | 82.8375 |  | 13.1333 | 81.32 | 83.8667 | 84.1667
8.6 |  | 3.3 | 77.9344 | 80.25 | 80.625 |  | 3.9 | 78.0333 | 80.7202 | 81.1 |  | 5.5 | 78.1333 | 81.16 | 81.5 |  | 9.3 | 79.28 | 82.2 | 82.6429 |  | 12.9 | 81.04 | 83.7 | 83.8
8.7 |  | 3.3 | 77.5595 | 80.1 | 80.3 |  | 3.8333 | 77.8 | 80.4801 | 80.75 |  | 5.3333 | 77.74 | 80.94 | 81.32 |  | 9.05 | 78.94 | 82 | 82.4 |  | 12.55 | 80.7 | 83.28 | 83.65
8.8 |  | 3.2 | 77.2845 | 79.75 | 80.05 |  | 3.7 | 77.4 | 80.1951 | 80.5 |  | 5.2667 | 77.4 | 80.78 | 81.0333 |  | 8.9 | 78.55 | 81.9 | 82.0667 |  | 12.3 | 80.2 | 83.1 | 83.42
8.9 |  | 3.2 | 76.8513 | 79.525 | 79.8 |  | 3.7 | 77.125 | 79.92 | 80.35 |  | 5.18 | 77 | 80.475 | 80.7286 |  | 8.625 | 78.2 | 81.6333 | 81.8 |  | 12.1 | 80.05 | 82.925 | 83.2
9 |  | 3.1 | 76.5347 | 79.3167 | 79.575 |  | 3.6 | 76.7 | 79.6199 | 80.025 |  | 5 | 76.6 | 80.16 | 80.4667 |  | 8.4333 | 77.92 | 81.3 | 81.5667 |  | 11.9 | 79.8 | 82.7833 | 82.96
9.1 |  | 3.1 | 76.0848 | 79 | 79.35 |  | 3.6 | 76.35 | 79.5199 | 79.76 |  | 5 | 76.2 | 79.92 | 80.3 |  | 8.3333 | 77.55 | 81.15 | 81.4 |  | 11.7 | 79.5333 | 82.58 | 82.66
9.2 |  | 3 | 75.7248 | 78.875 | 79.1333 |  | 3.5 | 75.925 | 79.2448 | 79.45 |  | 4.8667 | 75.9 | 79.675 | 79.9571 |  | 8.0333 | 77.2 | 80.925 | 81.1333 |  | 11.5 | 79.3 | 82.35 | 82.525
9.3 |  | 3 | 75.1849 | 78.7 | 78.9 |  | 3.4 | 75.45 | 78.9998 | 79.275 |  | 4.7667 | 75.525 | 79.46 | 79.7667 |  | 7.9 | 76.85 | 80.68 | 80.8833 |  | 11.2 | 79 | 82.12 | 82.3
9.4 |  | 2.9 | 74.6851 | 78.475 | 78.76 |  | 3.3333 | 75 | 78.7797 | 79.025 |  | 4.675 | 75.1 | 79.225 | 79.5167 |  | 7.75 | 76.5 | 80.3667 | 80.7 |  | 10.925 | 78.3333 | 81.9 | 82.1167
9.5 |  | 2.9 | 73.8602 | 78.24 | 78.6 |  | 3.3 | 74.44 | 78.5446 | 78.925 |  | 4.6 | 74.65 | 79.02 | 79.35 |  | 7.6 | 76.12 | 80.1 | 80.4286 |  | 10.6667 | 77.975 | 81.7333 | 81.84
9.6 |  | 2.9 | 72.8355 | 78.08 | 78.3667 |  | 3.2 | 73.9 | 78.2996 | 78.68 |  | 4.4667 | 74.25 | 78.8 | 79.16 |  | 7.6 | 75.68 | 79.9 | 80.25 |  | 10.5 | 77.72 | 81.54 | 81.575
9.7 |  | 2.8 | 71.0858 | 77.95 | 78.2 |  | 3.2 | 73.3 | 78.1696 | 78.4 |  | 4.35 | 73.8 | 78.54 | 78.8 |  | 7.45 | 75.4 | 79.5833 | 79.9714 |  | 10.4 | 77.46 | 81.28 | 81.4
9.8 |  | 2.8 | 68.3113 | 77.82 | 78 |  | 3.15 | 72.6 | 77.9595 | 78.225 |  | 4.3 | 73.4 | 78.32 | 78.5 |  | 7.2 | 75.05 | 79.4286 | 79.72 |  | 10.2667 | 77.12 | 81 | 81.1
9.9 |  | 2.8 | 63.3873 | 77.675 | 77.9 |  | 3.1 | 71.6 | 77.8195 | 78.0333 |  | 4.2 | 72.9 | 78.1 | 78.4 |  | 7.1 | 74.7 | 79.1571 | 79.5143 |  | 10.1 | 76.725 | 80.8 | 81
10 |  | 2.8 | 56.5887 | 77.5 | 77.8 |  | 3.1 | 70.05 | 77.7194 | 77.9 |  | 4.1 | 72.46 | 77.92 | 78.16 |  | 6.9667 | 74.32 | 78.9 | 79.3 |  | 9.9 | 76.3857 | 80.5286 | 80.76
10.1 |  |  | 49.3901 | 77.3 | 77.7 |  | 3 | 67.64 | 77.5194 | 77.775 |  | 4.1 | 71.84 | 77.68 | 77.9 |  | 6.8 | 73.9 | 78.7333 | 79 |  | 9.7667 | 76.06 | 80.2 | 80.54
10.2 |  |  | 40.9918 | 77.125 | 77.5 |  | 3 | 63.975 | 77.3193 | 77.62 |  | 4 | 71.16 | 77.4333 | 77.7714 |  | 6.7 | 73.5 | 78.5333 | 78.725 |  | 9.6 | 75.675 | 80 | 80.25
10.3 |  |  | 32.8735 | 77.075 | 77.4 |  |  | 58.5 | 77.1943 | 77.5 |  | 3.9 | 70.3 | 77.22 | 77.6 |  | 6.5 | 73.14 | 78.3143 | 78.525 |  | 9.3667 | 75.46 | 79.8 | 80
10.4 |  |  | 25.4549 | 76.95 | 77.2 |  |  | 51.75 | 77.0193 | 77.4 |  | 3.9 | 69.275 | 77.1 | 77.42 |  | 6.4 | 72.6 | 78.0667 | 78.2333 |  | 9.175 | 75.14 | 79.6333 | 79.9
10.5 |  |  | 19.4961 | 76.8 | 77.1 |  |  | 44.4 | 76.9192 | 77.26 |  |  | 67.875 | 76.9167 | 77.14 |  | 6.3667 | 72.1667 | 77.7857 | 78.0833 |  | 9 | 74.75 | 79.3667 | 79.68
10.6 |  |  | 15.6969 | 76.675 | 76.95 |  |  | 36.8 | 76.7692 | 77.05 |  |  | 66.1 | 76.775 | 76.96 |  | 6.2 | 71.725 | 77.5857 | 77.8 |  | 8.95 | 74.4 | 79.1 | 79.46
10.7 |  |  | 13.4973 | 76.525 | 76.8 |  |  | 29.66 | 76.6392 | 76.9 |  |  | 62.3 | 76.64 | 76.8667 |  | 6.1 | 71.22 | 77.3 | 77.6429 |  | 8.8 | 73.95 | 78.775 | 79.1714
10.8 |  |  | 12.1976 | 76.4 | 76.725 |  |  | 23.875 | 76.4191 | 76.7333 |  |  | 57.38 | 76.45 | 76.7 |  | 6 | 70.9 | 77.1 | 77.4667 |  | 8.6667 | 73.5333 | 78.65 | 79
10.9 |  |  | 11.5227 | 76.26 | 76.5 |  |  | 19.15 | 76.3191 | 76.58 |  |  | 51.5 | 76.3 | 76.6286 |  | 5.9 | 70.475 | 76.92 | 77.2167 |  | 8.5 | 73.28 | 78.4833 | 78.8
11 |  |  | 10.8978 | 76.1 | 76.36 |  |  | 16.3 | 76.1991 | 76.46 |  |  | 44.575 | 76.15 | 76.4143 |  |  | 69.85 | 76.7 | 76.95 |  | 8.4 | 72.9 | 78.2333 | 78.5333
11.1 |  |  | 10.4979 | 75.98 | 76.2 |  |  | 14.5 | 75.979 | 76.34 |  |  | 37.925 | 76 | 76.2143 |  |  | 69.25 | 76.4571 | 76.7429 |  | 8.275 | 72.5 | 77.84 | 78.32
11.2 |  |  | 10.198 | 75.85 | 76.08 |  |  | 13.45 | 75.7189 | 76.2 |  |  | 31.6 | 75.8 | 76.05 |  |  | 68.575 | 76.3143 | 76.55 |  | 8.2 | 72.14 | 77.65 | 78.1
11.3 |  |  | 9.9647 | 75.74 | 75.9 |  |  | 12.65 | 75.6939 | 76 |  |  | 26 | 75.6 | 75.8 |  |  | 67.8333 | 76.0571 | 76.35 |  | 8.1 | 71.72 | 77.5 | 77.8333
11.4 |  |  | 9.7231 | 75.5 | 75.8 |  |  | 12.08 | 75.5189 | 75.82 |  |  | 21.92 | 75.42 | 75.7 |  |  | 66.8 | 75.85 | 76.2 |  | 8 | 71.325 | 77.22 | 77.62
11.5 |  |  | 9.5731 | 75.4 | 75.625 |  |  | 11.58 | 75.3188 | 75.7 |  |  | 19 | 75.3 | 75.6 |  |  | 65.4 | 75.6875 | 76.0333 |  |  | 70.9 | 76.98 | 77.425
11.6 |  |  | 9.3381 | 75.2 | 75.5 |  |  | 11.35 | 75.2188 | 75.5 |  |  | 17.18 | 75.1 | 75.4 |  |  | 63.4 | 75.5 | 75.75 |  |  | 70.5 | 76.8 | 77.22
11.7 |  |  | 9.1982 | 75 | 75.35 |  |  | 11.05 | 75.0521 | 75.375 |  |  | 16 | 74.975 | 75.25 |  |  | 60.5 | 75.2833 | 75.5833 |  |  | 70 | 76.66 | 76.95
11.8 |  |  | 9.0982 | 74.9 | 75.2 |  |  | 10.775 | 74.9187 | 75.25 |  |  | 15.25 | 74.8 | 75.1 |  |  | 56.34 | 75.1286 | 75.4 |  |  | 69.6 | 76.42 | 76.72
11.9 |  |  | 8.8982 | 74.8 | 75.05 |  |  | 10.525 | 74.6987 | 75.1 |  |  | 14.65 | 74.675 | 74.8571 |  |  | 51.28 | 74.9667 | 75.1167 |  |  | 69.2 | 76.275 | 76.525
12 |  |  | 8.7982 | 74.5667 | 74.8 |  |  | 10.275 | 74.5937 | 74.9 |  |  | 14.05 | 74.45 | 74.7 |  |  | 45.55 | 74.8143 | 74.95 |  |  | 68.6 | 76.025 | 76.2667
12.1 |  |  | 8.6733 | 74.45 | 74.675 |  |  | 10.1 | 74.4186 | 74.8 |  |  | 13.7 | 74.225 | 74.5 |  |  | 39.7 | 74.64 | 74.725 |  |  | 67.9 | 75.7 | 76.05
12.2 |  |  | 8.5316 | 74.24 | 74.525 |  |  | 9.85 | 74.2186 | 74.65 |  |  | 13.4 | 74.1 | 74.3 |  |  | 33.8 | 74.4333 | 74.5571 |  |  | 67.34 | 75.6 | 75.85
12.3 |  |  | 8.3983 | 74.1 | 74.4 |  |  | 9.68 | 74.1185 | 74.5 |  |  | 13.06 | 73.975 | 74.12 |  |  | 29.1333 | 74.225 | 74.4 |  |  | 66.5857 | 75.3 | 75.72
12.4 |  |  | 8.2733 | 74 | 74.25 |  |  | 9.55 | 73.8685 | 74.26 |  |  | 12.775 | 73.76 | 73.9143 |  |  | 25.425 | 73.88 | 74.2429 |  |  | 65.46 | 75.2 | 75.48
12.5 |  |  | 8.1984 | 73.8 | 74.075 |  |  | 9.375 | 73.7851 | 74.05 |  |  | 12.54 | 73.66 | 73.7 |  |  | 22.8 | 73.7 | 74.08 |  |  | 64 | 74.85 | 75.3667
12.6 |  |  | 8.0984 | 73.65 | 73.8333 |  |  | 9.2 | 73.6184 | 73.9 |  |  | 12.2 | 73.35 | 73.52 |  |  | 21.2 | 73.5833 | 73.9 |  |  | 62.2 | 74.675 | 75.1
12.7 |  |  | 7.9984 | 73.525 | 73.6833 |  |  | 9.05 | 73.4684 | 73.7333 |  |  | 12 | 73.16 | 73.44 |  |  | 20.05 | 73.4 | 73.7333 |  |  | 59.325 | 74.575 | 74.9
12.8 |  |  | 7.7984 | 73.375 | 73.575 |  |  | 8.9 | 73.2183 | 73.5 |  |  | 11.82 | 73.04 | 73.2 |  |  | 19.34 | 73.1333 | 73.5 |  |  | 55.8 | 74.44 | 74.72
12.9 |  |  | 7.7485 | 73.125 | 73.4 |  |  | 8.725 | 73.0183 | 73.325 |  |  | 11.575 | 72.8 | 73.025 |  |  | 18.7 | 72.95 | 73.3167 |  |  | 51 | 74.3 | 74.475
13 |  |  | 7.6985 | 72.975 | 73.25 |  |  | 8.6333 | 72.8432 | 73.16 |  |  | 11.34 | 72.5667 | 72.8667 |  |  | 18.1 | 72.7667 | 73.0429 |  |  | 46.3 | 74.16 | 74.3
13.1 |  |  | 7.5985 | 72.84 | 73.1 |  |  | 8.5167 | 72.6932 | 73 |  |  | 11.1333 | 72.36 | 72.5571 |  |  | 17.64 | 72.6667 | 72.86 |  |  | 41.375 | 73.9 | 74.2
13.2 |  |  | 7.4985 | 72.7 | 72.94 |  |  | 8.35 | 72.5431 | 72.78 |  |  | 10.9 | 72.225 | 72.38 |  |  | 17.24 | 72.4167 | 72.675 |  |  | 36.5 | 73.7667 | 73.95
13.3 |  |  | 7.3985 | 72.6 | 72.7 |  |  | 8.2333 | 72.3181 | 72.65 |  |  | 10.7667 | 72 | 72.3 |  |  | 16.8 | 72.26 | 72.45 |  |  | 31.95 | 73.475 | 73.72
13.4 |  |  | 7.2985 | 72.4 | 72.5333 |  |  | 8.1 | 72.158 | 72.4 |  |  | 10.6 | 71.75 | 72.0167 |  |  | 16.56 | 72.0429 | 72.2857 |  |  | 28.5667 | 73.325 | 73.5
13.5 |  |  | 7.1486 | 72.2 | 72.4 |  |  | 8 | 71.958 | 72.225 |  |  | 10.45 | 71.62 | 71.8333 |  |  | 16.25 | 71.7333 | 72.1 |  |  | 26.1 | 73.04 | 73.34
13.6 |  |  | 7.0986 | 72 | 72.15 |  |  | 7.9 | 71.7179 | 72.0667 |  |  | 10.1833 | 71.4 | 71.58 |  |  | 16 | 71.625 | 71.9 |  |  | 24.36 | 72.86 | 73.2
13.7 |  |  | 6.9986 | 71.8 | 72 |  |  | 7.8 | 71.6179 | 71.9 |  |  | 10.1 | 71.25 | 71.3833 |  |  | 15.76 | 71.4286 | 71.7 |  |  | 23.1 | 72.65 | 72.88
13.8 |  |  | 6.9586 | 71.625 | 71.88 |  |  | 7.7 | 71.4179 | 71.675 |  |  | 9.925 | 70.98 | 71.2 |  |  | 15.525 | 71.1667 | 71.52 |  |  | 22.225 | 72.425 | 72.775
13.9 |  |  | 6.8236 | 71.375 | 71.65 |  |  | 7.6 | 71.1678 | 71.45 |  |  | 9.7667 | 70.7 | 71.02 |  |  | 15.2 | 70.9167 | 71.25 |  |  | 21.575 | 72.26 | 72.5
14 |  |  | 6.7986 | 71.2 | 71.4 |  |  | 7.475 | 70.9377 | 71.22 |  |  | 9.625 | 70.5333 | 70.75 |  |  | 14.98 | 70.74 | 71.05 |  |  | 21.075 | 71.9714 | 72.32
14.1 |  |  | 6.6987 | 71.025 | 71.125 |  |  | 7.4 | 70.7844 | 70.95 |  |  | 9.4 | 70.3 | 70.5667 |  |  | 14.7667 | 70.575 | 70.7667 |  |  | 20.575 | 71.8167 | 72.1
14.2 |  |  | 6.5987 | 70.675 | 70.9 |  |  | 7.3 | 70.5176 | 70.775 |  |  | 9.3 | 70.0667 | 70.2333 |  |  | 14.45 | 70.35 | 70.48 |  |  | 20.2 | 71.6 | 71.975
14.3 |  |  | 6.4987 | 70.425 | 70.55 |  |  | 7.2 | 70.2426 | 70.5333 |  |  | 9.1 | 69.84 | 70.04 |  |  | 14.2 | 70.0833 | 70.2833 |  |  | 19.9 | 71.4 | 71.825
14.4 |  |  | 6.4987 | 70.0667 | 70.25 |  |  | 7.1 | 70.0508 | 70.28 |  |  | 9 | 69.5833 | 69.76 |  |  | 14.05 | 69.8 | 70.1 |  |  | 19.6 | 71.1667 | 71.58
14.5 |  |  | 6.4987 | 69.6 | 69.8 |  |  | 7 | 69.7174 | 69.96 |  |  | 8.9 | 69.4167 | 69.56 |  |  | 13.8 | 69.62 | 69.84 |  |  | 19.24 | 71.025 | 71.35
14.6 |  |  | 6.3987 | 68.9 | 69 |  |  | 6.9 | 69.4174 | 69.7 |  |  | 8.78 | 69.15 | 69.3143 |  |  | 13.6 | 69.4 | 69.64 |  |  | 18.925 | 70.8 | 71.22
14.7 |  |  | 6.2987 | 67.9 | 68.15 |  |  | 6.9 | 69.1506 | 69.4 |  |  | 8.6667 | 68.9 | 69.1667 |  |  | 13.375 | 69.16 | 69.46 |  |  | 18.675 | 70.6 | 71
14.8 |  |  | 6.1988 | 66.4 | 66.6 |  |  | 6.8 | 68.7172 | 69.1 |  |  | 8.6 | 68.6 | 68.8 |  |  | 13.125 | 68.9667 | 69.2 |  |  | 18.3667 | 70.3 | 70.75
14.9 |  |  | 6.1988 | 63.8 | 64.1 |  |  | 6.7 | 68.2671 | 68.5 |  |  | 8.4 | 68.275 | 68.5667 |  |  | 13 | 68.7667 | 68.8857 |  |  | 18.175 | 70.08 | 70.5
15 |  |  | 6.0988 | 59.1 | 59.5 |  |  | 6.6 | 67.717 | 68.1 |  |  | 8.3 | 68.025 | 68.32 |  |  | 12.82 | 68.6 | 68.6833 |  |  | 17.94 | 69.9 | 70.3
15.1 |  |  | 5.9988 | 53.2 | 53.9 |  |  | 6.575 | 66.9167 | 67.275 |  |  | 8.2 | 67.775 | 68.02 |  |  | 12.6 | 68.3 | 68.5 |  |  | 17.72 | 69.6 | 70.025
15.2 |  |  |  | 46.55 | 47.22 |  |  | 6.5 | 65.9665 | 66.22 |  |  | 8.1 | 67.5 | 67.6143 |  |  | 12.42 | 68.1167 | 68.3 |  |  | 17.5 | 69.4 | 69.85
15.3 |  |  |  | 39.325 | 40.175 |  |  | 6.4 | 64.5162 | 64.8333 |  |  | 8 | 67.175 | 67.2667 |  |  | 12.3 | 67.8333 | 68.1 |  |  | 17.16 | 69.26 | 69.7
15.4 |  |  |  | 32.35 | 33.0667 |  |  | 6.375 | 62.1155 | 62.7 |  |  | 7.9 | 66.74 | 66.9333 |  |  | 12.1 | 67.5 | 67.8333 |  |  | 16.875 | 68.95 | 69.44
15.5 |  |  |  | 26.125 | 27.14 |  |  |  | 58.2896 | 58.875 |  |  | 7.82 | 66.35 | 66.64 |  |  | 11.9 | 67.3 | 67.54 |  |  | 16.6667 | 68.8 | 69.2
15.6 |  |  |  | 20.95 | 22.35 |  |  |  | 52.9732 | 53.6 |  |  | 7.72 | 65.9333 | 66.12 |  |  | 11.72 | 67.02 | 67.28 |  |  | 16.5 | 68.525 | 68.95
15.7 |  |  |  | 17.8 | 19.2 |  |  |  | 46.7117 | 47.8667 |  |  | 7.6 | 65.3667 | 65.62 |  |  | 11.6 | 66.75 | 67 |  |  | 16.3 | 68.3 | 68.72
15.8 |  |  |  | 16 | 17.425 |  |  |  | 40.085 | 41.6 |  |  | 7.5 | 64.8 | 65.05 |  |  | 11.5 | 66.46 | 66.7143 |  |  | 16.12 | 68.1 | 68.42
15.9 |  |  |  | 14.9333 | 16.5 |  |  |  | 33.9335 | 35 |  |  | 7.4 | 63.96 | 64.2 |  |  | 11.35 | 66.1714 | 66.5 |  |  | 15.875 | 67.8 | 68.2
16 |  |  |  | 14.25 | 15.7 |  |  |  | 28.157 | 29.475 |  |  |  | 62.72 | 63.0333 |  |  | 11.1 | 65.92 | 66.2125 |  |  | 15.64 | 67.62 | 67.975
16.1 |  |  |  | 13.7 | 15.28 |  |  |  | 22.9057 | 24.54 |  |  |  | 61.025 | 61.3 |  |  | 10.9 | 65.7143 | 65.95 |  |  | 15.44 | 67.3 | 67.74
16.2 |  |  |  | 13.38 | 14.95 |  |  |  | 19.755 | 21.1 |  |  |  | 58.3 | 58.6 |  |  | 10.8 | 65.4 | 65.5333 |  |  | 15.2 | 67.1 | 67.6
16.3 |  |  |  | 13.15 | 14.7 |  |  |  | 17.6795 | 19.32 |  |  |  | 54.4 | 54.7 |  |  | 10.7 | 65.075 | 65.32 |  |  | 15 | 66.875 | 67.375
16.4 |  |  |  | 12.94 | 14.45 |  |  |  | 16.5291 | 18.05 |  |  |  | 49 | 49.6667 |  |  | 10.56 | 64.7667 | 65.1 |  |  | 14.875 | 66.625 | 66.96
16.5 |  |  |  | 12.8 | 14.3 |  |  |  | 15.6372 | 17.25 |  |  |  | 43.725 | 44.1 |  |  | 10.5 | 64.5167 | 64.6833 |  |  | 14.7 | 66.4 | 66.8
16.6 |  |  |  | 12.6 | 14.15 |  |  |  | 15.0788 | 16.6 |  |  |  | 37.64 | 38.62 |  |  | 10.3 | 64.06 | 64.4333 |  |  | 14.48 | 66.2333 | 66.625
16.7 |  |  |  | 12.475 | 14.025 |  |  |  | 14.6537 | 16.175 |  |  |  | 32.46 | 33.12 |  |  | 10.2 | 63.78 | 64.14 |  |  | 14.275 | 65.9 | 66.36
16.8 |  |  |  | 12.325 | 13.9 |  |  |  | 14.3036 | 15.85 |  |  |  | 27.3 | 28.6 |  |  | 10.0667 | 63.38 | 63.82 |  |  | 14.1 | 65.7 | 66.1
16.9 |  |  |  | 12.2 | 13.8 |  |  |  | 14.0285 | 15.5 |  |  |  | 23.4 | 24.9 |  |  | 9.9333 | 63.15 | 63.36 |  |  | 13.9 | 65.36 | 65.875
17 |  |  |  | 12.1 | 13.625 |  |  |  | 13.7368 | 15.38 |  |  |  | 20.925 | 22.32 |  |  |  | 62.7 | 62.82 |  |  | 13.74 | 65.1 | 65.6
17.1 |  |  |  | 12 | 13.55 |  |  |  | 13.5701 | 15.15 |  |  |  | 19.36 | 20.86 |  |  |  | 62.2857 | 62.4571 |  |  | 13.625 | 64.88 | 65.28
17.2 |  |  |  | 11.9 | 13.42 |  |  |  | 13.4034 | 14.9 |  |  |  | 18.375 | 19.7 |  |  |  | 61.64 | 61.9333 |  |  | 13.36 | 64.625 | 65
17.3 |  |  |  | 11.8 | 13.3 |  |  |  | 13.27 | 14.8 |  |  |  | 17.6 | 18.98 |  |  |  | 61.1 | 61.32 |  |  | 13.2 | 64.36 | 64.6333
17.4 |  |  |  | 11.7 | 13.2 |  |  |  | 13.0833 | 14.6333 |  |  |  | 17.12 | 18.46 |  |  |  | 60.1571 | 60.54 |  |  | 13.05 | 63.98 | 64.35
17.5 |  |  |  | 11.6 | 13.1 |  |  |  | 12.8532 | 14.46 |  |  |  | 16.65 | 18.1667 |  |  |  | 59.06 | 59.3857 |  |  | 12.825 | 63.725 | 64.02
17.6 |  |  |  | 11.5 | 13.05 |  |  |  | 12.7032 | 14.3 |  |  |  | 16.25 | 17.72 |  |  |  | 57.6 | 57.9 |  |  | 12.7 | 63.6 | 63.85
17.7 |  |  |  | 11.4 | 12.9 |  |  |  | 12.5365 | 14.175 |  |  |  | 16 | 17.4667 |  |  |  | 54.9333 | 55.5 |  |  | 12.56 | 63.15 | 63.55
17.8 |  |  |  | 11.3 | 12.86 |  |  |  | 12.4031 | 14 |  |  |  | 15.725 | 17.2 |  |  |  | 51.3 | 52.12 |  |  | 12.325 | 62.875 | 63.25
17.9 |  |  |  | 11.25 | 12.7 |  |  |  | 12.3031 | 13.8 |  |  |  | 15.45 | 16.94 |  |  |  | 47.2667 | 47.92 |  |  | 12.2167 | 62.625 | 62.84
18 |  |  |  | 11.2 | 12.7 |  |  |  | 12.2031 | 13.7 |  |  |  | 15.2 | 16.6857 |  |  |  | 42.82 | 43.6 |  |  |  | 62.3 | 62.54
18.1 |  |  |  | 11.1 | 12.6 |  |  |  | 12.103 | 13.6 |  |  |  | 14.95 | 16.48 |  |  |  | 38.1 | 38.9667 |  |  |  | 62 | 62.275
18.2 |  |  |  | 11 | 12.5 |  |  |  | 11.903 | 13.475 |  |  |  | 14.8 | 16.3 |  |  |  | 33.42 | 34.7 |  |  |  | 61.58 | 61.9333
18.3 |  |  |  | 11 | 12.4 |  |  |  | 11.803 | 13.3 |  |  |  | 14.68 | 16.1 |  |  |  | 29.38 | 30.4833 |  |  |  | 61.32 | 61.6
18.4 |  |  |  | 10.9 | 12.4 |  |  |  | 11.7029 | 13.2 |  |  |  | 14.35 | 15.9 |  |  |  | 26.08 | 27.3667 |  |  |  | 60.9 | 61.28
18.5 |  |  |  | 10.8 | 12.3 |  |  |  | 11.6029 | 13.1 |  |  |  | 14.225 | 15.7167 |  |  |  | 23.9667 | 25.2833 |  |  |  | 60.45 | 61
18.6 |  |  |  | 10.725 | 12.3 |  |  |  | 11.5029 | 13 |  |  |  | 14 | 15.525 |  |  |  | 22.5 | 23.9333 |  |  |  | 60.025 | 60.325
18.7 |  |  |  | 10.7 | 12.2 |  |  |  | 11.4029 | 12.9 |  |  |  | 13.92 | 15.3 |  |  |  | 21.5857 | 23.04 |  |  |  | 59.5 | 59.9
18.8 |  |  |  | 10.64 | 12.1 |  |  |  | 11.3028 | 12.86 |  |  |  | 13.725 | 15.2167 |  |  |  | 20.9 | 22.3333 |  |  |  | 58.9 | 59.2
18.9 |  |  |  | 10.6 | 12 |  |  |  | 11.2028 | 12.8 |  |  |  | 13.58 | 15.04 |  |  |  | 20.34 | 21.7 |  |  |  | 57.9 | 58.38
19 |  |  |  | 10.5 | 11.9 |  |  |  | 11.1428 | 12.6333 |  |  |  | 13.4 | 14.84 |  |  |  | 19.9286 | 21.3333 |  |  |  | 56.8 | 57.4
19.1 |  |  |  | 10.4 | 11.8333 |  |  |  | 11.0228 | 12.5 |  |  |  | 13.3 | 14.65 |  |  |  | 19.5667 | 20.92 |  |  |  | 55.35 | 55.8667
19.2 |  |  |  | 10.3 | 11.775 |  |  |  | 10.9628 | 12.44 |  |  |  | 13.08 | 14.58 |  |  |  | 19.2 | 20.66 |  |  |  | 53.2 | 53.7
19.3 |  |  |  | 10.225 | 11.7 |  |  |  | 10.8027 | 12.4 |  |  |  | 12.9 | 14.4 |  |  |  | 18.86 | 20.4 |  |  |  | 50.125 | 50.95
19.4 |  |  |  | 10.2 | 11.6 |  |  |  | 10.8027 | 12.3 |  |  |  | 12.8 | 14.3 |  |  |  | 18.66 | 20.1333 |  |  |  | 46.525 | 47.48
19.5 |  |  |  | 10.1 | 11.5 |  |  |  | 10.7027 | 12.2 |  |  |  | 12.675 | 14.1333 |  |  |  | 18.4143 | 19.78 |  |  |  | 42.64 | 43.35
19.6 |  |  |  | 10 | 11.4 |  |  |  | 10.7027 | 12.1 |  |  |  | 12.55 | 14 |  |  |  | 18.18 | 19.56 |  |  |  | 38.475 | 39.82
19.7 |  |  |  | 10 | 11.4 |  |  |  | 10.5827 | 12 |  |  |  | 12.425 | 13.8833 |  |  |  | 17.98 | 19.375 |  |  |  | 34.9 | 36.3
19.8 |  |  |  | 9.9 | 11.275 |  |  |  | 10.5026 | 11.9 |  |  |  | 12.3 | 13.7 |  |  |  | 17.6667 | 19.0833 |  |  |  | 31.5 | 32.45
19.9 |  |  |  | 9.775 | 11.2 |  |  |  | 10.4026 | 11.8333 |  |  |  | 12.2 | 13.6833 |  |  |  | 17.4833 | 18.9 |  |  |  | 28.8 | 30.0333
20 |  |  |  | 9.7 | 11.1 |  |  |  | 10.3026 | 11.75 |  |  |  | 12.1 | 13.5 |  |  |  | 17.3 | 18.6333 |  |  |  | 26.7 | 28.1
20.1 |  |  |  | 9.6 | 11 |  |  |  | 10.2026 | 11.7 |  |  |  | 12 | 13.5 |  |  |  | 16.98 | 18.55 |  |  |  | 25.15 | 26.6
20.2 | [8 empty] | 10.1425 | 11.6 |  |  |  | 11.85 | 13.32 |  |  |  | 16.8 | 18.4 |  |  |  | 24.2 | 25.65
20.3 | [8 empty] | 10.0275 | 11.5 |  |  |  | 11.8 | 13.2 |  |  |  | 16.5333 | 18.1167 |  |  |  | 23.6 | 25.1
20.4 | [8 empty] | 9.9775 | 11.4 |  |  |  | 11.7 | 13.1 |  |  |  | 16.4 | 17.92 |  |  |  | 22.9 | 24.55
20.5 | [8 empty] | 9.9025 | 11.3 |  |  |  | 11.6 | 13 |  |  |  | 16.26 | 17.7333 |  |  |  | 22.45 | 24.15
20.6 | [8 empty] | 9.8025 | 11.2 |  |  |  | 11.4333 | 12.84 |  |  |  | 16.08 | 17.45 |  |  |  | 22.1 | 23.7
20.7 | [8 empty] | 9.7224 | 11.1 |  |  |  | 11.4 | 12.7 |  |  |  | 15.88 | 17.3 |  |  |  | 21.8 | 23.325
20.8 | [13 empty] | 11.3 | 12.66 |  |  |  | 15.7 | 17.16 |  |  |  | 21.45 | 23
20.9 | [13 empty] | 11.2 | 12.4857 |  |  |  | 15.4857 | 16.9 |  |  |  | 21.2 | 22.7
21 | [13 empty] | 11.075 | 12.4 |  |  |  | 15.32 | 16.8333 |  |  |  | 20.98 | 22.525
21.1 | [13 empty] | 11 | 12.3 |  |  |  | 15.1333 | 16.6667 |  |  |  | 20.74 | 22.4
21.2 | [13 empty] | 10.88 | 12.2 |  |  |  | 14.9714 | 16.48 |  |  |  | 20.425 | 22.1
21.3 | [13 empty] | 10.8 | 12.1 |  |  |  | 14.8 | 16.35 |  |  |  | 20.22 | 21.95
21.4 | [13 empty] | 10.7 | 12.0167 |  |  |  | 14.7 | 16.2 |  |  |  | 19.975 | 21.7
21.5 | [18 empty] | 14.5 | 16.0667 |  |  |  | 19.74 | 21.3833
21.6 | [18 empty] | 14.4 | 15.9 |  |  |  | 19.48 | 21.2
21.7 | [18 empty] | 14.3 | 15.7286 |  |  |  | 19.3 | 20.95
21.8 | [18 empty] | 14.12 | 15.55 |  |  |  | 19.2 | 20.7333
21.9 | [18 empty] | 14 | 15.4 |  |  |  | 18.95 | 20.5
22 | [18 empty] | 13.8 | 15.3167 |  |  |  | 18.72 | 20.36
22.1 | [18 empty] | 13.7 | 15.125 |  |  |  | 18.5 | 20.1333
22.2 | [18 empty] | 13.6 | 15.02 |  |  |  | 18.38 | 19.925
22.3 | [18 empty] | 13.5167 | 14.8375 |  |  |  | 18.25 | 19.8
22.4 | [18 empty] | 13.42 | 14.7 |  |  |  | 17.98 | 19.58
22.5 | [18 empty] | 13.25 | 14.58 |  |  |  | 17.775 | 19.4
22.6 | [18 empty] | 13.08 | 14.5 |  |  |  | 17.575 | 19.25
22.7 | [18 empty] | 12.9 | 14.3571 |  |  |  | 17.425 | 19.125
22.8 | [18 empty] | 12.82 | 14.28 |  |  |  | 17.275 | 18.875
22.9 | [18 empty] | 12.74 | 14.0833 |  |  |  | 17.12 | 18.8

## Wedge 30 Longitudinal Profiles
Longitudinal Profiles
Wedge 30
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 18.2
-30.8 |  |  |  |  | 18.3444
-30.7 |  |  |  |  | 18.5375
-30.6 |  |  |  |  | 18.6625
-30.5 |  |  |  |  | 18.7875
-30.4 |  |  |  |  | 18.9125
-30.3 |  |  |  |  | 19.075
-30.2 |  |  |  |  | 19.325
-30.1 |  |  |  |  | 19.4875
-30 |  |  |  |  | 19.625
-29.9 |  |  |  |  | 19.875
-29.8 |  |  |  |  | 20.0625
-29.7 |  |  |  |  | 20.1875
-29.6 |  |  |  |  | 20.45
-29.5 |  |  |  |  | 20.6571
-29.4 |  |  |  |  | 20.7667
-29.3 |  |  |  |  | 21
-29.2 |  |  |  |  | 21.2
-29.1 |  |  |  |  | 21.3375
-29 |  |  |  |  | 21.675
-28.9 |  |  |  | 15.4 | 21.8625
-28.8 |  |  |  | 15.45 | 22.1
-28.7 |  |  |  | 15.6 | 22.2
-28.6 |  |  |  | 15.7571 | 22.5
-28.5 |  |  |  | 15.975 | 22.6875
-28.4 |  |  |  | 16.2125 | 22.9
-28.3 |  |  |  | 16.3375 | 23.15
-28.2 |  |  |  | 16.4625 | 23.4
-28.1 |  |  |  | 16.5875 | 23.65
-28 |  |  |  | 16.8111 | 23.9
-27.9 |  |  |  | 16.925 | 24.125
-27.8 |  |  |  | 17.1143 | 24.4125
-27.7 |  |  |  | 17.3 | 24.725
-27.6 |  |  |  | 17.45 | 24.975
-27.5 |  |  |  | 17.65 | 25.2375
-27.4 |  |  |  | 17.8 | 25.6
-27.3 |  |  |  | 18.0333 | 25.9
-27.2 |  |  |  | 18.2 | 26.225
-27.1 |  |  |  | 18.4375 | 26.6857
-27 |  |  |  | 18.5833 | 27.1429
-26.9 |  |  | 12.6333 | 18.85 | 27.6
-26.8 |  |  | 12.8 | 19.075 | 28.125
-26.7 |  |  | 12.9 | 19.3286 | 28.9
-26.6 |  |  | 12.98 | 19.5667 | 30.05
-26.5 |  |  | 13.1 | 19.8143 | 31.6
-26.4 |  |  | 13.2 | 20 | 33.325
-26.3 |  |  | 13.3667 | 20.25 | 35.7625
-26.2 |  |  | 13.44 | 20.5333 | 38.825
-26.1 |  |  | 13.6 | 20.7143 | 42.3375
-26 |  |  | 13.78 | 20.9 | 46.425
-25.9 |  | 11.3 | 13.9 | 21.2 | 50.7875
-25.8 |  | 11.425 | 13.9833 | 21.4857 | 55.1625
-25.7 |  | 11.5 | 14.1667 | 21.8167 | 59.2
-25.6 |  | 11.575 | 14.325 | 22.0875 | 62.7
-25.5 |  | 11.6167 | 14.45 | 22.425 | 65.625
-25.4 |  | 11.8 | 14.5857 | 22.7125 | 68
-25.3 |  | 11.86 | 14.7333 | 23.025 | 69.7429
-25.2 | 10.8286 | 11.95 | 14.9 | 23.3571 | 70.9
-25.1 | 10.9 | 12.0143 | 15.0429 | 23.8 | 72.1
-25 | 10.95 | 12.1571 | 15.3 | 24.2 | 72.8429
-24.9 | 11.0333 | 12.2167 | 15.4143 | 24.8143 | 73.5429
-24.8 | 11.1167 | 12.3625 | 15.6 | 25.4 | 73.9333
-24.7 | 11.2 | 12.5 | 15.8143 | 26.4 | 74.3625
-24.6 | 11.3 | 12.5857 | 16 | 27.5333 | 74.9125
-24.5 | 11.4 | 12.65 | 16.26 | 29.2 | 75.2667
-24.4 | 11.4833 | 12.8 | 16.4333 | 31.75 | 75.7
-24.3 | 11.5667 | 12.95 | 16.5857 | 34.6857 | 76
-24.2 | 11.6333 | 13.0167 | 16.88 | 38.5714 | 76.3571
-24.1 | 11.7 | 13.1 | 17.04 | 43 | 76.6667
-24 | 11.8 | 13.2857 | 17.25 | 48 | 76.95
-23.9 | 11.9 | 13.4 | 17.5143 | 53.2667 | 77.3
-23.8 | 11.9571 | 13.5667 | 17.8167 | 58.6833 | 77.55
-23.7 | 12 | 13.7333 | 18.0429 | 63.14 | 77.7667
-23.6 | 12.1571 | 13.8857 | 18.2571 | 66.9 | 78.0333
-23.5 | 12.225 | 14.0167 | 18.6 | 70.36 | 78.3875
-23.4 | 12.35 | 14.1714 | 18.86 | 72.7857 | 78.675
-23.3 | 12.4 | 14.4143 | 19.18 | 74.55 | 78.9
-23.2 | 12.5714 | 14.5571 | 19.5857 | 75.95 | 79.1429
-23.1 | 12.6429 | 14.7333 | 19.9667 | 76.9 | 79.425
-23 | 12.8 | 14.9714 | 20.44 | 77.6429 | 79.8
-22.9 | 12.8833 | 15.1 | 20.9857 | 78.2857 | 80.05
-22.8 | 13 | 15.3667 | 21.6167 | 78.7333 | 80.3
-22.7 | 13.1286 | 15.6 | 22.6125 | 79.2125 | 80.5
-22.6 | 13.2625 | 15.7667 | 23.8714 | 79.6625 | 80.7333
-22.5 | 13.4333 | 16 | 25.6143 | 80.0375 | 80.975
-22.4 | 13.6 | 16.2667 | 28.3667 | 80.375 | 81.225
-22.3 | 13.6714 | 16.5714 | 32.1286 | 80.6875 | 81.4375
-22.2 | 13.8286 | 16.8333 | 37.1375 | 80.975 | 81.5833
-22.1 | 14 | 17.1667 | 42.7 | 81.2 | 81.825
-22 | 14.18 | 17.6143 | 49.0333 | 81.6 | 82.0857
-21.9 | 14.38 | 18.15 | 55.95 | 81.8857 | 82.4571
-21.8 | 14.5333 | 18.78 | 62.675 | 82.15 | 82.6286
-21.7 | 14.7167 | 19.8333 | 68.5 | 82.4667 | 82.7833
-21.6 | 14.925 | 21.0833 | 73.65 | 82.8 | 83
-21.5 | 15.175 | 23.5143 | 77.3143 | 83.1 | 83.1429
-21.4 | 15.5286 | 26.5857 | 80.1 | 83.2857 | 83.3714
-21.3 | 15.7667 | 31.3 | 81.8571 | 83.58 | 83.6333
-21.2 | 16.24 | 36.5 | 82.9 | 83.825 | 83.7857
-21.1 | 16.76 | 43.5571 | 83.8 | 84.0333 | 84.1
-21 | 17.6125 | 50.6 | 84.5429 | 84.3286 | 84.3167
-20.9 | 19.0333 | 58.6 | 85.1286 | 84.5571 | 84.4833
-20.8 | 21.3 | 66.3333 | 85.6429 | 84.7 | 84.7571
-20.7 | 24.7286 | 73 | 86.1 | 84.975 | 84.9625
-20.6 | 29.6 | 78.65 | 86.46 | 85.1714 | 85.175
-20.5 | 36.3333 | 82.575 | 86.8714 | 85.3 | 85.3111
-20.4 | 43.5286 | 85.0667 | 87.15 | 85.55 | 85.45
-20.3 | 51.9 | 86.7143 | 87.4333 | 85.85 | 85.65
-20.2 | 60.35 | 87.9 | 87.7167 | 86.15 | 85.775
-20.1 | 68.7 | 88.7 | 87.9667 | 86.4 | 86
-20 | 76.1167 | 89.3 | 88.1857 | 86.55 | 86.1833
-19.9 | 82.2833 | 89.7857 | 88.4714 | 86.725 | 86.3571
-19.8 | 85.9667 | 90.2 | 88.7333 | 86.9375 | 86.5
-19.7 | 88.6 | 90.4833 | 88.9833 | 87.0625 | 86.625
-19.6 | 90.0857 | 90.7857 | 89.1667 | 87.3 | 86.7571
-19.5 | 91.2 | 91.05 | 89.4 | 87.48 | 86.8875
-19.4 | 91.92 | 91.3333 | 89.5667 | 87.6667 | 87.1143
-19.3 | 92.4 | 91.5714 | 89.8 | 87.8571 | 87.4
-19.2 | 92.7 | 91.7667 | 89.9333 | 88.125 | 87.5375
-19.1 | 92.92 | 92.0143 | 90.1625 | 88.2875 | 87.6833
-19 | 93.2 | 92.18 | 90.3286 | 88.45 | 87.8333
-18.9 | 93.38 | 92.35 | 90.5667 | 88.6 | 87.9167
-18.8 | 93.55 | 92.5167 | 90.76 | 88.7286 | 88.0714
-18.7 | 93.6286 | 92.6714 | 90.9 | 88.9889 | 88.2
-18.6 | 93.8 | 92.8 | 91.1286 | 89.1625 | 88.3667
-18.5 | 93.9714 | 92.95 | 91.2833 | 89.3 | 88.625
-18.4 | 94.1 | 93.1 | 91.5 | 89.4429 | 88.78
-18.3 | 94.1714 | 93.2 | 91.7 | 89.6 | 88.9
-18.2 | 94.26 | 93.36 | 91.7571 | 89.75 | 88.975
-18.1 | 94.4333 | 93.5 | 91.8875 | 89.9143 | 89.1625
-18 | 94.6 | 93.76 | 92 | 90.0571 | 89.3
-17.9 | 94.7 | 93.9 | 92.1167 | 90.1778 | 89.3833
-17.8 | 94.7857 | 93.9571 | 92.2714 | 90.32 | 89.5444
-17.7 | 94.9286 | 94.02 | 92.4143 | 90.5 | 89.7
-17.6 | 95 | 94.1857 | 92.5667 | 90.6 | 89.7625
-17.5 | 95.1167 | 94.25 | 92.7333 | 90.7 | 89.975
-17.4 | 95.2 | 94.3875 | 92.9167 | 90.8333 | 90.1125
-17.3 | 95.3 | 94.4667 | 93 | 91.1 | 90.25
-17.2 | 95.4429 | 94.6143 | 93.125 | 91.175 | 90.4
-17.1 | 95.5857 | 94.7667 | 93.2667 | 91.3625 | 90.5143
-17 | 95.6 | 94.9 | 93.4333 | 91.4 | 90.6
-16.9 | 95.76 | 94.9667 | 93.6 | 91.5571 | 90.8
-16.8 | 95.8333 | 95.025 | 93.6571 | 91.7 | 90.925
-16.7 | 95.9 | 95.1 | 93.7 | 91.8 | 91.05
-16.6 | 96 | 95.22 | 93.8 | 91.8833 | 91.1667
-16.5 | 96.04 | 95.3857 | 93.9 | 92.0143 | 91.2875
-16.4 | 96.1125 | 95.4667 | 94 | 92.1571 | 91.4125
-16.3 | 96.2429 | 95.6 | 94.1667 | 92.2778 | 91.5375
-16.2 | 96.3 | 95.6 | 94.325 | 92.4 | 91.6625
-16.1 | 96.4 | 95.7 | 94.45 | 92.6 | 91.7875
-16 | 96.4833 | 95.8 | 94.5 | 92.6333 | 91.9111
-15.9 | 96.5429 | 95.8667 | 94.6 | 92.8 | 92.025
-15.8 | 96.6 | 96 | 94.675 | 92.8571 | 92.1
-15.7 | 96.7 | 96.1 | 94.9 | 93.075 | 92.2
-15.6 | 96.7 | 96.1 | 94.975 | 93.2 | 92.3714
-15.5 | 96.72 | 96.2 | 95.0125 | 93.25 | 92.5
-15.4 | 96.9 | 96.25 | 95.1 | 93.3875 | 92.6143
-15.3 | 97 | 96.3 | 95.1 | 93.4125 | 92.7667
-15.2 | 97.1 | 96.4167 | 95.2 | 93.575 | 92.8
-15.1 | 97.1 | 96.5 | 95.2833 | 93.7625 | 92.8714
-15 | 97.2 | 96.58 | 95.4167 | 93.82 | 92.9889
-14.9 | 97.3 | 96.6167 | 95.5 | 94 | 93.1286
-14.8 | 97.3 | 96.7833 | 95.6 | 94.1 | 93.2714
-14.7 | 97.4 | 96.8 | 95.7 | 94.1 | 93.34
-14.6 | 97.4 | 96.8857 | 95.8 | 94.2 | 93.4
-14.5 | 97.4286 | 96.9429 | 95.8857 | 94.3 | 93.6
-14.4 | 97.5 | 97 | 95.95 | 94.3833 | 93.6857
-14.3 | 97.5 | 97.0429 | 96.1 | 94.5429 | 93.8
-14.2 | 97.6 | 97.1 | 96.1 | 94.675 | 93.8
-14.1 | 97.6 | 97.16 | 96.1857 | 94.7222 | 93.9667
-14 | 97.7 | 97.2 | 96.3 | 94.8 | 94.0167
-13.9 | 97.8 | 97.2833 | 96.3667 | 94.8857 | 94.1714
-13.8 | 97.9 | 97.35 | 96.4286 | 94.95 | 94.3
-13.7 | 97.9 | 97.4875 | 96.5 | 95.0875 | 94.3333
-13.6 | 98 | 97.5 | 96.6 | 95.2 | 94.475
-13.5 | 98 | 97.6 | 96.6 | 95.2625 | 94.6
-13.4 | 98.1 | 97.6 | 96.7 | 95.3 | 94.6625
-13.3 | 98.1 | 97.7 | 96.8 | 95.4714 | 94.8
-13.2 | 98.2 | 97.7167 | 96.9 | 95.5 | 94.8
-13.1 | 98.3 | 97.8 | 96.95 | 95.55 | 94.95
-13 | 98.3 | 97.8 | 97.0167 | 95.7 | 95
-12.9 | 98.3 | 97.9 | 97.1 | 95.78 | 95.1
-12.8 | 98.4 | 97.9429 | 97.2 | 95.9 | 95.25
-12.7 | 98.4143 | 98 | 97.3 | 95.975 | 95.4
-12.6 | 98.5 | 98.1 | 97.3667 | 96.0125 | 95.4833
-12.5 | 98.6 | 98.1857 | 97.4 | 96.15 | 95.5
-12.4 | 98.6 | 98.2125 | 97.5 | 96.2 | 95.5714
-12.3 | 98.6 | 98.3 | 97.5 | 96.3 | 95.7
-12.2 | 98.7 | 98.3 | 97.6 | 96.3667 | 95.7571
-12.1 | 98.8 | 98.4 | 97.6 | 96.4 | 95.8125
-12 | 98.8429 | 98.4 | 97.7429 | 96.5 | 95.9
-11.9 | 98.9 | 98.4857 | 97.8 | 96.575 | 95.9
-11.8 | 98.9 | 98.5 | 97.825 | 96.6 | 96
-11.7 | 98.9333 | 98.6 | 97.9 | 96.6667 | 96.0875
-11.6 | 99 | 98.7 | 98 | 96.725 | 96.2
-11.5 | 99 | 98.78 | 98 | 96.8 | 96.225
-11.4 | 99.0571 | 98.8 | 98.0143 | 96.9 | 96.35
-11.3 | 99.1 | 98.8 | 98.1 | 97.025 | 96.475
-11.2 | 99.18 | 98.8 | 98.1167 | 97.1 | 96.5
-11.1 | 99.2 | 98.9 | 98.2 | 97.1714 | 96.55
-11 | 99.3 | 98.9 | 98.2571 | 97.3 | 96.7
-10.9 | 99.4 | 98.9 | 98.3 | 97.3714 | 96.7
-10.8 | 99.4 | 99 | 98.4 | 97.4375 | 96.775
-10.7 | 99.4 | 99 | 98.4 | 97.5556 | 96.8
-10.6 | 99.4571 | 99.0667 | 98.5 | 97.6 | 96.9375
-10.5 | 99.5 | 99.1 | 98.5714 | 97.7 | 97
-10.4 | 99.5 | 99.2 | 98.6 | 97.7571 | 97.1
-10.3 | 99.56 | 99.2333 | 98.6429 | 97.825 | 97.18
-10.2 | 99.6 | 99.3 | 98.7 | 97.9 | 97.2286
-10.1 | 99.6 | 99.3 | 98.7 | 98 | 97.3
-10 | 99.7 | 99.4 | 98.7833 | 98 | 97.4111
-9.9 | 99.7333 | 99.4429 | 98.9 | 98.0714 | 97.5286
-9.8 | 99.8 | 99.5 | 98.9 | 98.1 | 97.6
-9.7 | 99.8 | 99.5 | 99 | 98.1333 | 97.7
-9.6 | 99.8 | 99.6 | 99 | 98.2 | 97.725
-9.5 | 99.8429 | 99.6 | 99.0429 | 98.3 | 97.85
-9.4 | 99.9 | 99.7 | 99.2 | 98.3167 | 97.9
-9.3 | 99.9 | 99.7 | 99.2 | 98.4 | 97.9875
-9.2 | 100 | 99.7 | 99.2429 | 98.5 | 98.0286
-9.1 | 100 | 99.8 | 99.3 | 98.5 | 98.1
-9 | 100 | 99.8 | 99.3 | 98.6 | 98.1
-8.9 | 100 | 99.8 | 99.4 | 98.6 | 98.1875
-8.8 | 100.1 | 99.8 | 99.4 | 98.6375 | 98.2
-8.7 | 100.1167 | 99.8667 | 99.5 | 98.7 | 98.2
-8.6 | 100.2 | 99.94 | 99.5 | 98.7 | 98.225
-8.5 | 100.2 | 100 | 99.5 | 98.7625 | 98.3
-8.4 | 100.2833 | 100 | 99.5167 | 98.8 | 98.4
-8.3 | 100.3 | 100.1 | 99.6 | 98.9 | 98.4
-8.2 | 100.3 | 100.1 | 99.6429 | 98.9 | 98.5
-8.1 | 100.3 | 100.1 | 99.7 | 99.025 | 98.5286
-8 | 100.4 | 100.1571 | 99.7 | 99.1 | 98.6
-7.9 | 100.4 | 100.2 | 99.7667 | 99.1 | 98.65
-7.8 | 100.4 | 100.2 | 99.8 | 99.15 | 98.7875
-7.7 | 100.4 | 100.2 | 99.8 | 99.2 | 98.8167
-7.6 | 100.4 | 100.2857 | 99.85 | 99.3 | 98.9
-7.5 | 100.5 | 100.3286 | 99.9 | 99.3 | 98.9286
-7.4 | 100.5 | 100.4 | 99.9 | 99.3333 | 99
-7.3 | 100.6 | 100.4667 | 99.9 | 99.4 | 99.0333
-7.2 | 100.6 | 100.5 | 99.9833 | 99.4125 | 99.1
-7.1 | 100.6 | 100.5 | 100 | 99.5 | 99.2
-7 | 100.6 | 100.5 | 100 | 99.5286 | 99.2
-6.9 | 100.6 | 100.5 | 100.1 | 99.6 | 99.2
-6.8 | 100.7 | 100.5 | 100.1571 | 99.6 | 99.2
-6.7 | 100.7 | 100.5 | 100.2 | 99.6571 | 99.2
-6.6 | 100.7 | 100.5875 | 100.2 | 99.7 | 99.24
-6.5 | 100.7 | 100.6 | 100.2 | 99.7 | 99.3
-6.4 | 100.8 | 100.6 | 100.2 | 99.7 | 99.3
-6.3 | 100.8 | 100.6 | 100.25 | 99.7 | 99.3
-6.2 | 100.8 | 100.6 | 100.2 | 99.8 | 99.3625
-6.1 | 100.8 | 100.6 | 100.2714 | 99.8 | 99.4
-6 | 100.8 | 100.6 | 100.3 | 99.8 | 99.4
-5.9 | 100.9 | 100.7 | 100.32 | 99.85 | 99.4
-5.8 | 100.9 | 100.7 | 100.4 | 99.9 | 99.5
-5.7 | 100.8333 | 100.7 | 100.4 | 99.9 | 99.5
-5.6 | 100.86 | 100.7 | 100.4 | 99.9667 | 99.5
-5.5 | 100.9 | 100.7 | 100.4 | 99.9714 | 99.6
-5.4 | 100.9 | 100.7667 | 100.4 | 99.9 | 99.6
-5.3 | 100.9 | 100.8 | 100.45 | 99.9 | 99.6
-5.2 | 100.9 | 100.8 | 100.4 | 99.9 | 99.6
-5.1 | 100.9 | 100.8 | 100.4 | 99.9 | 99.6
-5 | 100.9 | 100.7 | 100.4 | 99.9 | 99.6
-4.9 | 100.9 | 100.7 | 100.4 | 99.9 | 99.5429
-4.8 | 100.8 | 100.7 | 100.3 | 99.9 | 99.6
-4.7 | 100.8 | 100.6143 | 100.3 | 99.9 | 99.6
-4.6 | 100.72 | 100.6 | 100.3 | 99.9 | 99.6
-4.5 | 100.7 | 100.6 | 100.3 | 99.9 | 99.65
-4.4 | 100.7 | 100.6 | 100.3 | 99.9 | 99.7
-4.3 | 100.7 | 100.6 | 100.3 | 99.9 | 99.7
-4.2 | 100.6 | 100.6 | 100.22 | 99.9 | 99.7
-4.1 | 100.6 | 100.5 | 100.2 | 100 | 99.7167
-4 | 100.6 | 100.5 | 100.2 | 100 | 99.8
-3.9 | 100.5 | 100.4333 | 100.2143 | 100 | 99.8
-3.8 | 100.5 | 100.4 | 100.3 | 100 | 99.8
-3.7 | 100.5 | 100.4 | 100.3 | 100 | 99.9
-3.6 | 100.4 | 100.4 | 100.3 | 100 | 99.9
-3.5 | 100.4 | 100.475 | 100.3 | 100 | 100
-3.4 | 100.4 | 100.4 | 100.3 | 100.1 | 100
-3.3 | 100.4 | 100.4 | 100.3 | 100.1 | 100
-3.2 | 100.4 | 100.3571 | 100.3 | 100.1 | 100
-3.1 | 100.4 | 100.3 | 100.3 | 100.2 | 100
-3 | 100.4 | 100.3 | 100.3 | 100.2 | 100
-2.9 | 100.4 | 100.4 | 100.3 | 100.2 | 100
-2.8 | 100.4 | 100.4 | 100.3429 | 100.2 | 100
-2.7 | 100.4 | 100.4 | 100.4 | 100.2 | 100
-2.6 | 100.4 | 100.4 | 100.4 | 100.1429 | 100
-2.5 | 100.4 | 100.4 | 100.3 | 100.1875 | 100
-2.4 | 100.4 | 100.4 | 100.3 | 100.2 | 100.1
-2.3 | 100.4 | 100.4 | 100.3 | 100.2 | 100.0286
-2.2 | 100.4 | 100.4 | 100.3 | 100.2 | 100
-2.1 | 100.4 | 100.4 | 100.3 | 100.2 | 100.0375
-2 | 100.4 | 100.4 | 100.3 | 100.2 | 100.1
-1.9 | 100.4 | 100.4 | 100.3 | 100.1167 | 100.1
-1.8 | 100.3167 | 100.4 | 100.3 | 100.1 | 100.0667
-1.7 | 100.3 | 100.4 | 100.3 | 100.1 | 100
-1.6 | 100.3 | 100.4 | 100.2 | 100.1 | 100
-1.5 | 100.3 | 100.3429 | 100.2 | 100.1 | 99.9143
-1.4 | 100.3 | 100.3 | 100.2 | 100.1 | 99.9286
-1.3 | 100.3 | 100.3 | 100.2 | 100.0125 | 100
-1.2 | 100.2167 | 100.3 | 100.2 | 100 | 100
-1.1 | 100.2 | 100.2 | 100.2 | 100 | 100.0714
-1 | 100.2 | 100.2 | 100.1 | 100 | 100.1
-0.9 | 100.2 | 100.1875 | 100.1 | 100 | 100.06
-0.8 | 100.1 | 100.1 | 100.1 | 100 | 100
-0.7 | 100.05 | 100.1 | 100.1 | 100 | 99.98
-0.6 | 100.1 | 100.1 | 100.1 | 100 | 99.9
-0.5 | 100 | 100.1 | 100.05 | 100 | 99.9167
-0.4 | 100 | 100 | 100 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100 | 100
0.4 | 100 | 100 | 100 | 100 | 100
0.5 | 100 | 100.1 | 100.0833 | 100 | 99.9
0.6 | 100.1 | 100.1 | 100.1 | 100 | 99.9
0.7 | 100.1 | 100.1 | 100.1 | 100 | 99.95
0.8 | 100.1 | 100.1 | 100.1 | 100 | 100
0.9 | 100.18 | 100.1875 | 100.1 | 100 | 100.0333
1 | 100.2 | 100.2 | 100.1 | 100 | 100.1
1.1 | 100.2 | 100.2 | 100.2 | 100 | 100.0875
1.2 | 100.2 | 100.3 | 100.2 | 100 | 100
1.3 | 100.3 | 100.3 | 100.2 | 100.0625 | 100
1.4 | 100.3 | 100.3 | 100.2 | 100.1 | 99.9571
1.5 | 100.3 | 100.3429 | 100.2 | 100.1 | 99.9
1.6 | 100.3 | 100.4 | 100.2 | 100.1 | 100
1.7 | 100.3 | 100.4 | 100.3 | 100.1 | 100
1.8 | 100.3 | 100.4 | 100.3 | 100.1 | 100.1
1.9 | 100.4 | 100.4 | 100.3 | 100.1 | 100.1
2 | 100.4 | 100.4 | 100.3 | 100.2 | 100.1
2.1 | 100.4 | 100.4 | 100.3 | 100.2 | 100.0625
2.2 | 100.4 | 100.4 | 100.3 | 100.2 | 100
2.3 | 100.4 | 100.4 | 100.3 | 100.2 | 100
2.4 | 100.4 | 100.4 | 100.3 | 100.2 | 100.1
2.5 | 100.4 | 100.4 | 100.3 | 100.2 | 100
2.6 | 100.4 | 100.4 | 100.3667 | 100.1125 | 100
2.7 | 100.4 | 100.4 | 100.4 | 100.2 | 100
2.8 | 100.4 | 100.4 | 100.4 | 100.2 | 100
2.9 | 100.4 | 100.4 | 100.3 | 100.2 | 100
3 | 100.4 | 100.3 | 100.3 | 100.2 | 100.0778
3.1 | 100.4 | 100.3 | 100.3 | 100.2 | 100
3.2 | 100.4 | 100.3571 | 100.3 | 100.125 | 100
3.3 | 100.4 | 100.4 | 100.3 | 100.1 | 100
3.4 | 100.4 | 100.4 | 100.3 | 100.05 | 100
3.5 | 100.4 | 100.475 | 100.3 | 100 | 100
3.6 | 100.4 | 100.4 | 100.3 | 100 | 99.9333
3.7 | 100.4833 | 100.4 | 100.3 | 100 | 99.9
3.8 | 100.5 | 100.4 | 100.3 | 100 | 99.825
3.9 | 100.5 | 100.4333 | 100.2 | 100 | 99.8
4 | 100.6 | 100.5 | 100.2 | 100 | 99.8
4.1 | 100.6 | 100.5 | 100.2 | 99.9833 | 99.7333
4.2 | 100.6 | 100.6 | 100.26 | 99.9 | 99.7
4.3 | 100.7 | 100.6 | 100.3 | 99.9 | 99.7
4.4 | 100.7 | 100.6 | 100.3 | 99.9 | 99.7
4.5 | 100.7 | 100.6 | 100.3 | 99.9 | 99.6833
4.6 | 100.7833 | 100.6 | 100.3 | 99.9 | 99.6
4.7 | 100.8 | 100.6143 | 100.3 | 99.9 | 99.6
4.8 | 100.8 | 100.7 | 100.3 | 99.9 | 99.6
4.9 | 100.85 | 100.7 | 100.4 | 99.9 | 99.5667
5 | 100.9 | 100.7 | 100.4 | 99.9 | 99.5857
5.1 | 100.9 | 100.8 | 100.4 | 99.9 | 99.6
5.2 | 100.9 | 100.8 | 100.4 | 99.9 | 99.6
5.3 | 100.9 | 100.8 | 100.4 | 99.9 | 99.6
5.4 | 100.9 | 100.7667 | 100.4 | 99.9 | 99.6
5.5 | 100.9 | 100.7 | 100.4 | 99.9167 | 99.6
5.6 | 100.8 | 100.7 | 100.4 | 100 | 99.5
5.7 | 100.8833 | 100.7 | 100.4 | 99.9 | 99.5
5.8 | 100.9 | 100.7 | 100.4 | 99.9 | 99.5
5.9 | 100.88 | 100.7 | 100.3 | 99.9 | 99.4
6 | 100.8 | 100.6 | 100.3 | 99.8 | 99.4
6.1 | 100.8 | 100.6 | 100.2429 | 99.8 | 99.4
6.2 | 100.8 | 100.6 | 100.2167 | 99.8 | 99.3875
6.3 | 100.8 | 100.6 | 100.3 | 99.7 | 99.3
6.4 | 100.7429 | 100.6 | 100.2 | 99.7 | 99.3
6.5 | 100.7 | 100.6 | 100.2 | 99.7 | 99.3
6.6 | 100.7 | 100.5875 | 100.2 | 99.7 | 99.2667
6.7 | 100.7 | 100.5 | 100.2 | 99.7 | 99.2
6.8 | 100.6333 | 100.5 | 100.1333 | 99.6 | 99.2
6.9 | 100.6 | 100.5 | 100.1 | 99.6 | 99.2
7 | 100.6 | 100.5 | 100.0429 | 99.5 | 99.2
7.1 | 100.6 | 100.5 | 100 | 99.5 | 99.2
7.2 | 100.6 | 100.5 | 99.95 | 99.4 | 99.1333
7.3 | 100.6 | 100.4667 | 99.9 | 99.4 | 99.0667
7.4 | 100.5 | 100.4 | 99.9 | 99.3833 | 99
7.5 | 100.5 | 100.3286 | 99.9 | 99.3 | 98.95
7.6 | 100.42 | 100.2857 | 99.9 | 99.2667 | 98.9
7.7 | 100.4 | 100.2 | 99.8 | 99.2 | 98.85
7.8 | 100.4 | 100.2 | 99.8 | 99.1 | 98.7222
7.9 | 100.4 | 100.2 | 99.8 | 99.1 | 98.68
8 | 100.4 | 100.1571 | 99.7 | 99.1 | 98.6
8.1 | 100.35 | 100.1 | 99.7 | 99 | 98.5571
8.2 | 100.3 | 100.1 | 99.6143 | 98.9375 | 98.5
8.3 | 100.3 | 100.1 | 99.6 | 98.9 | 98.4
8.4 | 100.3 | 100 | 99.5 | 98.8429 | 98.4
8.5 | 100.2 | 100 | 99.5 | 98.7125 | 98.3
8.6 | 100.2 | 99.94 | 99.5 | 98.7 | 98.25
8.7 | 100.1667 | 99.8667 | 99.4875 | 98.7 | 98.2
8.8 | 100.1 | 99.8 | 99.4 | 98.6 | 98.2
8.9 | 100.0333 | 99.8 | 99.4 | 98.6 | 98.2
9 | 100 | 99.8 | 99.3429 | 98.575 | 98.1
9.1 | 100 | 99.8 | 99.3 | 98.5 | 98.1
9.2 | 100 | 99.7 | 99.225 | 98.4667 | 98
9.3 | 99.9143 | 99.7 | 99.2 | 98.4 | 98
9.4 | 99.9 | 99.7 | 99.1833 | 98.3667 | 97.9
9.5 | 99.8 | 99.6 | 99.0167 | 98.3 | 97.875
9.6 | 99.8 | 99.6 | 99 | 98.2375 | 97.75
9.7 | 99.8 | 99.5 | 99 | 98.1857 | 97.7
9.8 | 99.8 | 99.5 | 98.9 | 98.1 | 97.6
9.9 | 99.7833 | 99.4429 | 98.9 | 98.0222 | 97.5571
10 | 99.7 | 99.4 | 98.8 | 98 | 97.4333
10.1 | 99.6 | 99.3 | 98.7 | 97.9375 | 97.3
10.2 | 99.6 | 99.3 | 98.7 | 97.9 | 97.2571
10.3 | 99.6 | 99.2333 | 98.6143 | 97.8 | 97.2
10.4 | 99.5 | 99.2 | 98.6 | 97.7 | 97.1
10.5 | 99.5 | 99.1 | 98.6 | 97.6667 | 96.9875
10.6 | 99.4125 | 99.0667 | 98.5 | 97.6 | 96.9
10.7 | 99.4 | 99 | 98.4 | 97.5 | 96.8286
10.8 | 99.4 | 99 | 98.3875 | 97.4 | 96.8
10.9 | 99.3571 | 98.9 | 98.3 | 97.4 | 96.7
11 | 99.3 | 98.9 | 98.2286 | 97.2714 | 96.7
11.1 | 99.2 | 98.9 | 98.2 | 97.2 | 96.575
11.2 | 99.12 | 98.8 | 98.2 | 97.1 | 96.5
11.3 | 99.1 | 98.8 | 98.1 | 96.975 | 96.4
11.4 | 99 | 98.8 | 98 | 96.9 | 96.375
11.5 | 99 | 98.78 | 98 | 96.8 | 96.25
11.6 | 99 | 98.7 | 97.9833 | 96.7714 | 96.2
11.7 | 98.9833 | 98.6 | 97.9 | 96.7 | 96.1111
11.8 | 98.9 | 98.5 | 97.9 | 96.6 | 95.9833
11.9 | 98.9 | 98.4857 | 97.8 | 96.6 | 95.9
12 | 98.8 | 98.4 | 97.7143 | 96.5 | 95.9
12.1 | 98.8 | 98.4 | 97.6333 | 96.4571 | 95.8286
12.2 | 98.7 | 98.3 | 97.5857 | 96.3143 | 95.7833
12.3 | 98.6167 | 98.3 | 97.5 | 96.3 | 95.7
12.4 | 98.6 | 98.2125 | 97.5 | 96.2 | 95.6
12.5 | 98.58 | 98.1857 | 97.4 | 96.1 | 95.5
12.6 | 98.5 | 98.1 | 97.4 | 96.0571 | 95.5
12.7 | 98.4 | 98 | 97.2714 | 96 | 95.34
12.8 | 98.4 | 97.9429 | 97.2 | 95.8857 | 95.2714
12.9 | 98.3 | 97.9 | 97.1 | 95.8 | 95.1
13 | 98.3 | 97.8 | 97.0833 | 95.65 | 95.0286
13.1 | 98.25 | 97.8 | 97 | 95.6 | 94.9667
13.2 | 98.2 | 97.7167 | 96.9 | 95.5 | 94.8
13.3 | 98.1167 | 97.7 | 96.8429 | 95.4286 | 94.8
13.4 | 98.1 | 97.6 | 96.7 | 95.36 | 94.6875
13.5 | 98.0167 | 97.6 | 96.6 | 95.3 | 94.6
13.6 | 98 | 97.5 | 96.5833 | 95.1857 | 94.5
13.7 | 97.9286 | 97.4875 | 96.5 | 95.0429 | 94.3667
13.8 | 97.9 | 97.35 | 96.5 | 95 | 94.3
13.9 | 97.7667 | 97.2833 | 96.3333 | 94.8429 | 94.2
14 | 97.7 | 97.2 | 96.2875 | 94.8 | 94.05
14.1 | 97.6286 | 97.16 | 96.2 | 94.7625 | 93.9
14.2 | 97.6 | 97.1 | 96.1167 | 94.6286 | 93.8
14.3 | 97.5 | 97.0429 | 96.0857 | 94.5 | 93.8
14.4 | 97.5 | 97 | 95.92 | 94.4333 | 93.7111
14.5 | 97.4833 | 96.9429 | 95.9 | 94.2714 | 93.5667
14.6 | 97.4 | 96.8857 | 95.8143 | 94.2 | 93.4143
14.7 | 97.4 | 96.8 | 95.6714 | 94.12 | 93.3667
14.8 | 97.3333 | 96.7833 | 95.6 | 94.05 | 93.3
14.9 | 97.25 | 96.6167 | 95.5 | 93.96 | 93.1571
15 | 97.2 | 96.58 | 95.3875 | 93.88 | 93.025
15.1 | 97.1 | 96.5 | 95.25 | 93.7 | 92.9
15.2 | 97.05 | 96.4167 | 95.2 | 93.5875 | 92.8
15.3 | 97 | 96.3 | 95.1 | 93.4625 | 92.7
15.4 | 96.9 | 96.25 | 95.1 | 93.3375 | 92.6286
15.5 | 96.7667 | 96.2 | 95.0857 | 93.3 | 92.45
15.6 | 96.7 | 96.1 | 95 | 93.1625 | 92.3
15.7 | 96.65 | 96.1 | 94.9143 | 93.0375 | 92.22
15.8 | 96.6 | 96 | 94.7571 | 92.9125 | 92.075
15.9 | 96.6 | 95.8667 | 94.625 | 92.8 | 91.95
16 | 96.5 | 95.8 | 94.5 | 92.6857 | 91.9
16.1 | 96.4 | 95.7 | 94.425 | 92.55 | 91.7
16.2 | 96.35 | 95.6 | 94.3 | 92.4 | 91.6
16.3 | 96.2 | 95.6 | 94.1333 | 92.3286 | 91.5625
16.4 | 96.1571 | 95.4667 | 94 | 92.2111 | 91.4375
16.5 | 96.0857 | 95.3857 | 93.9167 | 92.0714 | 91.3125
16.6 | 96 | 95.22 | 93.85 | 91.9286 | 91.2
16.7 | 95.9 | 95.1 | 93.7833 | 91.7875 | 91.075
16.8 | 95.8 | 95.025 | 93.7 | 91.7 | 90.9
16.9 | 95.7 | 94.9667 | 93.58 | 91.6143 | 90.725
17 | 95.62 | 94.9 | 93.4143 | 91.4429 | 90.6111
17.1 | 95.5429 | 94.7667 | 93.3 | 91.3125 | 90.4857
17.2 | 95.4 | 94.6143 | 93.1 | 91.1875 | 90.4
17.3 | 95.32 | 94.4667 | 93 | 91.0571 | 90.2714
17.4 | 95.24 | 94.3875 | 92.9 | 90.8286 | 90.1375
17.5 | 95.1 | 94.25 | 92.8 | 90.7 | 89.925
17.6 | 95.0286 | 94.1857 | 92.65 | 90.6 | 89.7875
17.7 | 94.9 | 94.02 | 92.4 | 90.5 | 89.7
17.8 | 94.8375 | 93.9571 | 92.3 | 90.2833 | 89.5889
17.9 | 94.7 | 93.9 | 92.2 | 90.2 | 89.4143
18 | 94.5667 | 93.76 | 92.0167 | 90.0222 | 89.2333
18.1 | 94.3833 | 93.5 | 91.8625 | 89.9 | 89.1
18.2 | 94.3 | 93.36 | 91.7375 | 89.7 | 89.025
18.3 | 94.2 | 93.2 | 91.6857 | 89.5429 | 88.8444
18.4 | 94.1 | 93.1 | 91.52 | 89.4 | 88.8
18.5 | 94.0167 | 92.95 | 91.2571 | 89.3 | 88.5857
18.6 | 93.8 | 92.8 | 91.1 | 89.2125 | 88.3833
18.7 | 93.68 | 92.6714 | 90.95 | 88.9778 | 88.2286
18.8 | 93.5 | 92.5167 | 90.8 | 88.7667 | 88.1
18.9 | 93.42 | 92.35 | 90.5667 | 88.6 | 87.94
19 | 93.2143 | 92.18 | 90.4 | 88.4667 | 87.86
19.1 | 92.9 | 92.0143 | 90.1429 | 88.2375 | 87.7
19.2 | 92.6429 | 91.7667 | 90 | 88.1125 | 87.5625
19.3 | 92.4 | 91.5714 | 89.8 | 87.8833 | 87.3333
19.4 | 91.86 | 91.3333 | 89.5429 | 87.7167 | 87.1667
19.5 | 91.15 | 91.05 | 89.38 | 87.5333 | 86.9167
19.6 | 90.2 | 90.7857 | 89.2143 | 87.25 | 86.7857
19.7 | 88.6 | 90.4833 | 88.9571 | 87.025 | 86.6
19.8 | 86.1571 | 90.2 | 88.7 | 86.8875 | 86.425
19.9 | 82.1667 | 89.7857 | 88.525 | 86.7143 | 86.3833
20 | 76.3143 | 89.3 | 88.2333 | 86.5 | 86.2167
20.1 | 68.6625 | 88.7 | 87.9333 | 86.4 | 85.9667
20.2 | 60.2 | 87.9 | 87.6833 | 86.15 | 85.8
20.3 | 52.0333 | 86.7143 | 87.4 | 85.9 | 85.65
20.4 | 43.6667 | 85.0667 | 87.14 | 85.55 | 85.45
20.5 | 35.7 | 82.575 | 86.8143 | 85.3 | 85.3333
20.6 | 29.7857 | 78.65 | 86.5 | 85.1143 | 85.125
20.7 | 24.875 | 73 | 86.1667 | 84.975 | 84.975
20.8 | 21.1857 | 66.3333 | 85.7 | 84.7 | 84.75
20.9 | 18.9833 | 58.6 | 85.1571 | 84.5125 | 84.4375
21 | 17.65 | 50.6 | 84.5857 | 84.325 | 84.3
21.1 | 16.8 | 43.5571 | 83.8 | 84.0333 | 84.0429
21.2 | 16.1667 | 36.5 | 82.9 | 83.8333 | 83.8167
21.3 | 15.78 | 31.3 | 81.9 | 83.5667 | 83.6143
21.4 | 15.4429 | 26.5857 | 80.1 | 83.34 | 83.4
21.5 | 15.175 | 23.5143 | 77.5 | 83.1 | 83.1667
21.6 | 14.925 | 21.0833 | 73.55 | 82.8 | 82.95
21.7 | 14.7 | 19.8333 | 68.5 | 82.4667 | 82.8
21.8 | 14.5667 | 18.78 | 62.6 | 82.1571 | 82.65
21.9 | 14.3167 | 18.15 | 56.1143 | 81.9 | 82.425
22 | 14.24 | 17.6143 | 49.2 | 81.5714 | 82.1429
22.1 | 14.04 | 17.1667 | 42.9143 | 81.25 | 81.875
22.2 | 13.8714 | 16.8333 | 37.05 | 80.9375 | 81.625
22.3 | 13.7286 | 16.5714 | 32.3667 | 80.725 | 81.4625
22.4 | 13.55 | 16.2667 | 28.4429 | 80.375 | 81.275
22.5 | 13.4 | 16 | 25.8 | 80.0875 | 81.025
22.6 | 13.3143 | 15.7667 | 23.8 | 79.7125 | 80.6778
22.7 | 13.1714 | 15.6 | 22.5875 | 79.1625 | 80.4714
22.8 | 13 | 15.3667 | 21.5857 | 78.7333 | 80.3286
22.9 | 12.9 | 15.1 | 21 | 78.3 | 80.0667
23 | 12.8 | 14.9714 | 20.42 | 77.6 | 79.775
23.1 | 12.6833 | 14.7333 | 19.9333 | 76.95 | 79.4
23.2 | 12.5167 | 14.5571 | 19.6 | 75.925 | 79.1333
23.3 | 12.46 | 14.4143 | 19.24 | 74.6 | 78.85
23.4 | 12.3 | 14.1714 | 18.9 | 72.8333 | 78.7111
23.5 | 12.2 | 14.0167 | 18.56 | 70.5286 | 78.3571
23.6 | 12.1167 | 13.8857 | 18.2 | 67.1 | 78.0429
23.7 | 12.0571 | 13.7333 | 18 | 63.2714 | 77.7714
23.8 | 12 | 13.5667 | 17.7714 | 58.6667 | 77.5
23.9 | 11.9 | 13.4 | 17.4667 | 53.3286 | 77.25
24 | 11.8 | 13.2857 | 17.3 | 47.9833 | 77
24.1 | 11.7375 | 13.1 | 17.0333 | 43.05 | 76.6333
24.2 | 11.6 | 13.0167 | 16.85 | 38.8 | 76.3143
24.3 | 11.5 | 12.95 | 16.6333 | 34.6 | 76.05
24.4 | 11.5 | 12.8 | 16.4143 | 31.7 | 75.7
24.5 | 11.3667 | 12.65 | 16.22 | 29.25 | 75.2625
24.6 | 11.3 | 12.5857 | 15.9714 | 27.5667 | 74.8875
24.7 | 11.25 | 12.5 | 15.775 | 26.4 | 74.3875
24.8 | 11.1 | 12.3625 | 15.5833 | 25.45 | 74
24.9 | 11.0833 | 12.2167 | 15.4286 | 24.7857 | 73.5
25 | 11 | 12.1571 | 15.2667 | 24.2143 | 72.8286
25.1 | 10.9 | 12.0143 | 15.1143 | 23.8 | 72.0667
25.2 | 10.8 | 11.95 | 14.8667 | 23.3714 | 70.9714
25.3 |  | 11.86 | 14.7143 | 23.0286 | 69.7
25.4 |  | 11.8 | 14.5571 | 22.7625 | 68.05
25.5 |  | 11.6167 | 14.4143 | 22.4143 | 65.65
25.6 |  | 11.575 | 14.3 | 22.0429 | 62.8667
25.7 |  | 11.5 | 14.1333 | 21.76 | 59.2
25.8 |  | 11.425 | 14 | 21.475 | 55.1625
25.9 |  | 11.3 | 13.9 | 21.25 | 50.75
26 |  |  | 13.74 | 20.9 | 46.375
26.1 |  |  | 13.5714 | 20.7667 | 42.3875
26.2 |  |  | 13.4167 | 20.4833 | 38.7875
26.3 |  |  | 13.34 | 20.2571 | 35.775
26.4 |  |  | 13.2 | 20 | 33.3375
26.5 |  |  | 13.1 | 19.7667 | 31.5333
26.6 |  |  | 12.95 | 19.5286 | 30
26.7 |  |  | 12.8889 | 19.3429 | 28.9
26.8 |  |  | 12.7667 | 19.075 | 28.1286
26.9 |  |  | 12.7 | 18.8 | 27.6
27 |  |  |  | 18.6333 | 27.15
27.1 |  |  |  | 18.3875 | 26.6333
27.2 |  |  |  | 18.25 | 26.2
27.3 |  |  |  | 18.0714 | 25.925
27.4 |  |  |  | 17.85 | 25.5667
27.5 |  |  |  | 17.65 | 25.275
27.6 |  |  |  | 17.5 | 25.025
27.7 |  |  |  | 17.3571 | 24.675
27.8 |  |  |  | 17.1 | 24.425
27.9 |  |  |  | 16.8875 | 24.1889
28 |  |  |  | 16.7625 | 23.875
28.1 |  |  |  | 16.6375 | 23.6
28.2 |  |  |  | 16.5125 | 23.375
28.3 |  |  |  | 16.2875 | 23.2
28.4 |  |  |  | 16.1625 | 22.875
28.5 |  |  |  | 15.9889 | 22.7111
28.6 |  |  |  | 15.8143 | 22.4833
28.7 |  |  |  | 15.6571 | 22.2
28.8 |  |  |  | 15.5 | 22.0333
28.9 |  |  |  | 15.4 | 21.7875
29 |  |  |  |  | 21.625
29.1 |  |  |  |  | 21.375
29.2 |  |  |  |  | 21.2
29.3 |  |  |  |  | 21.05
29.4 |  |  |  |  | 20.76
29.5 |  |  |  |  | 20.6
29.6 |  |  |  |  | 20.5
29.7 |  |  |  |  | 20.2286
29.8 |  |  |  |  | 20.075
29.9 |  |  |  |  | 19.825
30 |  |  |  |  | 19.6375
30.1 |  |  |  |  | 19.5125
30.2 |  |  |  |  | 19.2875
30.3 |  |  |  |  | 19.125
30.4 |  |  |  |  | 18.9375
30.5 |  |  |  |  | 18.8125
30.6 |  |  |  |  | 18.6875
30.7 |  |  |  |  | 18.525
30.8 |  |  |  |  | 18.3444
30.9 |  |  |  |  | 18.225

## Wedge 30 Output Factors
Output factors
Wedge 30 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.876 | 0.887 | 0.895 | 0.903 | 0.911 | 0.918 | 0.923 | 0.929
| 4 | 0.887 | 0.901 | 0.912 | 0.922 | 0.932 | 0.941 | 0.947 | 0.955
| 5 | 0.896 | 0.911 | 0.922 | 0.936 | 0.949 | 0.96 | 0.968 | 0.977
| 7 | 0.909 | 0.927 | 0.941 | 0.959 | 0.974 | 0.99 | 1.001 | 1.014
| 10 | 0.92 | 0.941 | 0.957 | 0.98 | 0.999 | 1.022 | 1.037 | 1.055
| 15 | 0.93 | 0.954 | 0.972 | 1 | 1.026 | 1.055 | 1.076 | 1.102
| 20 | 0.936 | 0.962 | 0.982 | 1.012 | 1.043 | 1.077 | 1.103 | 1.136
| 30 | 0.943 | 0.972 | 0.995 | 1.029 | 1.065 | 1.109 | 1.143 | 1.187
| 40 | 0.948 | 0.978 | 1.002 | 1.039 | 1.08 | 1.129 | 1.168 | 1.218

## Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 40.5 | 46 | 54.4 | 58.2
0.1 | 46.3667 | 51.26 | 59.54 | 62.4
0.2 | 55.2143 | 59.7143 | 66.65 | 67.6
0.3 | 66.3 | 69.7 | 75.1 | 74.45
0.4 | 74.6 | 77.6 | 82.15 | 80.825
0.5 | 81.85 | 83.8 | 87.14 | 86.12
0.6 | 86.46 | 88.4 | 91 | 90.425
0.7 | 90.5 | 91.8667 | 93.84 | 93.44
0.8 | 93.46 | 94.5 | 96.0167 | 95.75
0.9 | 95.8 | 96.34 | 97.5667 | 97.28
1 | 97.25 | 97.7 | 98.6 | 98.42
1.1 | 98.4167 | 98.6 | 99.2 | 99.2
1.2 | 99.1 | 99.4 | 99.7 | 99.6
1.3 | 99.66 | 99.825 | 99.95 | 99.85
1.4 | 99.9667 | 99.9 | 100 | 100
1.5 | 100 | 100 | 100 | 100
1.6 | 99.9714 | 99.98 | 99.9 | 99.9
1.7 | 99.8 | 99.875 | 99.74 | 99.675
1.8 | 99.575 | 99.7 | 99.58 | 99.52
1.9 | 99.2667 | 99.3667 | 99.26 | 99.36
2 | 99.0167 | 99.0667 | 98.75 | 99.02
2.1 | 98.55 | 98.75 | 98.6 | 98.75
2.2 | 98.3333 | 98.32 | 98.25 | 98.3
2.3 | 97.68 | 97.9 | 97.9 | 97.95
2.4 | 97.14 | 97.44 | 97.45 | 97.6
2.5 | 96.74 | 97.14 | 97.0667 | 97.275
2.6 | 96.2 | 96.75 | 96.86 | 96.975
2.7 | 95.66 | 96.3333 | 96.525 | 96.68
2.8 | 95.3 | 95.85 | 95.975 | 96.2
2.9 | 94.76 | 95.6 | 95.55 | 95.85
3 | 94.3 | 95.075 | 95.25 | 95.5
3.1 | 93.9167 | 94.76 | 94.84 | 95.2
3.2 | 93.2 | 94.3 | 94.6 | 94.7
3.3 | 92.9 | 93.94 | 94.08 | 94.34
3.4 | 92.4 | 93.5 | 93.7 | 93.9
3.5 | 91.9 | 92.98 | 93.38 | 93.56
3.6 | 91.45 | 92.58 | 93.05 | 93.1333
3.7 | 91 | 92.15 | 92.65 | 92.9
3.8 | 90.5 | 91.7667 | 92.25 | 92.54
3.9 | 90 | 91.35 | 91.95 | 92.1
4 | 89.58 | 91.06 | 91.54 | 91.78
4.1 | 89.02 | 90.64 | 91.2 | 91.35
4.2 | 88.45 | 90.225 | 90.85 | 91
4.3 | 88.06 | 89.85 | 90.38 | 90.75
4.4 | 87.6333 | 89.4 | 90 | 90.28
4.5 | 87.1 | 88.9 | 89.6667 | 90
4.6 | 86.675 | 88.44 | 89.2 | 89.5667
4.7 | 86.35 | 88.0667 | 88.95 | 89.3
4.8 | 85.8 | 87.825 | 88.6 | 88.8333
4.9 | 85.15 | 87.45 | 88.3 | 88.6
5 | 84.7 | 87.025 | 87.8667 | 88.3
5.1 | 84.2333 | 86.4667 | 87.4 | 87.84
5.2 | 83.94 | 86.1 | 87.06 | 87.4
5.3 | 83.36 | 85.7333 | 86.68 | 87.06
5.4 | 83 | 85.36 | 86.35 | 86.75
5.5 | 82.6333 | 84.98 | 85.9667 | 86.4
5.6 | 82 | 84.62 | 85.675 | 86
5.7 | 81.65 | 84.15 | 85.4 | 85.7
5.8 | 81 | 83.75 | 85.05 | 85.2667
5.9 | 80.65 | 83.32 | 84.5 | 85
6 | 80.18 | 82.925 | 84.2 | 84.6
6.1 | 79.7429 | 82.48 | 83.9667 | 84.16
6.2 | 79.2 | 82.25 | 83.55 | 83.8
6.3 | 78.94 | 81.85 | 83.16 | 83.5333
6.4 | 78.5 | 81.5 | 82.8 | 83.2
6.5 | 78.02 | 81.1 | 82.46 | 82.86
6.6 | 77.46 | 80.6 | 82.1 | 82.56
6.7 | 77.12 | 80.26 | 81.75 | 82.175
6.8 | 76.6833 | 79.85 | 81.4 | 81.85
6.9 | 76.275 | 79.52 | 81.1 | 81.525
7 | 75.8 | 79.2 | 80.64 | 81.125
7.1 | 75.3 | 78.6667 | 80.34 | 80.76
7.2 | 74.95 | 78.2 | 79.9667 | 80.48
7.3 | 74.52 | 77.8833 | 79.7333 | 80.15
7.4 | 74.0333 | 77.6 | 79.3429 | 79.8333
7.5 | 73.66 | 77.1667 | 79.1 | 79.46
7.6 | 73.3 | 76.8333 | 78.75 | 79.1667
7.7 | 72.86 | 76.35 | 78.36 | 78.7
7.8 | 72.52 | 76.0167 | 77.86 | 78.32
7.9 | 72.04 | 75.675 | 77.5667 | 78.06
8 | 71.6 | 75.38 | 77.2 | 77.8714
8.1 | 71.14 | 74.98 | 76.8 | 77.54
8.2 | 70.8 | 74.48 | 76.525 | 77.15
8.3 | 70.42 | 74.15 | 76.1667 | 76.8
8.4 | 69.98 | 73.78 | 75.8333 | 76.425
8.5 | 69.625 | 73.45 | 75.5667 | 76.22
8.6 | 69.3 | 73 | 75.28 | 75.92
8.7 | 68.8571 | 72.7 | 74.9 | 75.6
8.8 | 68.4 | 72.35 | 74.56 | 75.26
8.9 | 67.94 | 72.06 | 74.2667 | 74.9
9 | 67.52 | 71.72 | 73.84 | 74.52
9.1 | 67.225 | 71.35 | 73.52 | 74.14
9.2 | 66.75 | 71 | 73.24 | 73.8
9.3 | 66.45 | 70.55 | 72.9 | 73.54
9.4 | 65.95 | 70.14 | 72.6 | 73.3167
9.5 | 65.5 | 69.9 | 72.3 | 72.9333
9.6 | 65.225 | 69.56 | 72 | 72.58
9.7 | 64.9 | 69.22 | 71.7 | 72.2
9.8 | 64.5 | 68.85 | 71.3 | 72
9.9 | 64.12 | 68.35 | 70.875 | 71.65
10 | 63.7 | 68 | 70.65 | 71.32
10.1 | 63.25 | 67.7333 | 70.3 | 71.05
10.2 | 62.9 | 67.45 | 69.96 | 70.7
10.3 | 62.5667 | 67.1333 | 69.7 | 70.3667
10.4 | 62.175 | 66.8 | 69.4 | 70.1
10.5 | 61.9 | 66.4333 | 69.075 | 69.775
10.6 | 61.4 | 66.075 | 68.84 | 69.48
10.7 | 61.125 | 65.75 | 68.4333 | 69.14
10.8 | 60.8333 | 65.4 | 68.08 | 68.84
10.9 | 60.5 | 65.1 | 67.7 | 68.54
11 | 60.12 | 64.72 | 67.38 | 68.32
11.1 | 59.65 | 64.4 | 67.15 | 67.95
11.2 | 59.45 | 64.04 | 66.88 | 67.62
11.3 | 59.16 | 63.72 | 66.52 | 67.275
11.4 | 58.72 | 63.4 | 66.26 | 67.04
11.5 | 58.4 | 63.1 | 66 | 66.7
11.6 | 58.0333 | 62.75 | 65.76 | 66.4286
11.7 | 57.675 | 62.54 | 65.475 | 66.15
11.8 | 57.35 | 62.2333 | 65.15 | 65.825
11.9 | 56.8857 | 61.86 | 64.88 | 65.6
12 | 56.75 | 61.52 | 64.5333 | 65.2333
12.1 | 56.325 | 61.15 | 64.2286 | 64.94
12.2 | 56 | 60.86 | 63.825 | 64.64
12.3 | 55.6667 | 60.55 | 63.62 | 64.38
12.4 | 55.3 | 60.15 | 63.25 | 64.15
12.5 | 55.0667 | 59.8667 | 62.94 | 63.875
12.6 | 54.7 | 59.575 | 62.6333 | 63.58
12.7 | 54.34 | 59.2667 | 62.45 | 63.3
12.8 | 54.0833 | 58.9 | 62.1667 | 62.9
12.9 | 53.7667 | 58.5667 | 62 | 62.8167
13 | 53.4833 | 58.36 | 61.6 | 62.54
13.1 | 53.2333 | 58.06 | 61.18 | 62.08
13.2 | 52.88 | 57.7667 | 61.025 | 61.825
13.3 | 52.5667 | 57.475 | 60.74 | 61.56
13.4 | 52.2333 | 57.1 | 60.4 | 61.25
13.5 | 52.02 | 56.74 | 60.08 | 60.98
13.6 | 51.72 | 56.52 | 59.78 | 60.82
13.7 | 51.4 | 56.28 | 59.55 | 60.5
13.8 | 51.05 | 56.02 | 59.32 | 60.14
13.9 | 50.74 | 55.625 | 59 | 59.92
14 | 50.5 | 55.3333 | 58.74 | 59.6571
14.1 | 50.15 | 54.9667 | 58.4333 | 59.38
14.2 | 49.88 | 54.76 | 58.2 | 59.12
14.3 | 49.5667 | 54.5667 | 58 | 58.8
14.4 | 49.32 | 54.26 | 57.75 | 58.46
14.5 | 48.98 | 53.96 | 57.3833 | 58.275
14.6 | 48.6 | 53.6667 | 57.18 | 58.04
14.7 | 48.28 | 53.3333 | 56.8 | 57.7667
14.8 | 48.1 | 53.025 | 56.4833 | 57.4667
14.9 | 47.875 | 52.78 | 56.3 | 57.3
15 | 47.55 | 52.475 | 55.9667 | 57
15.1 | 47.3 | 52.2571 | 55.7167 | 56.7
15.2 | 47.06 | 51.95 | 55.525 | 56.3667
15.3 | 46.76 | 51.6667 | 55.22 | 56.16
15.4 | 46.44 | 51.4 | 55 | 55.925
15.5 | 46.3 | 51.2 | 54.64 | 55.72
15.6 | 45.96 | 50.92 | 54.5 | 55.475
15.7 | 45.675 | 50.675 | 54.26 | 55.15
15.8 | 45.4 | 50.38 | 53.975 | 54.9
15.9 | 45.2 | 50.1143 | 53.76 | 54.7
16 | 44.9 | 49.9 | 53.4 | 54.55
16.1 | 44.7 | 49.7 | 53.14 | 54.14
16.2 | 44.5 | 49.36 | 52.975 | 53.925
16.3 | 44.08 | 49.0833 | 52.74 | 53.7
16.4 | 43.9 | 48.8 | 52.48 | 53.45
16.5 | 43.6 | 48.54 | 52.3 | 53.2333
16.6 | 43.425 | 48.2333 | 52.02 | 53.04
16.7 | 43.175 | 48.05 | 51.8 | 52.8
16.8 | 42.9 | 47.8 | 51.52 | 52.55
16.9 | 42.6 | 47.56 | 51.1667 | 52.3
17 | 42.4 | 47.34 | 50.95 | 52.075
17.1 | 42.1 | 47.125 | 50.82 | 51.7333
17.2 | 41.9 | 46.8 | 50.6 | 51.3833
17.3 | 41.68 | 46.5429 | 50.25 | 51.2167
17.4 | 41.45 | 46.2 | 50.04 | 51.0833
17.5 | 41.15 | 46.1 | 49.75 | 50.8333
17.6 | 40.95 | 45.8333 | 49.5667 | 50.6333
17.7 | 40.72 | 45.55 | 49.3167 | 50.36
17.8 | 40.4667 | 45.3667 | 49.12 | 50.16
17.9 | 40.2714 | 45.1167 | 48.825 | 49.84
18 | 40.1167 | 44.9 | 48.66 | 49.65
18.1 | 39.8333 | 44.68 | 48.3667 | 49.425
18.2 | 39.55 | 44.4 | 48.2667 | 49.2333
18.3 | 39.3 | 44.175 | 47.925 | 49
18.4 | 39.1 | 43.88 | 47.6857 | 48.8
18.5 | 38.8667 | 43.675 | 47.46 | 48.52
18.6 | 38.625 | 43.45 | 47.2667 | 48.24
18.7 | 38.4 | 43.22 | 47.075 | 48.08
18.8 | 38.15 | 43.025 | 46.84 | 47.9143
18.9 | 37.9 | 42.74 | 46.5 | 47.625
19 | 37.75 | 42.5 | 46.2429 | 47.325
19.1 | 37.55 | 42.26 | 46.08 | 47.1
19.2 | 37.32 | 41.96 | 45.9 | 46.92
19.3 | 37.04 | 41.75 | 45.7 | 46.74
19.4 | 36.9333 | 41.55 | 45.475 | 46.5
19.5 | 36.66 | 41.36 | 45.2 | 46.2833
19.6 | 36.425 | 41.16 | 45.04 | 46.1167
19.7 | 36.28 | 40.94 | 44.84 | 45.8
19.8 | 36 | 40.7333 | 44.6 | 45.6
19.9 | 35.82 | 40.55 | 44.44 | 45.45
20 | 35.58 | 40.28 | 44.1333 | 45.3
20.1 | 35.4 | 40.16 | 43.9143 | 45.1
20.2 | 35.275 | 39.96 | 43.72 | 44.7667
20.3 | 35.0333 | 39.66 | 43.5 | 44.5667
20.4 | 34.8 | 39.4667 | 43.3 | 44.375
20.5 | 34.7 | 39.3 | 43.2 | 44.25
20.6 | 34.44 | 39.1 | 42.96 | 44.05
20.7 | 34.275 | 38.88 | 42.7 | 43.82
20.8 | 34.15 | 38.64 | 42.5 | 43.6
20.9 | 33.8571 | 38.45 | 42.3 | 43.44
21 | 33.68 | 38.15 | 42.1 | 43.15
21.1 | 33.5 | 38.0333 | 41.975 | 43.04
21.2 | 33.3143 | 37.82 | 41.725 | 42.7833
21.3 | 33.0333 | 37.6 | 41.45 | 42.5333
21.4 | 32.925 | 37.4 | 41.3 | 42.42
21.5 | 32.8 | 37.28 | 41.1 | 42.18
21.6 | 32.5833 | 37.1 | 40.98 | 41.94
21.7 | 32.4167 | 36.85 | 40.775 | 41.74
21.8 | 32.15 | 36.64 | 40.54 | 41.6
21.9 | 32.02 | 36.45 | 40.34 | 41.3833
22 | 31.8 | 36.36 | 40.14 | 41.1833
22.1 | 31.6 | 36.05 | 39.9667 | 41.1
22.2 | 31.48 | 35.9 | 39.8 | 40.8667
22.3 | 31.3 | 35.72 | 39.5833 | 40.68
22.4 | 31.12 | 35.525 | 39.375 | 40.55
22.5 | 30.94 | 35.3 | 39.2 | 40.325
22.6 | 30.8 | 35.15 | 38.975 | 40.1
22.7 | 30.6 | 34.92 | 38.7333 | 39.88
22.8 | 30.42 | 34.7333 | 38.58 | 39.675
22.9 | 30.2 | 34.5667 | 38.5 | 39.44
23 | 30.1 | 34.4 | 38.22 | 39.3
23.1 | 29.98 | 34.225 | 38.1 | 39.12
23.2 | 29.675 | 34 | 37.9 | 38.9333
23.3 | 29.525 | 33.84 | 37.7 | 38.7714
23.4 | 29.3167 | 33.7571 | 37.52 | 38.6
23.5 | 29.22 | 33.5667 | 37.3 | 38.475
23.6 | 29.04 | 33.45 | 37.1 | 38.3
23.7 | 28.9 | 33.2 | 37 | 38.15
23.8 | 28.7 | 32.98 | 36.8 | 37.9667
23.9 | 28.6 | 32.8 | 36.6 | 37.7
24 | 28.4 | 32.65 | 36.4333 | 37.58
24.1 | 28.18 | 32.44 | 36.2 | 37.36
24.2 | 28.06 | 32.3 | 36.05 | 37.12
24.3 | 27.85 | 32.1 | 35.9 | 37
24.4 | 27.725 | 31.88 | 35.7167 | 36.8833
24.5 | 27.6 | 31.7 | 35.6 | 36.7
24.6 | 27.4 | 31.56 | 35.36 | 36.5
24.7 | 27.3 | 31.44 | 35.225 | 36.3
24.8 | 27.1 | 31.2667 | 35.02 | 36.2
24.9 | 27 | 31.1 | 34.92 | 35.95
25 | 26.86 | 31 | 34.72 | 35.74
25.1 | 26.625 | 30.85 | 34.6 | 35.64
25.2 | 26.4833 | 30.64 | 34.375 | 35.5
25.3 | 26.4 | 30.5 | 34.25 | 35.375
25.4 | 26.225 | 30.3 | 34.1333 | 35.2
25.5 | 26.1 | 30.18 | 33.92 | 34.9333
25.6 | 25.9 | 30 | 33.8 | 34.82
25.7 | 25.8 | 29.8667 | 33.62 | 34.7
25.8 | 25.7 | 29.75 | 33.48 | 34.55
25.9 | 25.52 | 29.6 | 33.2 | 34.36
26 | 25.4 | 29.45 | 33.0833 | 34.2
26.1 | 25.2 | 29.2 | 32.9 | 34.0333
26.2 | 25.1 | 29.075 | 32.7833 | 33.825
26.3 | 24.88 | 28.96 | 32.6286 | 33.7
26.4 | 24.7667 | 28.8 | 32.44 | 33.525
26.5 | 24.7 | 28.66 | 32.2714 | 33.4
26.6 | 24.55 | 28.5 | 32.1 | 33.275
26.7 | 24.4 | 28.38 | 32 | 33.1
26.8 | 24.25 | 28.1 | 31.85 | 32.94
26.9 | 24.14 | 28.02 | 31.675 | 32.76
27 | 24 | 27.9 | 31.4833 | 32.625
27.1 | 23.85 | 27.7667 | 31.3 | 32.5
27.2 | 23.7667 | 27.5833 | 31.2333 | 32.3167
27.3 | 23.6 | 27.4167 | 31.1 | 32.2
27.4 | 23.5 | 27.34 | 30.9 | 32.1
27.5 | 23.3 | 27.06 | 30.8167 | 31.9333
27.6 | 23.16 | 26.95 | 30.64 | 31.7
27.7 | 23.1 | 26.82 | 30.5 | 31.5167
27.8 | 22.94 | 26.7 | 30.3833 | 31.3
27.9 | 22.8333 | 26.5333 | 30.1833 | 31.175
28 | 22.675 | 26.4 | 30.0167 | 31
28.1 | 22.6 | 26.3167 | 29.8333 | 30.9
28.2 | 22.5 | 26.15 | 29.7 | 30.7333
28.3 | 22.35 | 26.05 | 29.6 | 30.6
28.4 | 22.2 | 25.95 | 29.4714 | 30.46
28.5 | 22.1 | 25.74 | 29.3 | 30.3
28.6 | 21.98 | 25.6 | 29.1333 | 30.2
28.7 | 21.9 | 25.5 | 29 | 30.06
28.8 | 21.7 | 25.3 | 28.8667 | 29.9
28.9 | 21.62 | 25.22 | 28.74 | 29.8
29 | 21.5 | 25.1 | 28.6 | 29.6
29.1 | 21.3 | 24.925 | 28.4667 | 29.54
29.2 | 21.2 | 24.84 | 28.36 | 29.4
29.3 | 21.1 | 24.7 | 28.2 | 29.2
29.4 | 21 | 24.6 | 28.1 | 29.06
29.5 | 20.85 | 24.5 | 27.8833 | 28.95
29.6 | 20.8 | 24.3 | 27.78 | 28.7667
29.7 | 20.7 | 24.225 | 27.6 | 28.7
29.8 | 20.6 | 24.1 | 27.5167 | 28.52
29.9 | 20.42 | 24 | 27.4 | 28.4
30 | 20.3 | 23.85 | 27.28 | 28.25

## SSD80 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 40.9 | 47.2 | 58.8 | 66
0.1 | 46.475 | 52 | 63.775 | 70.45
0.2 | 54.2 | 60 | 70.875 | 76.4143
0.3 | 65.72 | 70.54 | 78.4 | 83.0333
0.4 | 74.7 | 78.55 | 85.04 | 88.3
0.5 | 81.7 | 84.9 | 89.9 | 92.2
0.6 | 87 | 89.45 | 93.1 | 94.9375
0.7 | 91 | 92.8 | 95.65 | 96.875
0.8 | 93.8 | 95.3 | 97.3 | 98.275
0.9 | 96.1 | 97.1333 | 98.5 | 99.1625
1 | 97.4 | 98.25 | 99.3 | 99.72
1.1 | 98.7 | 99.075 | 99.8 | 100
1.2 | 99.4667 | 99.725 | 100 | 100
1.3 | 99.85 | 99.975 | 100 | 100
1.4 | 100 | 100 | 100 | 99.9333
1.5 | 100 | 100 | 99.875 | 99.75
1.6 | 99.975 | 99.875 | 99.725 | 99.4
1.7 | 99.8 | 99.5375 | 99.275 | 99.1
1.8 | 99.25 | 99.2 | 98.975 | 98.66
1.9 | 98.925 | 98.95 | 98.65 | 98.35
2 | 98.55 | 98.6 | 98.2 | 97.925
2.1 | 98.14 | 98.15 | 97.78 | 97.55
2.2 | 97.62 | 97.725 | 97.35 | 97.1333
2.3 | 97.0333 | 97.2 | 96.9333 | 96.68
2.4 | 96.56 | 96.7 | 96.5 | 96.25
2.5 | 96.05 | 96.3667 | 96.05 | 95.8
2.6 | 95.4333 | 95.7 | 95.7 | 95.425
2.7 | 94.8 | 95.38 | 95.15 | 95
2.8 | 94.4 | 94.9 | 94.8 | 94.5
2.9 | 93.7 | 94.5 | 94.35 | 94
3 | 93.25 | 94.05 | 93.8 | 93.7
3.1 | 92.7 | 93.55 | 93.4 | 93.2
3.2 | 92.22 | 92.975 | 93 | 92.85
3.3 | 91.75 | 92.55 | 92.55 | 92.42
3.4 | 91.05 | 92.05 | 92.1 | 91.95
3.5 | 90.525 | 91.65 | 91.6 | 91.55
3.6 | 90.025 | 91.14 | 91.225 | 91.125
3.7 | 89.45 | 90.65 | 90.775 | 90.65
3.8 | 89.025 | 90.25 | 90.35 | 90.2
3.9 | 88.45 | 89.7 | 89.95 | 89.9
4 | 87.84 | 89.3 | 89.55 | 89.4
4.1 | 87.35 | 88.8 | 89.125 | 88.9333
4.2 | 86.76 | 88.4 | 88.65 | 88.6333
4.3 | 86.4 | 87.9 | 88.225 | 88.2
4.4 | 85.8 | 87.45 | 87.8286 | 87.75
4.5 | 85.3 | 86.975 | 87.35 | 87.275
4.6 | 84.75 | 86.55 | 86.95 | 86.9667
4.7 | 84.35 | 86 | 86.4667 | 86.5
4.8 | 83.6 | 85.675 | 86.2 | 86.025
4.9 | 83.15 | 85.15 | 85.7 | 85.65
5 | 82.6 | 84.75 | 85.2 | 85.2
5.1 | 82.16 | 84.35 | 84.8 | 84.8
5.2 | 81.7 | 83.8 | 84.35 | 84.45
5.3 | 81.125 | 83.3 | 83.8667 | 84
5.4 | 80.65 | 82.9 | 83.5 | 83.6
5.5 | 80.075 | 82.55 | 83.05 | 83.25
5.6 | 79.68 | 81.9 | 82.6 | 82.9
5.7 | 79.2 | 81.55 | 82.4 | 82.4
5.8 | 78.6 | 81.125 | 82 | 82.1
5.9 | 78.3 | 80.6667 | 81.55 | 81.65
6 | 77.6667 | 80.2 | 81.075 | 81.25
6.1 | 77.1667 | 79.75 | 80.75 | 80.85
6.2 | 76.8 | 79.4 | 80.44 | 80.45
6.3 | 76.45 | 79.05 | 80 | 80.05
6.4 | 75.88 | 78.575 | 79.6 | 79.675
6.5 | 75.35 | 77.96 | 79 | 79.325
6.6 | 74.775 | 77.55 | 78.775 | 78.8667
6.7 | 74.425 | 77.2 | 78.45 | 78.575
6.8 | 73.95 | 76.8 | 78.05 | 78.25
6.9 | 73.375 | 76.4 | 77.65 | 77.8
7 | 72.9 | 76 | 77.15 | 77.35
7.1 | 72.55 | 75.6 | 76.8333 | 77
7.2 | 72.15 | 75.1 | 76.4333 | 76.75
7.3 | 71.65 | 74.675 | 76.05 | 76.25
7.4 | 71.175 | 74.3333 | 75.5875 | 75.95
7.5 | 70.86 | 73.9625 | 75.2125 | 75.525
7.6 | 70.38 | 73.5714 | 74.95 | 75.3
7.7 | 69.8 | 73.18 | 74.54 | 74.8
7.8 | 69.5 | 72.7 | 74.075 | 74.4
7.9 | 69.1 | 72.3125 | 73.725 | 74.06
8 | 68.6 | 71.85 | 73.375 | 73.65
8.1 | 68.2 | 71.45 | 73.05 | 73.375
8.2 | 67.7 | 71.1 | 72.8 | 73
8.3 | 67.34 | 70.76 | 72.25 | 72.6
8.4 | 66.85 | 70.425 | 71.9333 | 72.3
8.5 | 66.45 | 69.925 | 71.6 | 71.9
8.6 | 65.975 | 69.45 | 71.2 | 71.55
8.7 | 65.65 | 69.075 | 70.9 | 71.22
8.8 | 65.2 | 68.7 | 70.525 | 70.8333
8.9 | 64.8 | 68.3 | 70.175 | 70.5
9 | 64.4 | 67.95 | 69.75 | 70.15
9.1 | 63.95 | 67.6 | 69.45 | 69.7875
9.2 | 63.7 | 67.22 | 69.05 | 69.4125
9.3 | 63.2333 | 66.8 | 68.675 | 69.15
9.4 | 62.7 | 66.42 | 68.35 | 68.7667
9.5 | 62.36 | 66.05 | 68 | 68.38
9.6 | 61.95 | 65.675 | 67.72 | 68.1
9.7 | 61.54 | 65.325 | 67.3 | 67.65
9.8 | 61.2571 | 65.05 | 66.9 | 67.35
9.9 | 60.84 | 64.7 | 66.575 | 67.0333
10 | 60.475 | 64.25 | 66.2 | 66.65
10.1 | 60.225 | 63.8 | 65.86 | 66.35
10.2 | 59.65 | 63.5333 | 65.525 | 66.025
10.3 | 59.2 | 63.1 | 65.2 | 65.65
10.4 | 58.9 | 62.7667 | 64.85 | 65.325
10.5 | 58.525 | 62.375 | 64.6 | 64.98
10.6 | 58.175 | 61.98 | 64.25 | 64.7
10.7 | 57.675 | 61.6625 | 63.9 | 64.4
10.8 | 57.375 | 61.28 | 63.6 | 63.95
10.9 | 57.05 | 60.95 | 63.2 | 63.7
11 | 56.7 | 60.6 | 62.85 | 63.45
11.1 | 56.4 | 60.3 | 62.6 | 63.125
11.2 | 56 | 59.94 | 62.25 | 62.75
11.3 | 55.6667 | 59.64 | 61.925 | 62.425
11.4 | 55.3 | 59.175 | 61.575 | 62.175
11.5 | 54.975 | 58.8714 | 61.25 | 61.85
11.6 | 54.725 | 58.5 | 60.975 | 61.5
11.7 | 54.275 | 58.2 | 60.65 | 61.25
11.8 | 53.9 | 57.925 | 60.3 | 60.9
11.9 | 53.58 | 57.475 | 60 | 60.6
12 | 53.175 | 57.15 | 59.75 | 60.3
12.1 | 52.85 | 56.875 | 59.4 | 59.95
12.2 | 52.5 | 56.55 | 59.1429 | 59.7
12.3 | 52.2 | 56.2 | 58.76 | 59.35
12.4 | 51.9 | 55.95 | 58.525 | 59.1
12.5 | 51.7 | 55.6 | 58.225 | 58.825
12.6 | 51.3 | 55.35 | 57.86 | 58.5
12.7 | 50.975 | 55 | 57.5 | 58.175
12.8 | 50.75 | 54.675 | 57.25 | 57.86
12.9 | 50.4 | 54.325 | 57 | 57.6
13 | 49.975 | 53.975 | 56.65 | 57.35
13.1 | 49.625 | 53.65 | 56.4 | 56.9429
13.2 | 49.4 | 53.5 | 56.175 | 56.76
13.3 | 49.1 | 53.1 | 55.825 | 56.525
13.4 | 48.85 | 52.8 | 55.55 | 56.2333
13.5 | 48.5 | 52.55 | 55.125 | 55.88
13.6 | 48.2 | 52.2 | 54.875 | 55.6
13.7 | 47.95 | 51.85 | 54.55 | 55.3
13.8 | 47.625 | 51.55 | 54.275 | 55.05
13.9 | 47.275 | 51.35 | 54.025 | 54.7
14 | 46.95 | 50.975 | 53.8 | 54.375
14.1 | 46.6 | 50.725 | 53.4333 | 54.15
14.2 | 46.425 | 50.45 | 53.2 | 53.875
14.3 | 46.2 | 50.1 | 52.95 | 53.625
14.4 | 45.85 | 49.825 | 52.7 | 53.375
14.5 | 45.5 | 49.575 | 52.3667 | 53.125
14.6 | 45.26 | 49.2875 | 52.125 | 52.85
14.7 | 44.9 | 48.9333 | 51.775 | 52.525
14.8 | 44.7 | 48.6286 | 51.525 | 52.25
14.9 | 44.4 | 48.4 | 51.275 | 51.95
15 | 44.025 | 48.075 | 51.025 | 51.7
15.1 | 43.875 | 47.8 | 50.7667 | 51.5
15.2 | 43.6 | 47.52 | 50.45 | 51.1
15.3 | 43.375 | 47.225 | 50.2 | 50.9333
15.4 | 43.025 | 47.075 | 49.925 | 50.65
15.5 | 42.775 | 46.75 | 49.6 | 50.4
15.6 | 42.525 | 46.4 | 49.3714 | 50.15
15.7 | 42.3 | 46.15 | 49.1333 | 49.94
15.8 | 41.95 | 45.9 | 48.875 | 49.625
15.9 | 41.7 | 45.7 | 48.625 | 49.45
16 | 41.475 | 45.4 | 48.4 | 49.225
16.1 | 41.225 | 45.2 | 48.1875 | 48.9
16.2 | 41.1 | 44.94 | 47.9 | 48.725
16.3 | 40.7333 | 44.6 | 47.6429 | 48.3857
16.4 | 40.45 | 44.4 | 47.3889 | 48.175
16.5 | 40.2 | 44.1 | 47.1286 | 47.95
16.6 | 39.95 | 43.9 | 46.86 | 47.7
16.7 | 39.72 | 43.6333 | 46.6 | 47.45
16.8 | 39.45 | 43.35 | 46.35 | 47.2
16.9 | 39.225 | 43.1 | 46.12 | 46.9
17 | 38.95 | 42.85 | 45.825 | 46.7
17.1 | 38.7 | 42.6 | 45.65 | 46.45
17.2 | 38.525 | 42.3 | 45.4 | 46.2
17.3 | 38.3333 | 42.125 | 45.2 | 46
17.4 | 37.975 | 41.875 | 44.95 | 45.825
17.5 | 37.825 | 41.625 | 44.7 | 45.475
17.6 | 37.55 | 41.45 | 44.45 | 45.225
17.7 | 37.3 | 41.225 | 44.275 | 45.06
17.8 | 37.1 | 40.9 | 44.05 | 44.925
17.9 | 36.95 | 40.7 | 43.8 | 44.575
18 | 36.7 | 40.525 | 43.55 | 44.325
18.1 | 36.45 | 40.2 | 43.3 | 44.1
18.2 | 36.2333 | 40 | 43.1 | 43.95
18.3 | 36 | 39.8 | 42.825 | 43.72
18.4 | 35.775 | 39.55 | 42.675 | 43.5
18.5 | 35.525 | 39.3 | 42.425 | 43.2667
18.6 | 35.3 | 39.075 | 42.175 | 42.975
18.7 | 35.1 | 38.925 | 41.925 | 42.7833
18.8 | 34.88 | 38.68 | 41.7 | 42.575
18.9 | 34.6571 | 38.45 | 41.525 | 42.325
19 | 34.5667 | 38.2 | 41.225 | 42.0875
19.1 | 34.2333 | 37.95 | 41.125 | 41.9625
19.2 | 34.05 | 37.8 | 40.875 | 41.7571
19.3 | 33.8 | 37.575 | 40.625 | 41.5
19.4 | 33.6 | 37.325 | 40.4 | 41.25
19.5 | 33.425 | 37.075 | 40.25 | 41.12
19.6 | 33.175 | 36.95 | 40 | 40.9
19.7 | 33.1 | 36.6857 | 39.8 | 40.725
19.8 | 32.775 | 36.5 | 39.5333 | 40.475
19.9 | 32.525 | 36.3 | 39.3 | 40.225
20 | 32.4 | 36 | 39.15 | 39.975
20.1 | 32.15 | 35.875 | 38.9 | 39.825
20.2 | 32 | 35.625 | 38.7 | 39.575
20.3 | 31.85 | 35.4 | 38.6 | 39.4
20.4 | 31.62 | 35.225 | 38.35 | 39.2
20.5 | 31.4286 | 35.075 | 38.1 | 39
20.6 | 31.3111 | 34.8 | 37.9 | 38.7667
20.7 | 31.075 | 34.65 | 37.675 | 38.55
20.8 | 30.8625 | 34.4 | 37.55 | 38.42
20.9 | 30.6125 | 34.275 | 37.3667 | 38.16
21 | 30.4222 | 34.025 | 37.2125 | 38.025
21.1 | 30.2875 | 33.8 | 36.8875 | 37.8
21.2 | 30.125 | 33.625 | 36.7571 | 37.625
21.3 | 30 | 33.475 | 36.575 | 37.4
21.4 | 29.7333 | 33.225 | 36.34 | 37.25
21.5 | 29.6 | 33.075 | 36.18 | 37
21.6 | 29.4 | 32.8625 | 35.95 | 36.85
21.7 | 29.2 | 32.7375 | 35.8 | 36.7
21.8 | 29 | 32.48 | 35.575 | 36.5
21.9 | 28.9 | 32.3 | 35.42 | 36.3
22 | 28.65 | 32.14 | 35.2 | 36.15
22.1 | 28.5 | 31.975 | 35.025 | 35.9
22.2 | 28.35 | 31.75 | 34.8 | 35.7
22.3 | 28.2 | 31.6 | 34.64 | 35.6
22.4 | 28 | 31.45 | 34.5 | 35.3667
22.5 | 27.9 | 31.2 | 34.3 | 35.2
22.6 | 27.75 | 31.05 | 34.1 | 35
22.7 | 27.5 | 30.9 | 33.95 | 34.9
22.8 | 27.35 | 30.75 | 33.8 | 34.675
22.9 | 27.2 | 30.5 | 33.6 | 34.5
23 | 27.05 | 30.4 | 33.4 | 34.3
23.1 | 26.84 | 30.2333 | 33.2 | 34.125
23.2 | 26.7 | 30 | 33.025 | 33.9
23.3 | 26.5714 | 29.86 | 32.875 | 33.75
23.4 | 26.3667 | 29.725 | 32.7 | 33.6
23.5 | 26.1857 | 29.46 | 32.575 | 33.45
23.6 | 26.05 | 29.325 | 32.325 | 33.3
23.7 | 25.9 | 29.2 | 32.2 | 33.1
23.8 | 25.775 | 29.025 | 32 | 32.9
23.9 | 25.6333 | 28.8 | 31.8 | 32.75
24 | 25.5 | 28.65 | 31.7 | 32.6
24.1 | 25.275 | 28.5 | 31.6 | 32.45
24.2 | 25.1 | 28.35 | 31.35 | 32.22
24.3 | 25 | 28.2 | 31.2 | 32.05
24.4 | 24.85 | 28.05 | 31.1 | 31.95
24.5 | 24.7 | 27.8 | 30.9 | 31.75
24.6 | 24.55 | 27.7 | 30.7 | 31.6143
24.7 | 24.4 | 27.6 | 30.5 | 31.475
24.8 | 24.2667 | 27.4 | 30.4 | 31.26
24.9 | 24.1 | 27.2 | 30.2 | 31.1
25 | 23.9 | 27.1 | 30.025 | 30.9714
25.1 | 23.8 | 26.95 | 29.9 | 30.8
25.2 | 23.7 | 26.7625 | 29.7 | 30.65
25.3 | 23.5 | 26.6 | 29.525 | 30.5
25.4 | 23.32 | 26.5 | 29.4 | 30.36
25.5 | 23.3 | 26.3667 | 29.25 | 30.125
25.6 | 23.1 | 26.1667 | 29.1 | 30
25.7 | 22.925 | 26 | 28.95 | 29.8333
25.8 | 22.9 | 25.825 | 28.8 | 29.7
25.9 | 22.7667 | 25.7 | 28.7 | 29.56
26 | 22.6 | 25.6 | 28.5333 | 29.4
26.1 | 22.4 | 25.475 | 28.35 | 29.275
26.2 | 22.3 | 25.3 | 28.2 | 29.12
26.3 | 22.24 | 25.125 | 28.0714 | 28.975
26.4 | 22.1 | 25 | 27.9 | 28.8
26.5 | 21.975 | 24.9 | 27.725 | 28.675
26.6 | 21.8 | 24.8 | 27.6 | 28.5
26.7 | 21.675 | 24.6 | 27.475 | 28.35
26.8 | 21.6 | 24.5 | 27.3 | 28.225
26.9 | 21.3875 | 24.3 | 27.2 | 28.1
27 | 21.3 | 24.2 | 27.05 | 27.925
27.1 | 21.2 | 24.1 | 26.9 | 27.775
27.2 | 21.1 | 23.9 | 26.75 | 27.625
27.3 | 20.95 | 23.8 | 26.62 | 27.5
27.4 | 20.8 | 23.675 | 26.475 | 27.325
27.5 | 20.7 | 23.54 | 26.325 | 27.2
27.6 | 20.56 | 23.4 | 26.275 | 27.1
27.7 | 20.425 | 23.25 | 26.1 | 26.98
27.8 | 20.3 | 23.1333 | 25.9 | 26.8
27.9 | 20.225 | 23 | 25.75 | 26.7
28 | 20.1 | 22.825 | 25.7 | 26.5
28.1 | 20 | 22.7 | 25.55 | 26.45
28.2 | 19.875 | 22.6 | 25.375 | 26.3
28.3 | 19.7 | 22.5 | 25.25 | 26.15
28.4 | 19.6 | 22.4 | 25.175 | 26
28.5 | 19.5 | 22.25 | 25 | 25.9
28.6 | 19.4 | 22.12 | 24.8833 | 25.7
28.7 | 19.3 | 21.9889 | 24.675 | 25.6
28.8 | 19.2 | 21.875 | 24.6 | 25.45
28.9 | 19.1 | 21.75 | 24.5 | 25.3
29 | 18.925 | 21.625 | 24.3 | 25.25
29.1 | 18.875 | 21.4667 | 24.2143 | 25.1
29.2 | 18.7333 | 21.4 | 24.1 | 25
29.3 | 18.6 | 21.26 | 24 | 24.875
29.4 | 18.4444 | 21.16 | 23.8 | 24.725
29.5 | 18.4 | 21.025 | 23.7 | 24.6
29.6 | 18.3 | 20.9 | 23.6 | 24.5
29.7 | 18.175 | 20.825 | 23.45 | 24.4
29.8 | 18.05 | 20.675 | 23.4 | 24.25
29.9 | 17.925 | 20.575 | 23.2 | 24.1
30 | 17.9 | 20.4 | 23.1 | 24

## SSD120 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 40.6 | 45.9 | 53.5 | 57.8
0.1 | 46.025 | 51.05 | 58.2 | 62.7
0.2 | 54.625 | 58.8 | 65.5 | 69.1
0.3 | 65.9 | 69.45 | 74.7 | 77.38
0.4 | 74.45 | 77.4 | 81.25 | 83.525
0.5 | 80.8 | 83.7 | 86.5 | 88.25
0.6 | 86 | 88 | 90.45 | 91.8333
0.7 | 89.9 | 91.4333 | 93.4 | 94.125
0.8 | 92.9 | 93.9 | 95.3667 | 96.3667
0.9 | 95.4 | 96 | 96.95 | 97.6
1 | 96.9 | 97.5 | 98.3333 | 98.7333
1.1 | 98.05 | 98.425 | 99.05 | 99.1667
1.2 | 98.9 | 99.225 | 99.525 | 99.55
1.3 | 99.6 | 99.725 | 99.875 | 99.875
1.4 | 99.9 | 99.975 | 99.9 | 100
1.5 | 100 | 100 | 100 | 100
1.6 | 100 | 100 | 100 | 100
1.7 | 99.9 | 99.9 | 99.9 | 99.8
1.8 | 99.8 | 99.8 | 99.8 | 99.65
1.9 | 99.5 | 99.5 | 99.45 | 99.4
2 | 99.175 | 99.25 | 99.15 | 99.1
2.1 | 98.825 | 98.9 | 98.92 | 98.8
2.2 | 98.375 | 98.6 | 98.75 | 98.45
2.3 | 98.1333 | 98.3 | 98.35 | 98.2
2.4 | 97.6 | 98 | 98.0625 | 97.8
2.5 | 97.0429 | 97.5 | 97.725 | 97.4
2.6 | 96.65 | 97.175 | 97.35 | 97.2
2.7 | 96.3333 | 96.925 | 96.85 | 96.95
2.8 | 95.8 | 96.55 | 96.5667 | 96.6
2.9 | 95.3667 | 96.05 | 96.2429 | 96.275
3 | 94.9 | 95.75 | 96 | 95.85
3.1 | 94.525 | 95.2667 | 95.675 | 95.575
3.2 | 94.1 | 94.8667 | 95.3 | 95.25
3.3 | 93.725 | 94.58 | 94.9 | 94.85
3.4 | 93.15 | 94.2333 | 94.5 | 94.525
3.5 | 92.68 | 93.9 | 94.25 | 94.35
3.6 | 92.475 | 93.4667 | 93.925 | 93.98
3.7 | 91.9 | 93.1 | 93.675 | 93.7
3.8 | 91.5 | 92.7 | 93.3375 | 93.3
3.9 | 91.05 | 92.3333 | 92.9625 | 92.9
4 | 90.625 | 92 | 92.6 | 92.65
4.1 | 90.08 | 91.6 | 92.18 | 92.3
4.2 | 89.75 | 91.2 | 91.8333 | 91.9
4.3 | 89.2 | 90.84 | 91.6 | 91.5
4.4 | 88.8 | 90.4667 | 91.35 | 91.2
4.5 | 88.2667 | 90.08 | 90.9 | 91
4.6 | 87.8667 | 89.76 | 90.5429 | 90.6
4.7 | 87.5714 | 89.4333 | 90.25 | 90.3
4.8 | 87.1 | 88.975 | 90.0333 | 89.95
4.9 | 86.6 | 88.6333 | 89.65 | 89.5
5 | 86.2 | 88.24 | 89.25 | 89.275
5.1 | 85.75 | 87.8 | 88.74 | 89.05
5.2 | 85.3 | 87.4 | 88.4 | 88.75
5.3 | 84.82 | 87.15 | 88.2 | 88.425
5.4 | 84.45 | 86.8 | 88 | 88
5.5 | 84.05 | 86.45 | 87.6 | 87.65
5.6 | 83.55 | 86.05 | 87.25 | 87.3
5.7 | 83.2 | 85.7 | 86.9 | 87.1
5.8 | 82.725 | 85.35 | 86.6333 | 86.825
5.9 | 82.3667 | 85 | 86.35 | 86.4333
6 | 81.9 | 84.75 | 85.95 | 86.1
6.1 | 81.55 | 84.28 | 85.7 | 85.86
6.2 | 81.0667 | 84 | 85.3 | 85.525
6.3 | 80.6375 | 83.6 | 85.0222 | 85.1
6.4 | 80.2556 | 83.25 | 84.65 | 84.8333
6.5 | 79.85 | 82.8 | 84.225 | 84.3667
6.6 | 79.55 | 82.45 | 83.88 | 84.1
6.7 | 79 | 82.0375 | 83.65 | 83.675
6.8 | 78.76 | 81.7625 | 83.3667 | 83.425
6.9 | 78.2429 | 81.425 | 83.1 | 83.2
7 | 77.74 | 80.9625 | 82.65 | 82.925
7.1 | 77.475 | 80.55 | 82.25 | 82.5667
7.2 | 77.125 | 80.1889 | 82 | 82.3167
7.3 | 76.55 | 79.95 | 81.8333 | 82.02
7.4 | 76.2 | 79.7 | 81.4667 | 81.475
7.5 | 75.8 | 79.2 | 80.95 | 81.325
7.6 | 75.35 | 78.9 | 80.7 | 80.975
7.7 | 74.9 | 78.4667 | 80.3333 | 80.65
7.8 | 74.55 | 78.225 | 79.9 | 80.28
7.9 | 74.04 | 77.9125 | 79.7 | 80.1
8 | 73.875 | 77.475 | 79.4 | 79.9
8.1 | 73.45 | 77.18 | 79.2 | 79.5
8.2 | 72.96 | 76.9 | 78.95 | 79.2
8.3 | 72.65 | 76.5 | 78.6667 | 78.875
8.4 | 72.2 | 76.1 | 78.2 | 78.6
8.5 | 71.8 | 75.7 | 77.9333 | 78.32
8.6 | 71.4 | 75.425 | 77.575 | 77.95
8.7 | 71.05 | 74.975 | 77.325 | 77.675
8.8 | 70.725 | 74.625 | 77.0333 | 77.35
8.9 | 70.36 | 74.4 | 76.725 | 77.175
9 | 69.9 | 74.1 | 76.475 | 76.85
9.1 | 69.6 | 73.7571 | 76.2 | 76.38
9.2 | 69.2 | 73.45 | 75.8 | 76.16
9.3 | 68.8 | 73.1 | 75.3 | 75.7333
9.4 | 68.3333 | 72.675 | 75.0667 | 75.55
9.5 | 68.025 | 72.3333 | 74.82 | 75.3
9.6 | 67.56 | 72.075 | 74.35 | 75.05
9.7 | 67.2 | 71.7875 | 74.2 | 74.7
9.8 | 66.9 | 71.35 | 73.86 | 74.36
9.9 | 66.45 | 70.9375 | 73.625 | 74.12
10 | 66.22 | 70.75 | 73.2 | 73.8333
10.1 | 65.775 | 70.5 | 72.9 | 73.425
10.2 | 65.425 | 70.12 | 72.5667 | 73.15
10.3 | 65.05 | 69.75 | 72.4 | 72.825
10.4 | 64.625 | 69.45 | 72.175 | 72.575
10.5 | 64.375 | 69.125 | 71.85 | 72.325
10.6 | 64.0333 | 68.75 | 71.5 | 71.96
10.7 | 63.825 | 68.36 | 71.15 | 71.6
10.8 | 63.25 | 68.1 | 70.8875 | 71.34
10.9 | 63 | 67.8 | 70.6375 | 71
11 | 62.65 | 67.46 | 70.3375 | 70.8
11.1 | 62.25 | 67.15 | 69.975 | 70.6
11.2 | 62.1 | 66.875 | 69.65 | 70.35
11.3 | 61.75 | 66.525 | 69.275 | 70
11.4 | 61.3 | 66.4 | 69.1125 | 69.75
11.5 | 61 | 65.95 | 68.7714 | 69.5
11.6 | 60.7333 | 65.58 | 68.5 | 69.15
11.7 | 60.4 | 65.3 | 68.2 | 68.8
11.8 | 59.9333 | 65.1 | 68 | 68.6
11.9 | 59.7 | 64.65 | 67.8 | 68.3
12 | 59.35 | 64.3 | 67.475 | 67.9
12.1 | 58.9 | 64.0333 | 67.15 | 67.6
12.2 | 58.6333 | 63.725 | 66.85 | 67.2667
12.3 | 58.375 | 63.475 | 66.5 | 67.025
12.4 | 58.05 | 63.225 | 66.25 | 66.775
12.5 | 57.775 | 63.05 | 65.95 | 66.4667
12.6 | 57.35 | 62.74 | 65.7 | 66.22
12.7 | 57.05 | 62.4 | 65.45 | 66.075
12.8 | 56.725 | 62.0429 | 65.2 | 65.725
12.9 | 56.4 | 61.775 | 64.82 | 65.4667
13 | 56.1667 | 61.4333 | 64.5 | 65.2375
13.1 | 55.95 | 61.14 | 64.35 | 65.05
13.2 | 55.6333 | 60.8667 | 64.125 | 64.675
13.3 | 55.2667 | 60.6 | 63.8 | 64.5125
13.4 | 54.975 | 60.3 | 63.625 | 64.2
13.5 | 54.8 | 59.9 | 63.3 | 63.95
13.6 | 54.3 | 59.72 | 62.96 | 63.6
13.7 | 54.125 | 59.375 | 62.75 | 63.35
13.8 | 53.775 | 59.025 | 62.5 | 63.025
13.9 | 53.625 | 58.88 | 62.1 | 62.875
14 | 53.25 | 58.6 | 61.8 | 62.525
14.1 | 53.025 | 58.4 | 61.575 | 62.3667
14.2 | 52.7 | 58.05 | 61.325 | 62
14.3 | 52.4 | 57.7 | 61.075 | 61.7667
14.4 | 51.9 | 57.45 | 60.7375 | 61.5
14.5 | 51.65 | 57.1 | 60.5625 | 61.24
14.6 | 51.35 | 56.775 | 60.1875 | 61
14.7 | 51.175 | 56.55 | 59.9 | 60.8
14.8 | 50.85 | 56.2 | 59.7 | 60.6
14.9 | 50.6 | 55.95 | 59.45 | 60.375
15 | 50.2333 | 55.7 | 59.2 | 60.02
15.1 | 50.0667 | 55.45 | 58.96 | 59.8
15.2 | 49.8 | 55.1 | 58.725 | 59.55
15.3 | 49.45 | 54.85 | 58.45 | 59.175
15.4 | 49.2 | 54.7 | 58.2 | 59
15.5 | 48.95 | 54.3857 | 57.975 | 58.775
15.6 | 48.66 | 54.1 | 57.725 | 58.425
15.7 | 48.425 | 53.875 | 57.48 | 58.22
15.8 | 48.14 | 53.7125 | 57.25 | 58.025
15.9 | 47.875 | 53.3667 | 57 | 57.8
16 | 47.625 | 52.975 | 56.75 | 57.625
16.1 | 47.4 | 52.7667 | 56.5 | 57.3
16.2 | 47.05 | 52.5 | 56.3 | 57.05
16.3 | 46.8 | 52.25 | 56 | 56.8
16.4 | 46.55 | 52 | 55.8 | 56.65
16.5 | 46.3 | 51.75 | 55.6 | 56.3
16.6 | 46.06 | 51.5 | 55.275 | 56.1
16.7 | 45.8 | 51.25 | 55.125 | 55.95
16.8 | 45.575 | 50.9 | 54.775 | 55.675
16.9 | 45.3143 | 50.76 | 54.525 | 55.425
17 | 45.05 | 50.6 | 54.28 | 55.075
17.1 | 44.8 | 50.375 | 54.1 | 54.9
17.2 | 44.575 | 50.2 | 53.9 | 54.8
17.3 | 44.325 | 49.825 | 53.6 | 54.5167
17.4 | 44.075 | 49.675 | 53.4 | 54.2
17.5 | 43.8833 | 49.425 | 53.15 | 54.05
17.6 | 43.5857 | 49.2 | 52.9 | 53.8
17.7 | 43.3 | 48.85 | 52.65 | 53.375
17.8 | 43.15 | 48.5 | 52.4 | 53.2
17.9 | 42.9 | 48.35 | 52.15 | 52.95
18 | 42.72 | 48.1667 | 52 | 52.775
18.1 | 42.5 | 47.9 | 51.75 | 52.525
18.2 | 42.275 | 47.6 | 51.5 | 52.375
18.3 | 41.95 | 47.4 | 51.3 | 52.1
18.4 | 41.8 | 47.15 | 51.025 | 51.9
18.5 | 41.44 | 46.9 | 50.875 | 51.75
18.6 | 41.3 | 46.675 | 50.625 | 51.6
18.7 | 41.05 | 46.425 | 50.5 | 51.35
18.8 | 40.8 | 46.275 | 50.25 | 51.1
18.9 | 40.55 | 46.025 | 50 | 50.9
19 | 40.3 | 45.78 | 49.76 | 50.7
19.1 | 40.2 | 45.55 | 49.525 | 50.325
19.2 | 40.025 | 45.3 | 49.275 | 50.1
19.3 | 39.7333 | 45.1 | 49.1 | 49.95
19.4 | 39.5 | 44.85 | 48.9 | 49.74
19.5 | 39.325 | 44.6 | 48.75 | 49.48
19.6 | 39.05 | 44.45 | 48.5 | 49.24
19.7 | 38.8 | 44.225 | 48.3 | 49.1111
19.8 | 38.6 | 44.075 | 48.0857 | 48.98
19.9 | 38.425 | 43.825 | 47.85 | 48.725
20 | 38.3 | 43.5 | 47.6 | 48.4333
20.1 | 38.1 | 43.4 | 47.35 | 48.25
20.2 | 37.8429 | 43.1333 | 47.1 | 48.05
20.3 | 37.675 | 43.025 | 47 | 47.85
20.4 | 37.45 | 42.8 | 46.8167 | 47.625
20.5 | 37.275 | 42.55 | 46.475 | 47.475
20.6 | 37.1 | 42.325 | 46.3 | 47.32
20.7 | 36.92 | 42.1333 | 46.1667 | 47.175
20.8 | 36.7 | 41.825 | 46 | 46.9
20.9 | 36.425 | 41.7 | 45.7 | 46.675
21 | 36.2 | 41.45 | 45.4 | 46.35
21.1 | 36.025 | 41.3 | 45.3 | 46.2
21.2 | 35.8 | 41.05 | 45.1 | 46.1
21.3 | 35.625 | 40.9 | 44.875 | 45.8
21.4 | 35.4 | 40.6571 | 44.7 | 45.7
21.5 | 35.3 | 40.5 | 44.375 | 45.5
21.6 | 35.075 | 40.2556 | 44.2 | 45.36
21.7 | 34.925 | 40.0625 | 44 | 45.125
21.8 | 34.7 | 39.9375 | 43.85 | 44.9
21.9 | 34.55 | 39.8125 | 43.7 | 44.725
22 | 34.4 | 39.4714 | 43.5 | 44.475
22.1 | 34.15 | 39.25 | 43.4 | 44.34
22.2 | 34 | 39.1 | 43.05 | 44.06
22.3 | 33.8 | 38.875 | 42.925 | 43.85
22.4 | 33.625 | 38.72 | 42.75 | 43.7
22.5 | 33.4 | 38.475 | 42.525 | 43.5
22.6 | 33.3 | 38.35 | 42.375 | 43.3
22.7 | 33.075 | 38.1 | 42.225 | 43.1
22.8 | 32.925 | 37.96 | 41.98 | 42.9778
22.9 | 32.775 | 37.8 | 41.8 | 42.825
23 | 32.525 | 37.675 | 41.625 | 42.575
23.1 | 32.4 | 37.425 | 41.4 | 42.4125
23.2 | 32.15 | 37.2 | 41.2 | 42.175
23.3 | 32.1 | 37.025 | 41 | 41.95
23.4 | 31.85 | 36.88 | 40.9 | 41.775
23.5 | 31.65 | 36.7 | 40.7 | 41.525
23.6 | 31.6 | 36.45 | 40.5 | 41.4
23.7 | 31.36 | 36.3 | 40.3 | 41.3
23.8 | 31.225 | 36.2 | 40.05 | 41.1
23.9 | 31.075 | 35.95 | 39.9 | 41
24 | 30.8 | 35.72 | 39.8 | 40.75
24.1 | 30.6 | 35.575 | 39.6 | 40.6
24.2 | 30.4 | 35.425 | 39.375 | 40.4
24.3 | 30.25 | 35.275 | 39.225 | 40.2
24.4 | 30.1 | 35.1 | 38.975 | 40.05
24.5 | 29.975 | 34.9 | 38.825 | 39.8
24.6 | 29.8 | 34.75 | 38.675 | 39.7
24.7 | 29.625 | 34.55 | 38.5333 | 39.575
24.8 | 29.5333 | 34.38 | 38.375 | 39.325
24.9 | 29.4 | 34.25 | 38.2 | 39.175
25 | 29.2 | 34 | 38 | 38.925
25.1 | 29.075 | 33.8 | 37.88 | 38.775
25.2 | 28.9429 | 33.7 | 37.7 | 38.625
25.3 | 28.725 | 33.6 | 37.45 | 38.5
25.4 | 28.6 | 33.375 | 37.3 | 38.35
25.5 | 28.425 | 33.225 | 37.15 | 38.2
25.6 | 28.175 | 33.075 | 37.02 | 37.9667
25.7 | 28.025 | 32.925 | 36.8 | 37.75
25.8 | 27.975 | 32.675 | 36.7 | 37.6333
25.9 | 27.825 | 32.54 | 36.575 | 37.5
26 | 27.68 | 32.35 | 36.325 | 37.275
26.1 | 27.5 | 32.2 | 36.2 | 37.0625
26.2 | 27.4 | 32 | 36.05 | 36.9375
26.3 | 27.2 | 31.925 | 35.8667 | 36.725
26.4 | 27.1 | 31.75 | 35.7 | 36.55
26.5 | 27 | 31.6 | 35.5 | 36.4125
26.6 | 26.8667 | 31.4 | 35.375 | 36.3
26.7 | 26.625 | 31.3 | 35.225 | 36.2
26.8 | 26.4333 | 31.175 | 34.975 | 36
26.9 | 26.3 | 31 | 34.825 | 35.9
27 | 26.2 | 30.875 | 34.7 | 35.6833
27.1 | 26.1 | 30.625 | 34.6 | 35.5
27.2 | 25.9 | 30.475 | 34.475 | 35.38
27.3 | 25.8 | 30.325 | 34.225 | 35.25
27.4 | 25.7 | 30.2 | 34.0333 | 35.04
27.5 | 25.6 | 30.025 | 33.9833 | 34.9
27.6 | 25.475 | 29.88 | 33.775 | 34.7667
27.7 | 25.2625 | 29.75 | 33.6625 | 34.5667
27.8 | 25.1375 | 29.6 | 33.475 | 34.425
27.9 | 24.98 | 29.425 | 33.3125 | 34.3
28 | 24.9 | 29.2875 | 33.175 | 34.1
28.1 | 24.8 | 29.14 | 32.9625 | 34
28.2 | 24.65 | 29 | 32.8375 | 33.85
28.3 | 24.48 | 28.9 | 32.7 | 33.7
28.4 | 24.3 | 28.65 | 32.5333 | 33.6
28.5 | 24.2 | 28.5 | 32.4 | 33.4
28.6 | 24.06 | 28.4 | 32.3 | 33.25
28.7 | 24 | 28.225 | 32.1667 | 33.1
28.8 | 23.9 | 28.175 | 32 | 32.95
28.9 | 23.7 | 28 | 31.8333 | 32.8
29 | 23.55 | 27.9 | 31.68 | 32.675
29.1 | 23.4 | 27.75 | 31.55 | 32.5
29.2 | 23.3 | 27.6 | 31.4 | 32.3667
29.3 | 23.2 | 27.5 | 31.35 | 32.2667
29.4 | 23.1 | 27.4 | 31.1333 | 32.1
29.5 | 23 | 27.275 | 31 | 31.975
29.6 | 22.9 | 27.025 | 30.85 | 31.8
29.7 | 22.8 | 26.9 | 30.7 | 31.64
29.8 | 22.6333 | 26.8143 | 30.5333 | 31.5
29.9 | 22.55 | 26.675 | 30.4 | 31.4
30 | 22.4 | 26.525 | 30.275 | 31.2667

## Wedge 45 Lateral Profiles
Lateral Profiles
Wedge 45
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.5 cm FieldSize:4x4 cm2 | Depth:1.5 cm FieldSize:10x10 cm2 | Depth:1.5 cm FieldSize:20x20 cm2 | Depth:1.5 cm FieldSize:40x20 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:40x20 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:40x20 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:40x20 cm2
-17.9 | [18 empty] | 16.5167 | 20.4167
-17.8 | [18 empty] | 16.8 | 20.5833
-17.7 | [18 empty] | 17.0333 | 20.8
-17.6 | [18 empty] | 17.22 | 21.1
-17.5 | [18 empty] | 17.4 | 21.4333
-17.4 | [18 empty] | 17.76 | 21.6333
-17.3 | [18 empty] | 17.95 | 21.9
-17.2 | [18 empty] | 18.3 | 22.0833
-17.1 | [18 empty] | 18.44 | 22.5
-17 | [18 empty] | 18.8 | 22.78
-16.9 | [14 empty] | 13.2694 | 16.7 |  |  | 19.05 | 23.04
-16.8 | [14 empty] | 13.4827 | 16.86 |  |  | 19.34 | 23.24
-16.7 | [14 empty] | 13.7027 | 17.06 |  |  | 19.68 | 23.525
-16.6 | [14 empty] | 13.8695 | 17.3333 |  |  | 19.9429 | 23.82
-16.5 | [14 empty] | 14.0628 | 17.55 |  |  | 20.2 | 24.16
-16.4 | [14 empty] | 14.3529 | 17.76 |  |  | 20.46 | 24.4333
-16.3 | [14 empty] | 14.5779 | 18.0167 |  |  | 20.82 | 24.64
-16.2 | [14 empty] | 14.7363 | 18.1833 |  |  | 21.125 | 25.05
-16.1 | [14 empty] | 14.983 | 18.5 |  |  | 21.38 | 25.34
-16 | [14 empty] | 15.228 | 18.8167 |  |  | 21.65 | 25.6
-15.9 | [10 empty] | 9.525 | 12.34 |  |  | 15.5231 | 19.025 |  |  | 22.08 | 26
-15.8 | [10 empty] | 9.625 | 12.5333 |  |  | 15.7281 | 19.25 |  |  | 22.4 | 26.3
-15.7 | [10 empty] | 9.8 | 12.7143 |  |  | 15.9432 | 19.5333 |  |  | 22.7167 | 26.74
-15.6 | [10 empty] | 10 | 12.9333 |  |  | 16.3533 | 19.8429 |  |  | 22.95 | 27.05
-15.5 | [10 empty] | 10.1667 | 13.06 |  |  | 16.5783 | 20.1 |  |  | 23.3667 | 27.36
-15.4 | [6 empty] | 7.6 | 10.3542 |  |  | 10.3333 | 13.3 |  |  | 16.8034 | 20.3667 |  |  | 23.7 | 27.65
-15.3 | [6 empty] | 7.74 | 10.4042 |  |  | 10.5 | 13.48 |  |  | 17.2034 | 20.7 |  |  | 24.1 | 28.04
-15.2 | [6 empty] | 7.825 | 10.6042 |  |  | 10.65 | 13.65 |  |  | 17.5035 | 20.9667 |  |  | 24.6 | 28.4333
-15.1 |  |  | 6.85 | 9.5 |  |  | 7.95 | 10.721 |  |  | 10.85 | 13.8 |  |  | 17.7435 | 21.2 |  |  | 24.925 | 28.9
-15 |  |  | 6.975 | 9.66 |  |  | 8.075 | 10.9044 |  |  | 11.0167 | 14 |  |  | 18.0286 | 21.55 |  |  | 25.2333 | 29.34
-14.9 |  |  | 7.1 | 9.7 |  |  | 8.2 | 10.9544 |  |  | 11.1833 | 14.18 |  |  | 18.3637 | 21.9 |  |  | 25.72 | 29.7
-14.8 |  |  | 7.2 | 9.8 |  |  | 8.4 | 11.1378 |  |  | 11.3667 | 14.38 |  |  | 18.6037 | 22.1667 |  |  | 26.14 | 29.98
-14.7 |  |  | 7.3 | 9.98 |  |  | 8.5 | 11.3045 |  |  | 11.58 | 14.625 |  |  | 18.9705 | 22.44 |  |  | 26.55 | 30.45
-14.6 |  |  | 7.4 | 10.1 |  |  | 8.6 | 11.4046 |  |  | 11.7333 | 14.85 |  |  | 19.3539 | 22.9 |  |  | 26.975 | 30.95
-14.5 |  |  | 7.5333 | 10.2333 |  |  | 8.8 | 11.5046 |  |  | 12 | 15 |  |  | 19.6539 | 23.2333 |  |  | 27.44 | 31.3333
-14.4 |  |  | 7.6286 | 10.4 |  |  | 8.8833 | 11.6618 |  |  | 12.225 | 15.2833 |  |  | 20.004 | 23.5667 |  |  | 27.95 | 31.76
-14.3 |  |  | 7.8 | 10.5 |  |  | 9.1 | 11.8381 |  |  | 12.46 | 15.54 |  |  | 20.4041 | 23.9286 |  |  | 28.6167 | 32.3
-14.2 |  |  | 7.8333 | 10.625 |  |  | 9.25 | 12.0298 |  |  | 12.7143 | 15.8 |  |  | 20.7841 | 24.3 |  |  | 29.15 | 32.9
-14.1 |  |  | 7.9167 | 10.72 |  |  | 9.3333 | 12.2249 |  |  | 12.95 | 16.05 |  |  | 21.2709 | 24.66 |  |  | 29.7 | 33.5
-14 |  |  | 8.0833 | 10.9 |  |  | 9.52 | 12.3249 |  |  | 13.2 | 16.275 |  |  | 21.6293 | 25.16 |  |  | 30.36 | 34.15
-13.9 |  |  | 8.2 | 11.04 |  |  | 9.7 | 12.5193 |  |  | 13.54 | 16.4833 |  |  | 21.9544 | 25.6 |  |  | 31.1 | 34.9
-13.8 |  |  | 8.35 | 11.15 |  |  | 9.8333 | 12.7651 |  |  | 13.75 | 16.8 |  |  | 22.3845 | 25.9333 |  |  | 32.3 | 36.1
-13.7 |  |  | 8.5 | 11.35 |  |  | 10.1 | 12.9719 |  |  | 14.04 | 17.12 |  |  | 22.8903 | 26.28 |  |  | 33.5 | 37.24
-13.6 |  |  | 8.6333 | 11.45 |  |  | 10.2 | 13.1386 |  |  | 14.275 | 17.34 |  |  | 23.3547 | 26.75 |  |  | 35.48 | 39
-13.5 |  |  | 8.725 | 11.6 |  |  | 10.425 | 13.4054 |  |  | 14.6 | 17.72 |  |  | 23.7648 | 27.22 |  |  | 38.2 | 41.6
-13.4 |  |  | 8.85 | 11.7 |  |  | 10.65 | 13.5254 |  |  | 14.9 | 18 |  |  | 24.3049 | 27.8 |  |  | 42.1333 | 45.18
-13.3 |  |  | 9 | 11.9 |  |  | 10.875 | 13.7555 |  |  | 15.22 | 18.3 |  |  | 24.785 | 28.3 |  |  | 47.8833 | 50.5
-13.2 |  |  | 9.2 | 12.02 |  |  | 11.025 | 14.0056 |  |  | 15.62 | 18.7 |  |  | 25.5051 | 28.72 |  |  | 56.05 | 57.9
-13.1 |  |  | 9.35 | 12.2 |  |  | 11.28 | 14.2257 |  |  | 15.9 | 19.04 |  |  | 26.0852 | 29.4 |  |  | 65.95 | 66.8
-13 |  |  | 9.54 | 12.4 |  |  | 11.55 | 14.3658 |  |  | 16.325 | 19.36 |  |  | 26.8653 | 30.15 |  |  | 76.32 | 77.1833
-12.9 |  |  | 9.66 | 12.575 |  |  | 11.775 | 14.6259 |  |  | 16.65 | 19.75 |  |  | 27.6056 | 31 |  |  | 87.06 | 86.6833
-12.8 |  |  | 9.78 | 12.7 |  |  | 12 | 14.906 |  |  | 17.0714 | 20.16 |  |  | 28.7391 | 32.05 |  |  | 96.5333 | 95.6
-12.7 |  |  | 9.9833 | 12.925 |  |  | 12.2 | 15.2228 |  |  | 17.5 | 20.6 |  |  | 30.146 | 33.42 |  |  | 104.9667 | 103.6
-12.6 |  |  | 10.2 | 13.075 |  |  | 12.5333 | 15.5062 |  |  | 17.9 | 20.95 |  |  | 32.1064 | 35 |  |  | 111.4 | 109.4667
-12.5 |  |  | 10.3333 | 13.3 |  |  | 12.75 | 15.7313 |  |  | 18.36 | 21.46 |  |  | 34.8736 | 37.7 |  |  | 115.5667 | 113.4667
-12.4 |  |  | 10.58 | 13.52 |  |  | 13.16 | 16.0564 |  |  | 18.84 | 21.9333 |  |  | 39.3079 | 42 |  |  | 118.2333 | 115.8
-12.3 |  |  | 10.8 | 13.625 |  |  | 13.4 | 16.3065 |  |  | 19.3 | 22.35 |  |  | 45.8925 | 48.3 |  |  | 120 | 117.3
-12.2 |  |  | 10.925 | 13.875 |  |  | 13.7 | 16.7067 |  |  | 19.8333 | 22.95 |  |  | 55.3311 | 56.82 |  |  | 121.1 | 118.4
-12.1 |  |  | 11.18 | 14.14 |  |  | 14.1 | 17.0468 |  |  | 20.6 | 23.54 |  |  | 66.7934 | 67.8 |  |  | 121.9667 | 119.2
-12 |  |  | 11.4 | 14.375 |  |  | 14.54 | 17.3736 |  |  | 21.34 | 24.3 |  |  | 78.0907 | 79.0167 |  |  | 122.4 | 119.65
-11.9 |  |  | 11.625 | 14.6 |  |  | 14.88 | 17.8471 |  |  | 22.1 | 25.1167 |  |  | 90.7181 | 90.92 |  |  | 122.76 | 120.05
-11.8 |  |  | 11.9 | 14.8333 |  |  | 15.3 | 18.1872 |  |  | 23.3 | 26.2 |  |  | 102.0205 | 101.9333 |  |  | 123.05 | 120.3
-11.7 |  |  | 12.15 | 15.1 |  |  | 15.72 | 18.7075 |  |  | 24.72 | 27.5 |  |  | 111.8024 | 110.15 |  |  | 123.225 | 120.5667
-11.6 |  |  | 12.375 | 15.34 |  |  | 16.32 | 19.2077 |  |  | 26.95 | 29.7 |  |  | 119.184 | 117.22 |  |  | 123.4 | 120.7
-11.5 |  |  | 12.725 | 15.675 |  |  | 16.9333 | 19.9079 |  |  | 30.2667 | 32.9 |  |  | 123.645 | 121.2333 |  |  | 123.48 | 120.8333
-11.4 |  |  | 12.975 | 16 |  |  | 17.8 | 20.6683 |  |  | 35.74 | 37.78 |  |  | 126.225 | 123.7 |  | 8.9 | 123.7 | 121
-11.3 |  |  | 13.4 | 16.35 |  |  | 18.8 | 21.6337 |  |  | 44.6 | 46.1 |  |  | 127.8658 | 125.2 |  | 9.1 | 123.775 | 121.0833
-11.2 |  |  | 13.85 | 16.85 |  |  | 20.38 | 23.2493 |  |  | 56.24 | 57.2 |  |  | 128.876 | 126.1667 |  | 9.3333 | 123.8 | 121.3
-11.1 |  |  | 14.4667 | 17.45 |  |  | 22.3 | 25.09 |  |  | 70.2 | 70.8 |  |  | 129.406 | 126.7333 |  | 9.5 | 123.9 | 121.4
-11 |  |  | 15.2 | 18.225 |  |  | 25.74 | 28.5314 |  |  | 86 | 85.7 |  |  | 129.876 | 127.22 |  | 9.7 | 124 | 121.4
-10.9 |  |  | 16.3 | 19.125 |  |  | 31.58 | 34.0936 |  |  | 99.74 | 98.8 |  | 6.9 | 130.1427 | 127.4 |  | 9.9 | 124.1 | 121.4
-10.8 |  |  | 18 | 20.8 |  |  | 41 | 43.2173 |  |  | 113.6333 | 112.125 |  | 6.9 | 130.326 | 127.6 |  | 10 | 124.02 | 121.5
-10.7 |  |  | 20.7 | 23.32 |  |  | 54.02 | 56.4226 |  |  | 124.525 | 122.75 |  | 7 | 130.3593 | 127.7833 |  | 10.1 | 124 | 121.5
-10.6 |  |  | 25.2 | 28.04 |  |  | 70.575 | 71.0884 |  |  | 131.85 | 129.74 |  | 7.1 | 130.476 | 127.8 |  | 10.2 | 124 | 121.5
-10.5 |  |  | 33.95 | 36.2 |  |  | 87 | 88.3353 |  |  | 136.1 | 133.6 |  | 7.3 | 130.586 | 127.9 |  | 10.45 | 123.975 | 121.5
-10.4 |  |  | 47.7 | 48.6 |  |  | 104.68 | 104.542 |  | 4.4 | 138.4167 | 136 |  | 7.44 | 130.626 | 128 |  | 10.6 | 123.9 | 121.5
-10.3 |  |  | 64.0667 | 65.62 |  |  | 119.88 | 118.6975 |  | 4.5 | 139.96 | 137.34 |  | 7.575 | 130.626 | 128 |  | 10.775 | 123.9 | 121.5
-10.2 |  |  | 82.375 | 83.1 |  | 3.4009 | 131.62 | 130.0018 |  | 4.6 | 140.5 | 137.9 |  | 7.7 | 130.626 | 128 |  | 10.9 | 123.9 | 121.5
-10.1 |  |  | 101.22 | 100.775 |  | 3.4009 | 139.5 | 137.7802 |  | 4.6 | 140.85 | 138.34 |  | 7.8667 | 130.526 | 128 |  | 11.2 | 123.7714 | 121.4167
-10 |  | 3 | 118.26 | 117.18 |  | 3.5009 | 144.22 | 141.7065 |  | 4.7 | 140.98 | 138.5 |  | 8.02 | 130.4427 | 128 |  | 11.4 | 123.7 | 121.375
-9.9 |  | 3.1 | 133.1 | 131.92 |  | 3.5009 | 146.5 | 144.258 |  | 4.8 | 141.1 | 138.5 |  | 8.1 | 130.346 | 128 |  | 11.575 | 123.65 | 121.3
-9.8 |  | 3.1 | 144.1 | 141.8 |  | 3.6009 | 147.62 | 145.158 |  | 4.9 | 141 | 138.5 |  | 8.3 | 130.2593 | 127.85 |  | 11.8 | 123.5833 | 121.25
-9.7 |  | 3.1 | 149.95 | 147.3 |  | 3.6009 | 148.1 | 145.633 |  | 5 | 141 | 138.44 |  | 8.45 | 130.146 | 127.8 |  | 11.95 | 123.475 | 121.1
-9.6 |  | 3.2 | 152.25 | 149.66 |  | 3.7009 | 148.3 | 145.758 |  | 5.1 | 140.84 | 138.24 |  | 8.5667 | 130.026 | 127.62 |  | 12.15 | 123.35 | 121.06
-9.5 |  | 3.2 | 153.425 | 150.5667 |  | 3.7009 | 148.3 | 145.758 |  | 5.2 | 140.58 | 138.125 |  | 8.8 | 129.9117 | 127.44 |  | 12.5 | 123.3 | 120.975
-9.4 |  | 3.3 | 153.64 | 150.9 |  | 3.8009 | 148.2 | 145.658 |  | 5.3 | 140.35 | 137.9167 |  | 8.95 | 129.746 | 127.25 |  | 12.7 | 123.2 | 120.88
-9.3 |  | 3.3 | 153.7 | 150.9 |  | 3.901 | 147.875 | 145.5247 |  | 5.4 | 140.1 | 137.675 |  | 9.1 | 129.486 | 127.06 |  | 12.95 | 123.1 | 120.75
-9.2 |  | 3.4 | 153.6 | 150.76 |  | 3.951 | 147.66 | 145.298 |  | 5.6 | 139.88 | 137.4667 |  | 9.28 | 129.326 | 126.96 |  | 13.225 | 123 | 120.7
-9.1 |  | 3.4 | 153.34 | 150.5333 |  | 4.001 | 147.25 | 144.898 |  | 5.7 | 139.6333 | 137.1667 |  | 9.525 | 129.201 | 126.8 |  | 13.45 | 122.8 | 120.6
-9 |  | 3.5 | 152.7 | 150.05 |  | 4.101 | 146.8667 | 144.633 |  | 5.8 | 139.3 | 136.8714 |  | 9.775 | 129.076 | 126.7 |  | 13.75 | 122.5333 | 120.425
-8.9 |  | 3.5 | 152.24 | 149.48 |  | 4.1344 | 146.5 | 144.208 |  | 5.9 | 139 | 136.5667 |  | 10 | 128.8593 | 126.6 |  | 14 | 122.35 | 120.3
-8.8 |  | 3.5667 | 151.5 | 148.775 |  | 4.2011 | 146 | 143.858 |  | 6.1 | 138.58 | 136.28 |  | 10.175 | 128.546 | 126.4333 |  | 14.2667 | 122.2 | 120.275
-8.7 |  | 3.7 | 150.85 | 148.3 |  | 4.3011 | 145.5 | 143.182 |  | 6.2 | 138.28 | 136.06 |  | 10.475 | 128.401 | 126.2 |  | 14.45 | 122.1 | 120.1
-8.6 |  | 3.7 | 150.15 | 147.5 |  | 4.4678 | 145 | 142.737 |  | 6.4 | 137.95 | 135.625 |  | 10.6667 | 128.076 | 125.9429 |  | 14.9333 | 121.9667 | 119.9
-8.5 |  | 3.8 | 149.52 | 146.975 |  | 4.6011 | 144.45 | 142.177 |  | 6.575 | 137.45 | 135.25 |  | 11 | 127.846 | 125.7 |  | 15.2 | 121.86 | 119.7
-8.4 |  | 3.9 | 148.875 | 146.475 |  | 4.6011 | 143.68 | 141.737 |  | 6.7 | 137 | 134.94 |  | 11.2 | 127.526 | 125.475 |  | 15.55 | 121.6833 | 119.5429
-8.3 |  | 3.9 | 148.1 | 145.6 |  | 4.7012 | 143.32 | 141.206 |  | 6.8333 | 136.7 | 134.4833 |  | 11.4333 | 127.1917 | 125.34 |  | 15.8 | 121.4 | 119.3833
-8.2 |  | 4.0333 | 147.425 | 144.95 |  | 4.8012 | 142.75 | 140.6893 |  | 7 | 136.15 | 134.2333 |  | 11.7 | 127.075 | 124.94 |  | 16.1 | 121.2 | 119.2167
-8.1 |  | 4.1 | 146.675 | 144.4 |  | 4.9012 | 142.175 | 140.106 |  | 7.2 | 135.7 | 133.76 |  | 11.95 | 126.775 | 124.76 |  | 16.525 | 120.95 | 119.125
-8 |  | 4.2 | 146 | 143.58 |  | 5.0346 | 141.5667 | 139.556 |  | 7.4 | 135.3 | 133.35 |  | 12.375 | 126.485 | 124.46 |  | 16.8333 | 120.7333 | 119
-7.9 |  | 4.2667 | 145.2 | 142.9 |  | 5.2013 | 140.92 | 139.056 |  | 7.625 | 134.8 | 132.9 |  | 12.6333 | 126.225 | 124.2714 |  | 17.15 | 120.46 | 118.7
-7.8 |  | 4.3 | 144.625 | 142.3 |  | 5.4013 | 140.28 | 138.43 |  | 7.875 | 134.44 | 132.6 |  | 12.9 | 125.845 | 124.0333 |  | 17.7 | 120.2 | 118.45
-7.7 |  | 4.4 | 143.9 | 141.7 |  | 5.5014 | 139.78 | 137.915 |  | 8.0333 | 133.96 | 132.2 |  | 13.3 | 125.6583 | 123.8 |  | 18.1667 | 120 | 118.22
-7.6 |  | 4.525 | 143.1 | 140.86 |  | 5.6014 | 139.3 | 137.455 |  | 8.3 | 133.5333 | 131.82 |  | 13.6333 | 125.375 | 123.5 | 4 | 18.6667 | 119.7333 | 118
-7.5 |  | 4.6 | 142.4 | 140.1 |  | 5.8014 | 138.55 | 136.6217 |  | 8.5 | 133.04 | 131.24 |  | 13.9333 | 124.965 | 123.175 | 3.9 | 19.2 | 119.6333 | 117.95
-7.4 |  | 4.75 | 141.66 | 139.725 |  | 6.0015 | 138 | 136.1744 |  | 8.7667 | 132.62 | 130.84 | 2.6 | 14.35 | 124.625 | 122.9333 | 3.7 | 19.95 | 119.46 | 117.9
-7.3 |  | 4.9333 | 140.92 | 138.975 |  | 6.1515 | 137.36 | 135.504 |  | 9.1 | 132.14 | 130.425 | 2.7 | 14.85 | 124.325 | 122.7 | 3.8 | 20.8 | 119.18 | 117.7333
-7.2 |  | 5 | 140.2 | 138.225 |  | 6.3516 | 136.75 | 134.914 |  | 9.375 | 131.7333 | 130 | 2.7 | 15.2 | 123.925 | 122.5 | 3.8333 | 21.95 | 118.9667 | 117.5
-7.1 |  | 5.15 | 139.52 | 137.575 | 1.2 | 6.5766 | 136.15 | 134.454 | 1.7 | 9.7 | 131.3 | 129.475 | 2.7 | 15.7667 | 123.65 | 122.075 | 3.9 | 23.5 | 118.7 | 117.32
-7 | 1.1 | 5.2667 | 138.74 | 137.04 | 1.2 | 6.7017 | 135.5 | 133.9938 | 1.7 | 10.025 | 130.7 | 129.05 | 2.7 | 16.2667 | 123.5 | 121.74 | 4.1 | 26.1 | 118.425 | 116.9714
-6.9 | 1.1 | 5.4 | 138.2 | 136.3 | 1.2 | 7.0018 | 134.78 | 133.193 | 1.7 | 10.25 | 130.15 | 128.5 | 2.8 | 17 | 123.175 | 121.5 | 4.1 | 29.0667 | 118.22 | 116.8167
-6.8 | 1.1 | 5.5333 | 137.3333 | 135.54 | 1.2 | 7.3018 | 134.2 | 132.653 | 1.7 | 10.6667 | 129.72 | 128.04 | 2.8 | 17.9 | 122.685 | 121.26 | 4.1 | 34.3 | 118.0333 | 116.6
-6.7 | 1.1 | 5.7 | 136.7 | 134.85 | 1.3 | 7.6019 | 133.78 | 132.003 | 1.8 | 11.1667 | 129.25 | 127.66 | 2.9 | 19.05 | 122.399 | 120.98 | 4.2 | 41.5667 | 117.8 | 116.3833
-6.6 | 1.1 | 5.9 | 136.2 | 134.24 | 1.3 | 7.8686 | 132.8833 | 131.603 | 1.8 | 11.5667 | 128.7 | 127.3 | 2.9 | 20.65 | 122.124 | 120.6833 | 4.2 | 50.1 | 117.52 | 116.1333
-6.5 | 1.1 | 6.1 | 135.24 | 133.65 | 1.3 | 8.102 | 132.54 | 130.952 | 1.8 | 12.2 | 128.275 | 126.8667 | 3 | 23.1333 | 121.824 | 120.35 | 4.3 | 60.95 | 117.24 | 115.94
-6.4 | 1.1 | 6.25 | 134.58 | 132.92 | 1.3 | 8.5021 | 131.76 | 130.452 | 1.9 | 12.8 | 127.72 | 126.4 | 3 | 26.8 | 121.524 | 120 | 4.35 | 70.15 | 116.85 | 115.7
-6.3 | 1.1 | 6.5 | 133.775 | 132.2333 | 1.4 | 8.9689 | 131.2 | 129.852 | 1.9 | 13.7 | 127.2 | 125.8 | 3.1 | 32.4 | 121.124 | 119.7714 | 4.5 | 80.075 | 116.725 | 115.475
-6.2 | 1.1 | 6.8 | 133.14 | 131.66 | 1.4 | 9.5524 | 130.48 | 129.252 | 2 | 14.8667 | 126.6 | 125.3 | 3.2 | 40.95 | 120.7573 | 119.44 | 4.5 | 88.8333 | 116.48 | 115.2667
-6.1 | 1.1 | 7.1 | 132.5 | 130.98 | 1.4 | 10.2026 | 129.92 | 128.7114 | 2 | 16.425 | 126.22 | 124.92 | 3.3 | 50.9 | 120.524 | 119.1 | 4.6 | 95.625 | 116.3 | 115.1143
-6 | 1.2 | 7.6 | 131.6 | 130.3 | 1.4 | 11.1028 | 129.375 | 128.0796 | 2 | 18.85 | 125.86 | 124.4 | 3.4 | 61.1667 | 120.164 | 118.74 | 4.7 | 100 | 116.08 | 114.9
-5.9 | 1.2 | 8.15 | 131.1 | 129.6 | 1.5 | 12.5532 | 128.8 | 127.5367 | 2.1 | 23 | 125.2667 | 123.95 | 3.5 | 72.3 | 119.804 | 118.3667 | 4.8 | 102.78 | 115.8 | 114.68
-5.8 | 1.2 | 9.0333 | 130.45 | 128.975 | 1.5 | 14.7037 | 128.15 | 127.051 | 2.1 | 29.5667 | 124.8667 | 123.36 | 3.5 | 83.2 | 119.474 | 118.1667 | 4.9 | 104.5 | 115.46 | 114.5
-5.7 | 1.2 | 10.4 | 129.85 | 128.4 | 1.5 | 17.7378 | 127.575 | 126.4758 | 2.2 | 38.7 | 124.3 | 123.0667 | 3.6 | 91.7333 | 119.124 | 117.95 | 5 | 105.6 | 115.25 | 114.26
-5.6 | 1.2 | 12.5 | 129.06 | 127.7 | 1.5 | 23.3392 | 126.9 | 125.79 | 2.2 | 50.375 | 123.64 | 122.64 | 3.8 | 98.7 | 118.724 | 117.4 | 5.2 | 106.7333 | 114.925 | 114.04
-5.5 | 1.3 | 16.3333 | 128.4 | 127.2 | 1.6 | 31.908 | 126.44 | 125.175 | 2.3 | 62.2 | 123.3667 | 122.2 | 3.9 | 103.3 | 118.384 | 117.16 | 5.3 | 107.05 | 114.7 | 113.8
-5.4 | 1.3 | 23.3 | 127.8 | 126.5333 | 1.6 | 44.6112 | 125.68 | 124.51 | 2.35 | 75.5 | 122.82 | 121.625 | 4 | 105.8 | 118.074 | 116.775 | 5.4 | 107.4 | 114.42 | 113.58
-5.3 | 1.3 | 33.1 | 127.1 | 125.9 | 1.6 | 56.6641 | 125.16 | 124.05 | 2.4 | 87.2 | 122.26 | 121.06 | 4.05 | 107.2 | 117.7573 | 116.5333 | 5.5 | 107.6 | 114.2 | 113.25
-5.2 | 1.4 | 46.5 | 126.45 | 125.16 | 1.7 | 70.4176 | 124.5714 | 123.549 | 2.5 | 96.6333 | 121.9 | 120.6 | 4.2 | 108.125 | 117.383 | 116.2 | 5.7 | 107.775 | 113.925 | 113.125
-5.1 | 1.4 | 61 | 125.8 | 124.6 | 1.7 | 84.121 | 123.9 | 122.949 | 2.6 | 103.5 | 121.38 | 120.18 | 4.2 | 108.775 | 117.073 | 115.7833 | 5.8 | 107.9333 | 113.64 | 112.8
-5 | 1.4 | 76.6333 | 125.14 | 124 | 1.8 | 95.824 | 123.4333 | 122.499 | 2.7 | 107.9 | 120.95 | 119.68 | 4.4 | 109.2 | 116.748 | 115.6 | 5.9 | 108 | 113.275 | 112.6
-4.9 | 1.5 | 90 | 124.56 | 123.4 | 1.8 | 104.076 | 123 | 121.899 | 2.7 | 109.9 | 120.4 | 119.32 | 4.45 | 109.4 | 116.273 | 115.2333 | 6.05 | 108 | 113.05 | 112.425
-4.8 | 1.5 | 101.6333 | 124 | 122.7 | 1.9 | 109.4775 | 122.3 | 121.3082 | 2.8 | 111.3 | 119.975 | 118.82 | 4.7 | 109.5 | 115.9897 | 114.98 | 6.2 | 108 | 112.8333 | 112.075
-4.7 | 1.5 | 109.8 | 123.26 | 122.2 | 1.9 | 112.628 | 121.825 | 120.748 | 2.9 | 111.975 | 119.46 | 118.32 | 4.8 | 109.5 | 115.723 | 114.6 | 6.4 | 108.1 | 112.625 | 111.9
-4.6 | 1.6 | 114.26 | 122.66 | 121.68 | 2 | 114.329 | 121.22 | 120.248 | 3 | 112.5 | 118.8333 | 117.92 | 4.95 | 109.5 | 115.223 | 114.3333 | 6.6 | 108.1 | 112.28 | 111.6333
-4.5 | 1.6 | 116.55 | 122.14 | 121.12 | 2.05 | 115.129 | 120.68 | 119.668 | 3.1 | 112.7 | 118.38 | 117.4857 | 5.1 | 109.6 | 114.848 | 114 | 6.65 | 108.1 | 112 | 111.4333
-4.4 | 1.6 | 117.6 | 121.4 | 120.5333 | 2.15 | 115.429 | 120.05 | 119.048 | 3.2 | 112.8 | 117.8667 | 117.15 | 5.25 | 109.5 | 114.523 | 113.6667 | 6.8 | 108.1 | 111.9 | 111.18
-4.3 | 1.7 | 117.8333 | 120.98 | 119.94 | 2.2 | 115.429 | 119.5 | 118.6137 | 3.3 | 112.8 | 117.46 | 116.6 | 5.4 | 109.5 | 114.1897 | 113.4429 | 7.1 | 108.1 | 111.64 | 110.9667
-4.2 | 1.75 | 117.9 | 120.36 | 119.36 | 2.3 | 115.429 | 118.85 | 118.167 | 3.5 | 112.7 | 117 | 116.18 | 5.6 | 109.4 | 113.7897 | 113.1333 | 7.3 | 107.9 | 111.4 | 110.6333
-4.1 | 1.8 | 117.9 | 119.82 | 118.78 | 2.4 | 115.329 | 118.58 | 117.597 | 3.6 | 112.55 | 116.625 | 115.7 | 5.8 | 109.3333 | 113.4563 | 112.7 | 7.4 | 107.86 | 111.0333 | 110.3
-4 | 1.9 | 117.7 | 119.16 | 118.24 | 2.5 | 115.204 | 118 | 117.097 | 3.8 | 112.425 | 116.1 | 115.35 | 6.05 | 109.3 | 113.023 | 112.275 | 7.6 | 107.8 | 110.775 | 110.0833
-3.9 | 1.9 | 117.4 | 118.475 | 117.58 | 2.6 | 114.954 | 117.44 | 116.597 | 4 | 112.2 | 115.68 | 114.9333 | 6.2 | 109.2 | 112.7897 | 112 | 7.9 | 107.8 | 110.525 | 109.9286
-3.8 | 2 | 117.05 | 117.94 | 117.24 | 2.75 | 114.6957 | 116.8667 | 115.996 | 4.2 | 112.05 | 115.14 | 114.45 | 6.4 | 109.12 | 112.4555 | 111.775 | 8.1667 | 107.6 | 110.3333 | 109.65
-3.7 | 2.1 | 116.7 | 117.275 | 116.62 | 2.8 | 114.229 | 116.4 | 115.586 | 4.3 | 111.8 | 114.66 | 114.08 | 6.65 | 108.9 | 112.162 | 111.48 | 8.5667 | 107.5333 | 110.05 | 109.44
-3.6 | 2.2 | 116.2 | 116.775 | 116.1 | 3 | 114.129 | 115.76 | 115.166 | 4.5 | 111.6 | 114.34 | 113.6 | 6.9 | 108.775 | 111.922 | 111.2 | 9.15 | 107.4 | 109.625 | 109.2333
-3.5 | 2.25 | 115.9 | 116.34 | 115.46 | 3.2 | 113.728 | 115.32 | 114.6793 | 4.8 | 111.5 | 113.98 | 113.2 | 7.3 | 108.6 | 111.522 | 110.7714 | 9.5 | 107.3 | 109.4667 | 109
-3.4 | 2.35 | 115.2667 | 115.75 | 114.94 | 3.4 | 113.528 | 114.95 | 114.196 | 5 | 111.275 | 113.475 | 112.78 | 7.8 | 108.5 | 111.222 | 110.5333 | 10.45 | 107.2 | 109.2167 | 108.68
-3.3 | 2.5 | 114.9333 | 115.32 | 114.5 | 3.6 | 112.8947 | 114.375 | 113.7457 | 5.3 | 111.075 | 112.9 | 112.5 | 8.3 | 108.3 | 110.8887 | 110.34 | 11.5 | 107.05 | 109 | 108.45
-3.2 | 2.6 | 114.5333 | 114.775 | 114.05 | 3.8 | 112.5947 | 113.94 | 113.295 | 5.8 | 110.7 | 112.6 | 112.05 | 8.95 | 108.2 | 110.522 | 110 | 12.4 | 106.9 | 108.775 | 108.1
-3.1 | 2.75 | 113.95 | 114.225 | 113.55 | 4.3 | 112.428 | 113.3 | 112.7117 | 6.15 | 110.4667 | 112.14 | 111.5833 | 9.9 | 108 | 110.222 | 109.6 | 14.5 | 106.9 | 108.46 | 107.94
-3 | 3 | 113.5333 | 113.75 | 113.12 | 4.55 | 112.1613 | 112.94 | 112.325 | 6.8 | 110.2 | 111.7 | 111.275 | 11.25 | 107.9 | 109.882 | 109.3333 | 17.8 | 106.7 | 108.2333 | 107.7
-2.9 | 3.4 | 113.15 | 113.16 | 112.62 | 5.25 | 111.828 | 112.52 | 111.905 | 7.7 | 110 | 111.3 | 110.78 | 13.3 | 107.72 | 109.522 | 109.0667 | 22.2 | 106.6 | 107.9 | 107.48
-2.8 | 3.9 | 112.7333 | 112.6667 | 112.12 | 6.2 | 111.228 | 111.9 | 111.495 | 9 | 109.725 | 110.9 | 110.5 | 16.45 | 107.5 | 109.172 | 108.75 | 28.5 | 106.5 | 107.6667 | 107.12
-2.7 | 4.75 | 112.35 | 112.2 | 111.65 | 7.4 | 110.828 | 111.55 | 111.064 | 11.2 | 109.4 | 110.6667 | 110.0571 | 21.75 | 107.3333 | 108.822 | 108.3333 | 38.1 | 106.3333 | 107.2667 | 106.9
-2.6 | 6.1 | 111.8667 | 111.7 | 111.25 | 9.5 | 110.528 | 111.1 | 110.594 | 14.3 | 109.2 | 110.25 | 109.8 | 29.15 | 107.2333 | 108.482 | 108.04 | 47.4333 | 106.075 | 107.0833 | 106.7167
-2.5 | 8.2 | 111.4 | 111.25 | 110.7 | 12.7 | 110.178 | 110.65 | 110.264 | 19.7 | 108.9 | 109.6667 | 109.36 | 38 | 106.9667 | 108.1506 | 107.8 | 58.4 | 105.825 | 106.8333 | 106.575
-2.4 | 12.2 | 110.9667 | 110.94 | 110.25 | 18.6 | 109.852 | 110.26 | 109.844 | 28.5 | 108.4667 | 109.34 | 108.88 | 49.9667 | 106.7 | 107.842 | 107.525 | 70 | 105.575 | 106.6 | 106.3714
-2.3 | 19 | 110.45 | 110.275 | 109.88 | 27.2 | 109.552 | 109.85 | 109.384 | 39.1 | 108.2 | 108.84 | 108.46 | 61.2 | 106.5 | 107.642 | 107.2 | 79.3 | 105.4 | 106.3833 | 106.1
-2.2 | 29.6 | 110.0333 | 109.8 | 109.38 | 37.7 | 109.027 | 109.34 | 108.8032 | 50.7333 | 107.8 | 108.5 | 108.1 | 73.1 | 106.15 | 107.346 | 106.98 | 87.6 | 105.2 | 106.1333 | 105.8333
-2.1 | 41.6 | 109.5333 | 109.2333 | 108.95 | 51.6 | 108.5937 | 108.84 | 108.418 | 63.45 | 107.55 | 108.2 | 107.6 | 83 | 105.925 | 106.971 | 106.62 | 93.9667 | 105.0333 | 105.9 | 105.5333
-2 | 55.1 | 109.1 | 108.875 | 108.475 | 64.4 | 108.227 | 108.4667 | 108.043 | 75 | 107.2 | 107.8 | 107.24 | 90.7 | 105.775 | 106.6877 | 106.22 | 97.1 | 104.825 | 105.65 | 105.3
-1.9 | 69.7 | 108.6333 | 108.38 | 108.12 | 75.6 | 107.827 | 108.1333 | 107.6763 | 85.2 | 106.9 | 107.24 | 107.025 | 96.1 | 105.6 | 106.421 | 105.9 | 99.45 | 104.6 | 105.3 | 105
-1.8 | 81.6 | 108.2333 | 107.94 | 107.6 | 86.5 | 107.627 | 107.7 | 107.343 | 92.75 | 106.5 | 106.88 | 106.6 | 98.8 | 105.3 | 106.021 | 105.7333 | 100.9 | 104.4 | 105.06 | 104.8333
-1.7 | 90.95 | 107.7333 | 107.5 | 107.2 | 94.2 | 107.0937 | 107.2 | 106.863 | 97.25 | 106.175 | 106.5 | 106.24 | 100.7 | 105 | 105.621 | 105.44 | 101.4 | 104.2333 | 104.7 | 104.6429
-1.6 | 97.7 | 107.25 | 107.15 | 106.76 | 98.35 | 106.627 | 106.7 | 106.4097 | 99.9 | 105.925 | 106.25 | 105.9 | 101.7 | 104.8 | 105.221 | 105.1 | 102.1 | 104.1 | 104.3333 | 104.35
-1.5 | 101.45 | 106.8 | 106.6 | 106.35 | 100.8 | 106.227 | 106.25 | 106.102 | 101.5 | 105.5667 | 105.8 | 105.5667 | 102.3 | 104.6 | 104.9877 | 104.74 | 102.3 | 103.925 | 104.06 | 104.125
-1.4 | 102.9 | 106.35 | 106.06 | 105.9 | 102.1 | 105.826 | 105.875 | 105.682 | 102.3 | 105.25 | 105.4 | 105.1 | 102.7 | 104.35 | 104.6543 | 104.475 | 102.4 | 103.675 | 103.7 | 103.825
-1.3 | 103.7 | 105.95 | 105.8 | 105.5 | 102.8 | 105.4593 | 105.5 | 105.192 | 102.7 | 104.9 | 105 | 104.76 | 102.9 | 104.075 | 104.321 | 104.15 | 102.3 | 103.28 | 103.55 | 103.54
-1.2 | 104 | 105.5 | 105.2 | 105 | 103 | 105.0927 | 105 | 104.792 | 102.7 | 104.4 | 104.66 | 104.35 | 102.9 | 103.725 | 103.946 | 103.7333 | 102.3 | 103.1 | 103.2667 | 103.1667
-1.1 | 104 | 105 | 104.7333 | 104.6 | 103.1 | 104.626 | 104.6 | 104.342 | 102.8 | 104.1 | 104.22 | 104 | 102.85 | 103.35 | 103.671 | 103.34 | 102.2 | 102.8667 | 103 | 102.95
-1 | 104 | 104.5 | 104.3 | 104.1 | 103 | 104.001 | 104.1333 | 103.992 | 102.7 | 103.725 | 103.72 | 103.6 | 102.8 | 102.95 | 103.3877 | 102.96 | 102.1 | 102.6 | 102.75 | 102.6333
-0.9 | 103.8 | 104 | 103.9 | 103.7 | 102.8 | 103.726 | 103.7333 | 103.591 | 102.6 | 103.4333 | 103.4667 | 103.3 | 102.6 | 102.55 | 103.0781 | 102.75 | 102 | 102.35 | 102.3667 | 102.3
-0.8 | 103.3 | 103.55 | 103.54 | 103.34 | 102.65 | 103.226 | 103.25 | 103.241 | 102.35 | 102.9667 | 103.0667 | 102.9857 | 102.3 | 102.4 | 102.646 | 102.35 | 101.9 | 102.1333 | 102.1167 | 102.06
-0.7 | 102.9 | 103.0667 | 103.02 | 102.9 | 102.4 | 102.926 | 102.9 | 102.821 | 102.1 | 102.6 | 102.7 | 102.42 | 102.1 | 102.15 | 102.24 | 102.1833 | 101.7 | 101.9333 | 101.8 | 101.775
-0.6 | 102.6 | 102.6333 | 102.6 | 102.325 | 102.15 | 102.2593 | 102.46 | 102.441 | 101.9 | 102.3 | 102.24 | 102.175 | 101.8 | 101.8 | 101.9 | 102 | 101.5 | 101.6667 | 101.5714 | 101.5333
-0.5 | 102.2 | 102.05 | 102.08 | 101.94 | 101.9 | 101.825 | 102.0667 | 101.941 | 101.7 | 101.8667 | 101.8 | 101.75 | 101.6 | 101.6 | 101.595 | 101.48 | 101.3 | 101.4 | 101.2667 | 101.275
-0.4 | 101.8 | 101.7 | 101.68 | 101.6667 | 101.45 | 101.325 | 101.68 | 101.641 | 101.4 | 101.5 | 101.44 | 101.3 | 101.3 | 101.2 | 101.22 | 101.2 | 101 | 101.1 | 100.9667 | 101.0667
-0.3 | 101.3 | 101.3 | 101.3 | 101.16 | 101.1 | 101.025 | 101.2 | 101.24 | 101 | 101.0667 | 101.06 | 100.875 | 101 | 100.9333 | 100.995 | 101 | 100.7667 | 100.85 | 100.7 | 100.8167
-0.2 | 100.9 | 100.875 | 100.82 | 100.8 | 100.7 | 100.7917 | 100.8 | 100.8067 | 100.7 | 100.75 | 100.66 | 100.62 | 100.7 | 100.6 | 100.74 | 100.7 | 100.5 | 100.5667 | 100.46 | 100.5
-0.1 | 100.6 | 100.5 | 100.36 | 100.42 | 100.35 | 100.4583 | 100.34 | 100.44 | 100.3 | 100.3 | 100.38 | 100.325 | 100.35 | 100.2333 | 100.37 | 100.3429 | 100.2 | 100.3 | 100.2 | 100.25
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.6 | 99.5667 | 99.55 | 99.6333 | 99.5 | 99.6499 | 99.7 | 99.6899 | 99.5 | 99.6 | 99.66 | 99.725 | 99.65 | 99.75 | 99.78 | 99.8 | 99.7 | 99.75 | 99.65 | 99.84
0.2 | 99.1 | 99.1667 | 99.2 | 99.35 | 99.1 | 99.0915 | 99.26 | 99.4398 | 99.1 | 99.2 | 99.3 | 99.325 | 99.4 | 99.4667 | 99.5699 | 99.55 | 99.3 | 99.425 | 99.38 | 99.5
0.3 | 98.8 | 98.8333 | 98.8667 | 98.875 | 98.8 | 98.8247 | 99 | 99.0729 | 98.8 | 98.95 | 98.8667 | 98.825 | 99.1 | 99.0667 | 99.1598 | 99.12 | 99 | 99.3 | 99.1 | 99.28
0.4 | 98.25 | 98.4 | 98.56 | 98.6333 | 98.4 | 98.5046 | 98.6 | 98.6895 | 98.5 | 98.6 | 98.55 | 98.6 | 98.6 | 98.675 | 98.7697 | 98.8 | 98.85 | 99.04 | 98.9 | 99
0.5 | 97.9 | 98 | 98.12 | 98.28 | 97.8 | 98.1245 | 98.1 | 98.3593 | 98 | 98.2667 | 98.24 | 98.4333 | 98.4 | 98.4 | 98.4447 | 98.6 | 98.45 | 98.95 | 98.6 | 98.7333
0.6 | 97.6 | 97.55 | 97.7 | 97.775 | 97.4 | 97.8245 | 97.82 | 97.9392 | 97.5 | 98 | 97.98 | 98.075 | 98.1 | 98.2 | 98.1863 | 98.3667 | 98.2 | 98.55 | 98.42 | 98.56
0.7 | 97.25 | 97.2 | 97.3667 | 97.48 | 97 | 97.3577 | 97.48 | 97.6391 | 97.3 | 97.725 | 97.62 | 97.7 | 97.65 | 97.9 | 97.8196 | 98.0333 | 97.9 | 98.2333 | 98.1 | 98.36
0.8 | 96.8 | 96.875 | 97.04 | 97.125 | 96.5 | 96.9242 | 97.08 | 97.389 | 96.75 | 97.35 | 97.36 | 97.3 | 97.3 | 97.55 | 97.6195 | 97.65 | 97.7 | 97.95 | 97.8 | 98.08
0.9 | 96.4 | 96.65 | 96.675 | 96.725 | 96 | 96.6242 | 96.9 | 96.8787 | 96.3 | 96.925 | 97.02 | 97.15 | 96.85 | 97.2 | 97.4195 | 97.4 | 97.1 | 97.6667 | 97.48 | 97.82
1 | 95.7667 | 96.2667 | 96.425 | 96.5 | 95.5 | 96.3241 | 96.46 | 96.6387 | 95.9 | 96.55 | 96.72 | 96.82 | 96.5 | 96.95 | 97.0594 | 97.1 | 96.7 | 97.4 | 97.26 | 97.6
1.1 | 95.1 | 95.9 | 96 | 96.175 | 94.85 | 95.8573 | 96.0667 | 96.3586 | 95.4 | 96.15 | 96.2667 | 96.46 | 95.9 | 96.55 | 96.786 | 96.9333 | 96.35 | 97.025 | 97.025 | 97.3333
1.2 | 94.3 | 95.55 | 95.65 | 95.82 | 93.9 | 95.4739 | 95.7 | 96.0184 | 94.5 | 95.7667 | 96.1 | 96.16 | 95.4 | 96.2667 | 96.5693 | 96.5 | 96 | 96.65 | 96.8333 | 97.0167
1.3 | 93.3 | 95.0667 | 95.325 | 95.54 | 93.1 | 95.2238 | 95.4 | 95.6383 | 93.8 | 95.425 | 95.7 | 95.85 | 95 | 96.0333 | 96.2592 | 96.25 | 95.5 | 96.425 | 96.46 | 96.78
1.4 | 91.8 | 94.7 | 95 | 95.2 | 91.7 | 94.8237 | 95.2 | 95.3381 | 92.75 | 95.125 | 95.4 | 95.5 | 94.4 | 95.7667 | 96.0192 | 96.1 | 95.1 | 96.225 | 96.1 | 96.52
1.5 | 89.8 | 94.3667 | 94.6 | 94.8667 | 90 | 94.4569 | 94.8667 | 94.918 | 91.5 | 94.825 | 94.975 | 95.24 | 93.4 | 95.275 | 95.8392 | 95.75 | 94.5 | 95.975 | 95.8 | 96.3333
1.6 | 86.1 | 94 | 94.2 | 94.6 | 87.3 | 94.1986 | 94.52 | 94.6879 | 89.7 | 94.35 | 94.74 | 94.8667 | 92.4 | 95 | 95.5191 | 95.425 | 93.7 | 95.5667 | 95.6333 | 96
1.7 | 79.6 | 93.5667 | 93.96 | 94.12 | 83.4 | 93.8735 | 94.225 | 94.3777 | 87.1 | 94 | 94.5 | 94.56 | 91.3 | 94.65 | 95.159 | 95.18 | 92.9 | 95.3 | 95.4667 | 95.76
1.8 | 71 | 93.25 | 93.56 | 93.88 | 76.4 | 93.3483 | 93.82 | 94.0626 | 83.4 | 93.7 | 94.2 | 94.2333 | 89.3 | 94.25 | 94.839 | 94.95 | 92.3 | 95.05 | 95.2 | 95.4
1.9 | 60.5 | 92.9 | 93.2667 | 93.55 | 67.5 | 92.9566 | 93.4 | 93.8175 | 77 | 93.3 | 93.8333 | 94 | 86.8 | 93.9 | 94.5439 | 94.64 | 90.45 | 94.7667 | 94.9429 | 95.1
2 | 48.9 | 92.5 | 93.0167 | 93.15 | 57.1 | 92.5731 | 93.075 | 93.4374 | 68.2 | 92.8667 | 93.36 | 93.7667 | 82.9 | 93.5 | 94.1588 | 94.26 | 89.05 | 94.5 | 94.68 | 94.8667
2.1 | 37.3 | 92.1 | 92.54 | 92.8 | 45.75 | 92.1897 | 92.65 | 93.0372 | 57.6 | 92.5 | 93.05 | 93.3 | 77 | 93.2333 | 93.8521 | 94.08 | 86.9 | 93.95 | 94.3857 | 94.64
2.2 | 26.2 | 91.7667 | 92.2 | 92.38 | 34.8 | 91.8897 | 92.425 | 92.6704 | 47.1 | 92.1333 | 92.78 | 93.0333 | 68.7 | 92.9 | 93.552 | 93.875 | 83.05 | 93.6 | 94.1667 | 94.42
2.3 | 17 | 91.2 | 91.85 | 91.98 | 24.9 | 91.4229 | 92 | 92.2702 | 36.2 | 91.7333 | 92.38 | 92.6 | 58.8 | 92.5 | 93.2436 | 93.5333 | 77.3 | 93.35 | 93.85 | 94.225
2.4 | 10.6 | 90.8 | 91.46 | 91.7 | 17.15 | 91.0561 | 91.7 | 91.8968 | 27.05 | 91.2667 | 91.9667 | 92.3667 | 48.95 | 92.1 | 92.9519 | 93.14 | 70.5 | 93.04 | 93.6 | 94.025
2.5 | 7.4 | 90.4 | 91.08 | 91.425 | 11.7 | 90.5726 | 91.175 | 91.6567 | 19.1 | 90.95 | 91.72 | 92 | 38.95 | 91.6667 | 92.6585 | 92.8667 | 62.3 | 92.65 | 93.3 | 93.72
2.6 | 5.4 | 89.95 | 90.62 | 91 | 8.6 | 90.1892 | 90.92 | 91.2365 | 13.8 | 90.36 | 91.38 | 91.68 | 29.7 | 91.3333 | 92.4185 | 92.6 | 52.9 | 92.325 | 93.05 | 93.44
2.7 | 4.3 | 89.5 | 90.26 | 90.6 | 6.8 | 89.7225 | 90.54 | 90.9864 | 10.7 | 90 | 91.04 | 91.42 | 22.25 | 90.975 | 92.0384 | 92.34 | 42.7 | 92.0333 | 92.68 | 93.18
2.8 | 3.5 | 89.1 | 89.95 | 90.1667 | 5.5 | 89.189 | 90.32 | 90.6562 | 8.5 | 89.6 | 90.7 | 91.05 | 16.75 | 90.5 | 91.7583 | 92 | 34 | 91.8 | 92.475 | 92.925
2.9 | 3.1 | 88.7 | 89.52 | 89.92 | 4.75 | 88.8555 | 89.86 | 90.136 | 7.4 | 89.1 | 90.4 | 90.625 | 13.25 | 90.1333 | 91.3182 | 91.66 | 25.9 | 91.25 | 92.2 | 92.7
3 | 2.7 | 88.25 | 89.18 | 89.54 | 4.1 | 88.4221 | 89.46 | 89.8159 | 6.45 | 88.6 | 89.975 | 90.35 | 11.0333 | 89.6333 | 91.0849 | 91.46 | 19.8 | 90.975 | 91.86 | 92.35
3.1 | 2.5 | 87.85 | 88.65 | 89.1667 | 3.7 | 87.9887 | 89.14 | 89.3757 | 5.8 | 88.2 | 89.6333 | 90 | 9.6 | 89.2667 | 90.8682 | 91.05 | 15.95 | 90.58 | 91.6 | 91.8857
3.2 | 2.3 | 87.3333 | 88.4 | 88.78 | 3.4667 | 87.4719 | 88.7667 | 89.1357 | 5.4 | 87.8 | 89.3 | 89.6 | 8.6 | 88.8667 | 90.4781 | 90.7 | 13.6 | 90.2667 | 91.35 | 91.5667
3.3 | 2.1667 | 86.9 | 88.0667 | 88.42 | 3.3 | 86.8967 | 88.4333 | 88.7855 | 5 | 87.35 | 88.925 | 89.4 | 7.8 | 88.3333 | 90.138 | 90.4 | 11.5 | 89.7 | 91.12 | 91.4
3.4 | 2 | 86.5333 | 87.7 | 88.1 | 3.1 | 86.4883 | 88.1 | 88.4854 | 4.6 | 86.95 | 88.6 | 88.96 | 7.3667 | 88 | 89.8013 | 90.1 | 10.75 | 89.3333 | 90.74 | 91.12
3.5 | 2 | 85.95 | 87.35 | 87.7667 | 2.9 | 86.1549 | 87.7 | 88.0352 | 4.4 | 86.4333 | 88.2 | 88.5667 | 6.9 | 87.7 | 89.6346 | 89.64 | 9.6 | 88.9 | 90.5 | 90.8286
3.6 | 1.9 | 85.6 | 87.05 | 87.4 | 2.7 | 85.5881 | 87.4333 | 87.8101 | 4.2 | 85.8667 | 87.85 | 88.2333 | 6.6667 | 87.15 | 89.3179 | 89.42 | 9.1 | 88.525 | 90.26 | 90.48
3.7 | 1.8 | 85.2333 | 86.8 | 87.14 | 2.6 | 85.1213 | 87 | 87.535 | 4 | 85.5 | 87.6 | 87.95 | 6.4 | 86.675 | 88.9178 | 89.14 | 8.75 | 88.15 | 89.9714 | 90.3143
3.8 | 1.8 | 84.7333 | 86.4 | 86.68 | 2.5 | 84.6212 | 86.7 | 87.1549 | 3.8 | 84.875 | 87.26 | 87.725 | 6.1 | 86.2 | 88.5377 | 88.75 | 8.35 | 87.7 | 89.64 | 90.02
3.9 | 1.7 | 84.3 | 85.975 | 86.4 | 2.35 | 84.121 | 86.375 | 86.7347 | 3.6 | 84.4 | 86.95 | 87.3 | 5.8 | 85.8 | 88.2176 | 88.5 | 7.95 | 87.225 | 89.24 | 89.68
4 | 1.6 | 83.8 | 85.7333 | 86.18 | 2.2 | 83.7876 | 86.12 | 86.5346 | 3.45 | 83.9 | 86.62 | 86.95 | 5.7 | 85.1 | 87.9176 | 88.26 | 7.7 | 86.625 | 89 | 89.3
4.1 | 1.5 | 83.1667 | 85.46 | 85.82 | 2.1 | 82.9207 | 85.8 | 86.1845 | 3.3 | 83.275 | 86.35 | 86.74 | 5.4 | 84.8 | 87.5675 | 87.9 | 7.3 | 86.325 | 88.6333 | 89.1
4.2 | 1.5 | 82.5 | 85.15 | 85.52 | 2.1 | 82.3956 | 85.45 | 85.8343 | 3.2 | 82.9667 | 85.98 | 86.4 | 5.3 | 84.425 | 87.2674 | 87.6 | 7.1 | 85.8333 | 88.28 | 88.86
4.3 | 1.4 | 81.8 | 84.68 | 85.16 | 2 | 81.8872 | 85.125 | 85.5509 | 3.05 | 82.4 | 85.66 | 86.05 | 5.15 | 83.75 | 86.9374 | 87.25 | 6.9 | 85.4 | 88 | 88.575
4.4 | 1.4 | 80.7 | 84.42 | 84.9 | 1.9 | 80.9453 | 84.9 | 85.2141 | 2.9 | 81.7 | 85.25 | 85.66 | 5 | 83.35 | 86.7173 | 86.96 | 6.8 | 85 | 87.76 | 88.34
4.5 | 1.4 | 79.4667 | 84.12 | 84.6 | 1.9 | 80.2201 | 84.575 | 84.909 | 2.85 | 80.9 | 84.95 | 85.36 | 4.8 | 82.8333 | 86.4173 | 86.65 | 6.6 | 84.5333 | 87.48 | 88.0571
4.6 | 1.3 | 76.9 | 83.8667 | 84.24 | 1.8 | 78.8447 | 84.18 | 84.7139 | 2.8 | 80.2 | 84.6833 | 85.14 | 4.6 | 82.1667 | 86.1572 | 86.52 | 6.3333 | 84 | 87.2286 | 87.775
4.7 | 1.3 | 73.35 | 83.45 | 83.85 | 1.7 | 77.1526 | 83.925 | 84.3671 | 2.6 | 79.2 | 84.4333 | 84.9 | 4.5 | 81.6 | 85.8372 | 86.18 | 6.3 | 83.55 | 87.04 | 87.5
4.8 | 1.3 | 67.4 | 83.1 | 83.6333 | 1.6 | 74.5186 | 83.58 | 84.0336 | 2.6 | 78.375 | 84.0667 | 84.55 | 4.4 | 81.2 | 85.5371 | 85.95 | 6.1 | 83.2 | 86.82 | 87.1667
4.9 | 1.2 | 59.1 | 82.8 | 83.38 | 1.6 | 70.6177 | 83.2714 | 83.7669 | 2.5 | 77 | 83.7667 | 84.24 | 4.3 | 80.6 | 85.317 | 85.48 | 5.85 | 82.7 | 86.5 | 86.95
5 | 1.2 | 49.7333 | 82.5 | 82.94 | 1.6 | 64.5161 | 82.9667 | 83.3333 | 2.4 | 74.6 | 83.48 | 83.9 | 4.2 | 79.8333 | 84.9884 | 85.22 | 5.7 | 82.1 | 86.1333 | 86.6333
5.1 | 1.2 | 40 | 82.2 | 82.6 | 1.5 | 56.4141 | 82.48 | 83.1333 | 2.3 | 71.8 | 83.14 | 83.62 | 4 | 79.1 | 84.7169 | 84.8571 | 5.55 | 81.6667 | 85.8 | 86.36
5.2 | 1.1 | 30.7 | 81.825 | 82.275 | 1.5 | 47.5452 | 82.22 | 82.7131 | 2.3 | 66.9 | 82.8333 | 83.28 | 4 | 78.1 | 84.2768 | 84.5333 | 5.35 | 81.2 | 85.4667 | 86.02
5.3 | 1.1 | 22.4 | 81.475 | 81.975 | 1.45 | 38.4096 | 81.9 | 82.3579 | 2.2 | 60.6 | 82.5 | 82.95 | 3.9 | 77.1 | 83.9568 | 84.32 | 5.3 | 80.7667 | 85.2167 | 85.6833
5.4 | 1.1 | 15.7 | 81.125 | 81.575 | 1.4 | 30.0575 | 81.58 | 82.1079 | 2.1 | 53.1 | 82.2 | 82.64 | 3.7 | 75.65 | 83.6967 | 83.95 | 5.2 | 80.175 | 84.8667 | 85.5
5.5 | 1.1 | 11.3333 | 80.7 | 81.25 | 1.4 | 22.6057 | 81.225 | 81.7127 | 2.1 | 44.3 | 81.85 | 82.4 | 3.6 | 73.8 | 83.2767 | 83.72 | 5 | 79.4333 | 84.66 | 85.3
5.6 | 1 | 9 | 80.25 | 80.825 | 1.3 | 16.7376 | 80.84 | 81.3325 | 2 | 36 | 81.45 | 82.075 | 3.5 | 70.75 | 82.8916 | 83.4833 | 4.9 | 78.7 | 84.4 | 85
5.7 | 1 | 7.6 | 79.9 | 80.48 | 1.3 | 12.8532 | 80.575 | 80.9824 | 1.9 | 28.5333 | 81.14 | 81.6 | 3.4 | 66.3333 | 82.6832 | 83.2333 | 4.8 | 77.8 | 84.08 | 84.725
5.8 | 1 | 6.8 | 79.66 | 80.2 | 1.2 | 10.7361 | 80.2 | 80.6656 | 1.9 | 22.2667 | 80.825 | 81.275 | 3.3 | 60.95 | 82.4165 | 82.85 | 4.7 | 76.6667 | 83.8 | 84.36
5.9 | 1 | 6.05 | 79.35 | 79.8 | 1.2 | 9.3023 | 79.9 | 80.3822 | 1.8 | 17.3 | 80.6 | 81 | 3.2 | 54.4 | 82.0664 | 82.4333 | 4.6 | 75.4 | 83.6667 | 83.98
6 | 1 | 5.8 | 78.96 | 79.5167 | 1.2 | 8.4021 | 79.5333 | 79.982 | 1.8 | 14.4 | 80.2333 | 80.6667 | 3.1 | 46.4 | 81.7163 | 82.2 | 4.5 | 73.6 | 83.4167 | 83.7
6.1 | 1 | 5.4 | 78.6 | 79.15 | 1.2 | 7.7353 | 79.08 | 79.5652 | 1.7 | 12.7333 | 79.85 | 80.22 | 3 | 38.8 | 81.4496 | 81.8667 | 4.4 | 70.9333 | 83 | 83.4
6.2 | 1 | 5.175 | 78.225 | 78.75 | 1.2 | 7.2351 | 78.775 | 79.2117 | 1.7 | 11.5333 | 79.4714 | 79.95 | 2.9 | 32 | 81.1329 | 81.5 | 4.4 | 66.85 | 82.7333 | 83.12
6.3 | 0.9 | 4.9667 | 77.92 | 78.44 | 1.2 | 6.8017 | 78.48 | 78.9316 | 1.7 | 10.675 | 79.1 | 79.6 | 2.9 | 26 | 80.7828 | 81.22 | 4.3 | 61.7 | 82.4 | 82.8
6.4 | 0.9 | 4.8 | 77.6 | 78.06 | 1.1 | 6.4683 | 78.12 | 78.5314 | 1.6 | 10 | 78.7667 | 79.175 | 2.8 | 21.8333 | 80.4411 | 80.95 | 4.2 | 55 | 82.08 | 82.45
6.5 | 0.9 | 4.6333 | 77.22 | 77.7 | 1.1 | 6.1015 | 77.7 | 78.2313 | 1.6 | 9.4667 | 78.46 | 78.85 | 2.7 | 18.85 | 80.116 | 80.6 | 4.05 | 48.6333 | 81.8 | 82.2
6.6 | 0.9 | 4.5 | 76.825 | 77.3667 | 1.1 | 6.0015 | 77.45 | 77.9312 | 1.6 | 9.05 | 78.0667 | 78.5333 | 2.7 | 16.8 | 79.696 | 80.22 | 4 | 41.3667 | 81.5 | 81.975
6.7 | 0.9 | 4.4 | 76.475 | 77.1167 | 1.1 | 5.8014 | 77.05 | 77.506 | 1.5 | 8.7667 | 77.7 | 78.1667 | 2.6 | 15.625 | 79.2825 | 79.86 | 3.95 | 34.9667 | 81.2 | 81.725
6.8 | 0.9 | 4.3 | 76.1 | 76.7 | 1.1 | 5.5014 | 76.58 | 77.2309 | 1.5 | 8.4 | 77.35 | 77.86 | 2.6 | 14.7333 | 78.9758 | 79.625 | 3.9 | 29.7 | 80.8333 | 81.4667
6.9 | 0.9 | 4.2 | 75.75 | 76.3 | 1 | 5.3763 | 76.275 | 76.8507 | 1.5 | 8.0667 | 77.08 | 77.4667 | 2.5 | 14.0333 | 78.6157 | 79.26 | 3.8 | 25.1 | 80.3 | 81.1333
7 | 0.9 | 4.025 | 75.4 | 75.94 | 1 | 5.2013 | 75.94 | 76.5306 | 1.4 | 7.8333 | 76.64 | 77.2 | 2.5 | 13.4 | 78.3157 | 78.86 | 3.8 | 22.15 | 79.975 | 80.75
7.1 |  | 3.925 | 75.1 | 75.55 | 1 | 5.1213 | 75.6 | 76.0505 | 1.4 | 7.5667 | 76.3 | 76.85 | 2.5 | 12.8667 | 77.9489 | 78.55 | 3.6 | 20.05 | 79.68 | 80.3143
7.2 |  | 3.9 | 74.7 | 75.225 |  | 4.9262 | 75.2667 | 75.7303 | 1.4 | 7.3667 | 75.85 | 76.3667 | 2.4 | 12.5 | 77.6655 | 78.1667 | 3.5 | 18.9 | 79.4 | 79.9
7.3 |  | 3.8 | 74.3 | 74.9 |  | 4.7512 | 74.9333 | 75.4552 |  | 7.175 | 75.5 | 75.98 | 2.3 | 12.1 | 77.1821 | 77.92 | 3.4 | 17.7 | 79.15 | 79.6
7.4 |  | 3.7 | 73.95 | 74.5 |  | 4.7012 | 74.5 | 75.1101 |  | 7 | 75.125 | 75.625 |  | 11.8 | 76.8487 | 77.6 | 3.4 | 16.9667 | 78.8 | 79.32
7.5 |  | 3.6 | 73.55 | 74.12 |  | 4.6011 | 74.12 | 74.7298 |  | 6.75 | 74.8 | 75.3 |  | 11.5 | 76.6153 | 77.24 | 3.4 | 16.3667 | 78.42 | 79.075
7.6 |  | 3.58 | 73.2 | 73.78 |  | 4.4011 | 73.86 | 74.4298 |  | 6.6 | 74.46 | 75 |  | 11.28 | 76.2953 | 76.9 | 3.2 | 16 | 78.1 | 78.85
7.7 |  | 3.475 | 72.86 | 73.45 |  | 4.3011 | 73.48 | 73.9963 |  | 6.4 | 74.1 | 74.6 |  | 10.9 | 75.9152 | 76.5 |  | 15.5 | 77.6833 | 78.55
7.8 |  | 3.4 | 72.5333 | 73.1 |  | 4.2011 | 73.1 | 73.6294 |  | 6.3 | 73.7 | 74.175 |  | 10.6667 | 75.5551 | 76.1 |  | 15.0333 | 77.5 | 78.1333
7.9 |  | 3.3 | 72.075 | 72.7333 |  | 4.101 | 72.7 | 73.2893 |  | 6.1 | 73.3667 | 73.85 |  | 10.4333 | 75.135 | 75.6667 |  | 14.6 | 77.14 | 77.75
8 |  | 3.2 | 71.825 | 72.35 |  | 4.001 | 72.375 | 72.8791 |  | 5.9 | 73 | 73.5 |  | 10.2 | 74.835 | 75.325 |  | 14.3333 | 76.7667 | 77.4
8.1 |  | 3.2 | 71.45 | 72.025 |  | 3.901 | 71.92 | 72.529 |  | 5.7667 | 72.6 | 73.16 |  | 9.9 | 74.4816 | 75.02 |  | 14.1 | 76.42 | 77.04
8.2 |  | 3.1 | 71.0333 | 71.725 |  | 3.8676 | 71.6 | 72.1789 |  | 5.6 | 72.2667 | 72.78 |  | 9.6 | 74.1648 | 74.6 |  | 13.8333 | 76.0333 | 76.8
8.3 |  | 3.1 | 70.725 | 71.225 |  | 3.7009 | 71.2 | 71.7537 |  | 5.425 | 71.9 | 72.4 |  | 9.5 | 73.7747 | 74.16 |  | 13.4333 | 75.7286 | 76.45
8.4 |  | 3.0667 | 70.3 | 70.8 |  | 3.7009 | 70.825 | 71.4286 |  | 5.28 | 71.45 | 72.02 |  | 9.275 | 73.3647 | 73.92 |  | 13.225 | 75.4 | 76.025
8.5 |  | 3 | 70 | 70.4667 |  | 3.6009 | 70.3667 | 70.9484 |  | 5.2 | 71.1 | 71.58 |  | 9.1 | 72.9479 | 73.55 |  | 12.9333 | 75.0667 | 75.65
8.6 |  | 2.9 | 69.48 | 69.98 |  | 3.5009 | 70.0333 | 70.5532 |  | 5.1 | 70.6833 | 71.22 |  | 8.8333 | 72.6345 | 73.15 |  | 12.6 | 74.8167 | 75.3333
8.7 |  | 2.9 | 69.05 | 69.7 |  | 3.5009 | 69.55 | 70.2681 |  | 5 | 70.34 | 70.84 |  | 8.6 | 72.2644 | 72.725 |  | 12.4 | 74.45 | 75.02
8.8 |  | 2.9 | 68.6667 | 69.26 |  | 3.4009 | 69.225 | 69.7479 |  | 4.9 | 69.9 | 70.4 |  | 8.4333 | 71.8644 | 72.3667 |  | 12 | 74.12 | 74.62
8.9 |  | 2.8 | 68.3 | 68.8 |  | 3.3008 | 68.84 | 69.3077 |  | 4.8 | 69.5 | 69.9667 |  | 8.3 | 71.4476 | 71.95 |  | 11.8667 | 73.7667 | 74.2
9 |  | 2.8 | 67.85 | 68.4 |  | 3.2758 | 68.35 | 68.9276 |  | 4.675 | 68.98 | 69.5 |  | 8.1 | 70.9942 | 71.7 |  | 11.65 | 73.4 | 73.975
9.1 |  | 2.7 | 67.3 | 67.975 |  | 3.2008 | 67.94 | 68.5274 |  | 4.525 | 68.7 | 69.15 |  | 7.9333 | 70.7141 | 71.35 |  | 11.5 | 73.1 | 73.68
9.2 |  | 2.7 | 66.8 | 67.44 |  | 3.1008 | 67.475 | 67.9272 |  | 4.5 | 68.2 | 68.75 |  | 7.7 | 70.2997 | 70.9333 |  | 11.3 | 72.7 | 73.25
9.3 |  | 2.7 | 66.3 | 66.82 |  | 3.1008 | 67.04 | 67.527 |  | 4.4 | 67.78 | 68.36 |  | 7.6 | 69.8807 | 70.4667 |  | 11.2 | 72.2667 | 73
9.4 |  | 2.6 | 65.6 | 66.28 |  | 3.0008 | 66.45 | 67.0268 |  | 4.4 | 67.38 | 67.95 |  | 7.4 | 69.5282 | 70.1333 |  | 10.85 | 72.04 | 72.6333
9.5 |  | 2.6 | 64.7 | 65.35 |  | 3.0008 | 65.95 | 66.4266 |  | 4.25 | 67 | 67.5 |  | 7.3 | 69.2138 | 69.675 |  | 10.675 | 71.6667 | 72.22
9.6 |  | 2.6 | 63.6 | 64.275 |  | 2.9007 | 65.325 | 65.9764 |  | 4.1333 | 66.425 | 66.95 |  | 7.2 | 68.6537 | 69.425 |  | 10.5 | 71.25 | 71.9
9.7 |  | 2.575 | 61.6 | 62.48 |  | 2.9007 | 64.575 | 65.2094 |  | 4.05 | 65.88 | 66.54 |  | 7.025 | 68.2737 | 68.875 |  | 10.3 | 70.86 | 71.4714
9.8 |  | 2.5 | 58.7 | 59.62 |  | 2.9007 | 63.825 | 64.5058 |  | 3.95 | 65.48 | 66 |  | 6.9 | 67.9469 | 68.55 |  | 10.1333 | 70.5667 | 71.06
9.9 |  | 2.475 | 53.8167 | 55.3 |  | 2.8007 | 62.775 | 63.5654 |  | 3.9 | 64.96 | 65.46 |  | 6.8 | 67.4735 | 68.1857 |  | 10 | 70.225 | 70.75
10 |  | 2.4 | 47.95 | 49.05 |  | 2.8007 | 61.35 | 62.0248 |  | 3.8 | 64.42 | 64.9143 |  | 6.7 | 67.0534 | 67.76 |  | 9.9 | 69.72 | 70.4333
10.1 |  |  | 41.025 | 42.98 |  | 2.7674 | 59.05 | 59.964 |  | 3.8 | 63.85 | 64.45 |  | 6.525 | 66.6333 | 67.35 |  | 9.75 | 69.3333 | 70.15
10.2 |  |  | 33.7 | 35.85 |  | 2.7007 | 55.8 | 56.4893 |  | 3.7 | 63.05 | 63.5714 |  | 6.4 | 66.2883 | 66.9 |  | 9.6 | 68.96 | 69.68
10.3 |  |  | 27.1 | 29 |  |  | 50.75 | 51.6957 |  | 3.6667 | 62.04 | 62.825 |  | 6.3 | 65.8799 | 66.525 |  | 9.5 | 68.725 | 69.42
10.4 |  |  | 20.85 | 23.22 |  |  | 44.6 | 45.8383 |  | 3.6 | 61.18 | 61.76 |  | 6.2 | 65.4131 | 66 |  | 9.3667 | 68.35 | 69.025
10.5 |  |  | 16.1 | 18.74 |  |  | 38.5 | 39.5825 |  |  | 59.6333 | 60.44 |  | 6.1 | 64.9797 | 65.55 |  | 9.2 | 67.95 | 68.7333
10.6 |  |  | 13.1 | 15.8667 |  |  | 31.5 | 33.6135 |  |  | 57.5 | 58.4 |  | 6 | 64.5129 | 65 |  | 9.1 | 67.5667 | 68.2333
10.7 |  |  | 11.475 | 14.1 |  |  | 25.46 | 27.511 |  |  | 54.16 | 55.225 |  | 5.9 | 63.9878 | 64.7 |  | 8.9 | 67.05 | 67.8
10.8 |  |  | 10.4 | 13.2 |  |  | 20.25 | 22.6341 |  |  | 49.5 | 50.9 |  | 5.8 | 63.4727 | 64.16 |  | 8.7 | 66.7 | 67.45
10.9 |  |  | 9.7 | 12.5 |  |  | 16.34 | 18.9325 |  |  | 44 | 45.7429 |  | 5.7 | 62.9876 | 63.6 |  | 8.6 | 66.25 | 66.96
11 |  |  | 9.36 | 12.1333 |  |  | 13.98 | 16.6066 |  |  | 38 | 40 |  |  | 62.3724 | 63.16 |  | 8.5 | 65.8143 | 66.66
11.1 |  |  | 9.08 | 11.8 |  |  | 12.55 | 15.2061 |  |  | 32.28 | 34.62 |  |  | 61.9124 | 62.6 |  | 8.36 | 65.4333 | 66.25
11.2 |  |  | 8.85 | 11.55 |  |  | 11.625 | 14.3057 |  |  | 26.6167 | 29.2 |  |  | 61.2623 | 62 |  | 8.2 | 65.1 | 65.8333
11.3 |  |  | 8.65 | 11.4 |  |  | 10.98 | 13.7055 |  |  | 21.95 | 24.5 |  |  | 60.4621 | 61.1667 |  | 8 | 64.66 | 65.5
11.4 |  |  | 8.5 | 11.28 |  |  | 10.5 | 13.1803 |  |  | 18.5 | 21.22 |  |  | 59.5119 | 60.2 |  | 7.925 | 64.2 | 65.1
11.5 |  |  | 8.3667 | 11.1 |  |  | 10.15 | 12.8451 |  |  | 16.3667 | 19.02 |  |  | 58.2916 | 58.925 |  |  | 63.78 | 64.6
11.6 |  |  | 8.275 | 11.025 |  |  | 9.9 | 12.605 |  |  | 14.9 | 17.725 |  |  | 56.2112 | 57.2 |  |  | 63.275 | 64.1
11.7 |  |  | 8.2 | 10.9 |  |  | 9.68 | 12.405 |  |  | 14.05 | 16.86 |  |  | 53.6108 | 54.4833 |  |  | 62.9 | 63.66
11.8 |  |  | 8.1 | 10.8667 |  |  | 9.45 | 12.1048 |  |  | 13.4 | 16.26 |  |  | 49.6099 | 50.95 |  |  | 62.42 | 63.28
11.9 |  |  | 8.1 | 10.8 |  |  | 9.22 | 11.9798 |  |  | 12.88 | 15.7 |  |  | 45.109 | 46.6 |  |  | 61.9 | 62.84
12 |  |  | 8.0333 | 10.7 |  |  | 9.1 | 11.8047 |  |  | 12.48 | 15.325 |  |  | 40.4081 | 41.7833 |  |  | 61.42 | 62.25
12.1 |  |  | 8 | 10.7 |  |  | 9 | 11.7047 |  |  | 12.2 | 14.96 |  |  | 35.127 | 37.3 |  |  | 60.9 | 61.74
12.2 |  |  | 7.9 | 10.6167 |  |  | 8.85 | 11.5046 |  |  | 11.88 | 14.7 |  |  | 30.6061 | 32.52 |  |  | 60.125 | 61.02
12.3 |  |  | 7.9 | 10.6 |  |  | 8.7 | 11.4046 |  |  | 11.65 | 14.4333 |  |  | 26.2852 | 28.7 |  |  | 59.44 | 60.4
12.4 |  |  | 7.825 | 10.575 |  |  | 8.65 | 11.3045 |  |  | 11.38 | 14.225 |  |  | 22.9546 | 25.64 |  |  | 58.425 | 59.4
12.5 |  |  | 7.8 | 10.5 |  |  | 8.6 | 11.2045 |  |  | 11.2 | 13.98 |  |  | 20.5241 | 23.4 |  |  | 57.3167 | 58.2333
12.6 |  |  | 7.8 | 10.5 |  |  | 8.5 | 11.1044 |  |  | 11 | 13.78 |  |  | 19.1288 | 22.18 |  |  | 55.32 | 56.7
12.7 |  |  | 7.7 | 10.4 |  |  | 8.4 | 11.1044 |  |  | 10.8 | 13.6 |  |  | 18.0636 | 21.3 |  |  | 52.9 | 54.15
12.8 |  |  | 7.7 | 10.3 |  |  | 8.4 | 10.9794 |  |  | 10.7 | 13.4 |  |  | 17.5035 | 20.55 |  |  | 49.5 | 51.2
12.9 |  |  | 7.6 | 10.3 |  |  | 8.26 | 10.9044 |  |  | 10.475 | 13.3 |  |  | 16.8284 | 20.05 |  |  | 45.5833 | 47.32
13 |  |  | 7.6 | 10.2 |  |  | 8.2 | 10.8443 |  |  | 10.3 | 13.125 |  |  | 16.2832 | 19.5667 |  |  | 41.34 | 43.225
13.1 |  |  | 7.5 | 10.2 |  |  | 8.1 | 10.7243 |  |  | 10.2 | 13 |  |  | 15.9699 | 19.2 |  |  | 36.825 | 39.34
13.2 |  |  | 7.5 | 10.1 |  |  | 8.02 | 10.6042 |  |  | 10.1 | 12.84 |  |  | 15.7031 | 18.825 |  |  | 32.375 | 35.3
13.3 |  |  | 7.4 | 10 |  |  | 8 | 10.6042 |  |  | 9.975 | 12.72 |  |  | 15.4364 | 18.5 |  |  | 28.76 | 31.8
13.4 |  |  | 7.3 | 9.9 |  |  | 7.9 | 10.5042 |  |  | 9.85 | 12.625 |  |  | 15.083 | 18.1333 |  |  | 25.65 | 29.02
13.5 |  |  | 7.2667 | 9.8 |  |  | 7.8 | 10.4042 |  |  | 9.75 | 12.46 |  |  | 14.878 | 17.925 |  |  | 23.4 | 26.9
13.6 |  |  | 7.2 | 9.8 |  |  | 7.7 | 10.3375 |  |  | 9.6 | 12.325 |  |  | 14.5829 | 17.7571 |  |  | 22.1 | 25.5167
13.7 |  |  | 7.05 | 9.675 |  |  | 7.6167 | 10.3041 |  |  | 9.48 | 12.2 |  |  | 14.3229 | 17.6 |  |  | 21.1 | 24.6
13.8 |  |  | 7 | 9.6 |  |  | 7.6 | 10.2041 |  |  | 9.35 | 12.1 |  |  | 14.0778 | 17.32 |  |  | 20.3667 | 23.8667
13.9 |  |  | 6.9 | 9.5 |  |  | 7.5 | 10.004 |  |  | 9.3 | 12 |  |  | 13.9028 | 17.04 |  |  | 19.825 | 23.3333
14 |  |  | 6.8 | 9.4 |  |  | 7.4 | 10.004 |  |  | 9.175 | 11.9 |  |  | 13.7027 | 16.8333 |  |  | 19.325 | 22.9
14.1 |  |  | 6.8 | 9.32 |  |  | 7.4 | 9.904 |  |  | 9.05 | 11.74 |  |  | 13.5777 | 16.64 |  |  | 19.04 | 22.56
14.2 |  |  | 6.7 | 9.24 |  |  | 7.3 | 9.8039 |  |  | 8.96 | 11.65 |  |  | 13.3827 | 16.5 |  |  | 18.5 | 22.275
14.3 |  |  | 6.625 | 9.1 |  |  | 7.2 | 9.7039 |  |  | 8.84 | 11.5 |  |  | 13.2627 | 16.3 |  |  | 18.24 | 21.96
14.4 |  |  | 6.6 | 9.1 |  |  | 7.1 | 9.6038 |  |  | 8.78 | 11.36 |  |  | 13.0626 | 16.14 |  |  | 17.9333 | 21.625
14.5 |  |  | 6.45 | 8.96 |  |  | 7.075 | 9.5038 |  |  | 8.6 | 11.3 |  |  | 12.8693 | 15.92 |  |  | 17.8 | 21.35
14.6 |  |  | 6.4 | 8.9 |  |  | 6.975 | 9.5038 |  |  | 8.5 | 11.2 |  |  | 12.7025 | 15.75 |  |  | 17.62 | 21.0833
14.7 |  |  | 6.3 | 8.8 |  |  | 6.85 | 9.4038 |  |  | 8.4 | 11.1 |  |  | 12.5358 | 15.6 |  |  | 17.4 | 20.8333
14.8 |  |  | 6.2571 | 8.7 |  |  | 6.7333 | 9.3037 |  |  | 8.3 | 11 |  |  | 12.3625 | 15.475 |  |  | 17.075 | 20.55
14.9 |  |  | 6.18 | 8.625 |  |  | 6.7 | 9.2037 |  |  | 8.2 | 10.8 |  |  | 12.1624 | 15.25 |  |  | 16.75 | 20.375
15 |  |  | 6.1 | 8.6 |  |  | 6.6 | 9.0369 |  |  | 8.075 | 10.7 |  |  | 12.0024 | 15.06 |  |  | 16.575 | 20.12
15.1 |  |  | 6 | 8.5 |  |  | 6.5 | 8.9036 |  |  | 7.96 | 10.6 |  |  | 11.9024 | 14.9 |  |  | 16.325 | 19.96
15.2 | [6 empty] | 6.4 | 8.9036 |  |  | 7.875 | 10.5 |  |  | 11.7423 | 14.76 |  |  | 16.075 | 19.7
15.3 | [6 empty] | 6.4 | 8.8035 |  |  | 7.8 | 10.4 |  |  | 11.6023 | 14.6 |  |  | 15.9 | 19.4667
15.4 | [6 empty] | 6.3 | 8.7035 |  |  | 7.7 | 10.325 |  |  | 11.5023 | 14.44 |  |  | 15.7 | 19.2667
15.5 | [10 empty] | 7.6 | 10.2 |  |  | 11.3023 | 14.325 |  |  | 15.5167 | 19.075
15.6 | [10 empty] | 7.5 | 10.08 |  |  | 11.1689 | 14.12 |  |  | 15.26 | 18.9
15.7 | [10 empty] | 7.4 | 10 |  |  | 11.0822 | 14.0167 |  |  | 15.04 | 18.75
15.8 | [10 empty] | 7.3333 | 9.9 |  |  | 11.0022 | 13.875 |  |  | 14.94 | 18.5333
15.9 | [10 empty] | 7.24 | 9.9 |  |  | 10.745 | 13.7667 |  |  | 14.8 | 18.3667
16 | [14 empty] | 10.6771 | 13.58 |  |  | 14.625 | 18.175
16.1 | [14 empty] | 10.5271 | 13.4 |  |  | 14.42 | 17.92
16.2 | [14 empty] | 10.4021 | 13.3 |  |  | 14.3 | 17.7
16.3 | [14 empty] | 10.242 | 13.2 |  |  | 14.1 | 17.6
16.4 | [14 empty] | 10.152 | 13.1 |  |  | 14 | 17.4
16.5 | [14 empty] | 10.022 | 12.9 |  |  | 13.86 | 17.18
16.6 | [14 empty] | 9.952 | 12.7 |  |  | 13.7 | 17.1
16.7 | [14 empty] | 9.802 | 12.6 |  |  | 13.55 | 16.95
16.8 | [14 empty] | 9.6019 | 12.5 |  |  | 13.4 | 16.775
16.9 | [14 empty] | 9.5019 | 12.4 |  |  | 13.2833 | 16.56
17 | [18 empty] | 13.08 | 16.35
17.1 | [18 empty] | 12.96 | 16.3
17.2 | [18 empty] | 12.7833 | 16.08
17.3 | [18 empty] | 12.6286 | 15.98
17.4 | [18 empty] | 12.4333 | 15.8
17.5 | [18 empty] | 12.2667 | 15.6667
17.6 | [18 empty] | 12.2 | 15.525
17.7 | [18 empty] | 12 | 15.3333
17.8 | [18 empty] | 11.95 | 15.14
17.9 | [18 empty] | 11.8 | 14.925

## Wedge 45 Longitudinal Profiles
Longitudinal Profiles
Wedge 45
SSD=100cm
Field Size=40x20cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 14.0833
-30.8 |  |  |  |  | 14.2375
-30.7 |  |  |  |  | 14.3625
-30.6 |  |  |  |  | 14.575
-30.5 |  |  |  |  | 14.7125
-30.4 |  |  |  |  | 14.8375
-30.3 |  |  |  |  | 15.025
-30.2 |  |  |  |  | 15.1778
-30.1 |  |  |  |  | 15.4
-30 |  |  |  |  | 15.525
-29.9 |  |  |  |  | 15.65
-29.8 |  |  |  |  | 15.775
-29.7 |  |  |  |  | 16
-29.6 |  |  |  |  | 16.25
-29.5 |  |  |  |  | 16.45
-29.4 |  |  |  |  | 16.5857
-29.3 |  |  |  |  | 16.8
-29.2 |  |  |  |  | 17.02
-29.1 |  |  |  |  | 17.2
-29 |  |  |  |  | 17.375
-28.9 |  |  |  | 11.7875 | 17.6667
-28.8 |  |  |  | 11.9125 | 17.925
-28.7 |  |  |  | 12.0375 | 18.0667
-28.6 |  |  |  | 12.1625 | 18.2286
-28.5 |  |  |  | 12.2875 | 18.4667
-28.4 |  |  |  | 12.4125 | 18.7333
-28.3 |  |  |  | 12.5375 | 18.9
-28.2 |  |  |  | 12.7111 | 19.1143
-28.1 |  |  |  | 12.875 | 19.3143
-28 |  |  |  | 13.025 | 19.6
-27.9 |  |  |  | 13.15 | 19.875
-27.8 |  |  |  | 13.35 | 20.1875
-27.7 |  |  |  | 13.5 | 20.475
-27.6 |  |  |  | 13.68 | 20.8222
-27.5 |  |  |  | 13.8429 | 21.075
-27.4 |  |  |  | 14.0714 | 21.45
-27.3 |  |  |  | 14.25 | 21.825
-27.2 |  |  |  | 14.4286 | 22.1
-27.1 |  |  |  | 14.6571 | 22.5
-27 |  |  |  | 14.875 | 23
-26.9 |  |  | 9.275 | 15.0167 | 23.4
-26.8 |  |  | 9.3167 | 15.2429 | 24.1857
-26.7 |  |  | 9.4833 | 15.4333 | 25.1
-26.6 |  |  | 9.6 | 15.6833 | 26.2
-26.5 |  |  | 9.6375 | 15.8 | 27.7
-26.4 |  |  | 9.7556 | 16.05 | 29.9
-26.3 |  |  | 9.8 | 16.3333 | 32.6333
-26.2 |  |  | 9.9429 | 16.54 | 36.3
-26.1 |  |  | 10 | 16.8286 | 40.2
-26 |  |  | 10.125 | 17.1143 | 44.5571
-25.9 |  | 8 | 10.2571 | 17.3286 | 49.2
-25.8 |  | 8 | 10.4 | 17.6143 | 54.0143
-25.7 |  | 8.1 | 10.5667 | 17.9 | 58.3167
-25.6 |  | 8.1714 | 10.625 | 18.1429 | 62.2571
-25.5 |  | 8.2 | 10.7667 | 18.4333 | 65.4571
-25.4 |  | 8.3 | 10.9 | 18.7875 | 67.75
-25.3 |  | 8.4 | 11 | 19.2333 | 69.5
-25.2 | 7.5875 | 8.5 | 11.1714 | 19.5429 | 70.9625
-25.1 | 7.6571 | 8.5 | 11.3286 | 19.9286 | 72
-25 | 7.7 | 8.6 | 11.4714 | 20.4429 | 72.8125
-24.9 | 7.7625 | 8.7 | 11.6333 | 21.0625 | 73.525
-24.8 | 7.8 | 8.7667 | 11.8 | 21.725 | 74.1
-24.7 | 7.8833 | 8.8286 | 11.9714 | 22.6625 | 74.56
-24.6 | 7.9286 | 8.9714 | 12.2143 | 24.05 | 74.9714
-24.5 | 8 | 9 | 12.3571 | 25.75 | 75.5
-24.4 | 8.1 | 9.1375 | 12.6167 | 28.35 | 75.7333
-24.3 | 8.1143 | 9.2 | 12.825 | 31.7 | 76.12
-24.2 | 8.2 | 9.3 | 13 | 35.8571 | 76.54
-24.1 | 8.3 | 9.4429 | 13.2125 | 40.6 | 76.9
-24 | 8.3 | 9.5333 | 13.3857 | 45.95 | 77.1667
-23.9 | 8.4 | 9.7 | 13.6714 | 51.8 | 77.3857
-23.8 | 8.4143 | 9.85 | 13.9286 | 56.9667 | 77.7
-23.7 | 8.5 | 10 | 14.0833 | 62 | 78.075
-23.6 | 8.6 | 10.1167 | 14.32 | 66.4286 | 78.3875
-23.5 | 8.6 | 10.2 | 14.6333 | 69.9 | 78.7333
-23.4 | 8.7 | 10.35 | 14.9 | 72.5714 | 79.0571
-23.3 | 8.7286 | 10.5286 | 15.2333 | 74.5 | 79.2556
-23.2 | 8.8 | 10.7 | 15.6667 | 75.8429 | 79.45
-23.1 | 8.9333 | 10.9143 | 15.9714 | 76.8 | 79.8286
-23 | 9.0143 | 11.05 | 16.4667 | 77.5667 | 80.1143
-22.9 | 9.1571 | 11.2167 | 17 | 78.15 | 80.3286
-22.8 | 9.3 | 11.3833 | 17.7429 | 78.7 | 80.55
-22.7 | 9.4 | 11.6 | 18.6333 | 79.2 | 80.7714
-22.6 | 9.5 | 11.82 | 20.1333 | 79.6857 | 81.05
-22.5 | 9.6167 | 12.02 | 22 | 80.0429 | 81.2
-22.4 | 9.7625 | 12.3 | 25.0857 | 80.38 | 81.4857
-22.3 | 9.9 | 12.5667 | 29.2 | 80.7 | 81.6857
-22.2 | 10.05 | 12.84 | 34.4143 | 81.1333 | 81.9286
-22.1 | 10.2 | 13.15 | 39.9571 | 81.4286 | 82.2143
-22 | 10.2833 | 13.5429 | 47.1286 | 81.6667 | 82.4
-21.9 | 10.4667 | 14.1 | 53.9 | 81.9 | 82.625
-21.8 | 10.6 | 14.8 | 61.3667 | 82.3667 | 82.7667
-21.7 | 10.8333 | 15.8 | 67.4 | 82.6333 | 82.94
-21.6 | 11.0167 | 17.3714 | 72.8 | 82.9111 | 83.25
-21.5 | 11.1833 | 19.7 | 76.6143 | 83.1714 | 83.5
-21.4 | 11.52 | 23 | 79.4 | 83.3 | 83.7286
-21.3 | 11.8429 | 27.85 | 81.1667 | 83.675 | 83.9
-21.2 | 12.2 | 33.6833 | 82.3 | 84 | 84.05
-21.1 | 12.8333 | 40.9 | 83.26 | 84.25 | 84.2333
-21 | 13.7857 | 48.1 | 83.98 | 84.4 | 84.525
-20.9 | 15.2 | 56.0286 | 84.6429 | 84.7 | 84.6875
-20.8 | 17.5667 | 64.1333 | 85.24 | 85 | 84.9125
-20.7 | 21.3 | 71.2833 | 85.6 | 85.24 | 85.075
-20.6 | 26.55 | 77.1 | 85.98 | 85.3875 | 85.2625
-20.5 | 33.1 | 80.9667 | 86.28 | 85.7125 | 85.4556
-20.4 | 40.7 | 83.62 | 86.6333 | 85.8375 | 85.6286
-20.3 | 49.8 | 85.3667 | 86.8833 | 86.025 | 85.9
-20.2 | 58.4 | 86.7 | 87.1 | 86.1875 | 86.075
-20.1 | 67.2667 | 87.58 | 87.3333 | 86.4 | 86.2
-20 | 74.8167 | 88.0667 | 87.6429 | 86.575 | 86.3375
-19.9 | 80.7333 | 88.8 | 87.9571 | 86.9143 | 86.5667
-19.8 | 84.7167 | 89.0857 | 88.175 | 87.1167 | 86.7429
-19.7 | 87.1571 | 89.4 | 88.5286 | 87.2714 | 86.8857
-19.6 | 88.7333 | 89.6429 | 88.78 | 87.5 | 87.0333
-19.5 | 89.8 | 90 | 88.9667 | 87.66 | 87.22
-19.4 | 90.5 | 90.3 | 89.15 | 87.8667 | 87.3857
-19.3 | 91.0667 | 90.5333 | 89.4143 | 88.1 | 87.575
-19.2 | 91.3286 | 90.7833 | 89.6167 | 88.2571 | 87.7625
-19.1 | 91.6333 | 90.9286 | 89.7714 | 88.4143 | 87.9333
-19 | 91.9571 | 91.0714 | 90 | 88.58 | 88.0167
-18.9 | 92.1 | 91.3333 | 90.1 | 88.66 | 88.1
-18.8 | 92.28 | 91.5 | 90.2375 | 88.8333 | 88.2571
-18.7 | 92.4 | 91.6 | 90.3625 | 89.0833 | 88.425
-18.6 | 92.6333 | 91.7333 | 90.5286 | 89.2667 | 88.55
-18.5 | 92.8 | 91.9 | 90.7571 | 89.4286 | 88.7333
-18.4 | 92.8571 | 92.0429 | 90.9 | 89.625 | 88.8875
-18.3 | 93.02 | 92.1857 | 91.1429 | 89.7167 | 89.0125
-18.2 | 93.1857 | 92.3167 | 91.3143 | 89.8833 | 89.1375
-18.1 | 93.2571 | 92.4833 | 91.4667 | 90.05 | 89.3429
-18 | 93.3 | 92.6 | 91.6 | 90.2 | 89.5333
-17.9 | 93.4333 | 92.7333 | 91.725 | 90.2857 | 89.7
-17.8 | 93.5 | 92.9 | 91.8667 | 90.58 | 89.8429
-17.7 | 93.7 | 93.02 | 92 | 90.7 | 90.05
-17.6 | 93.8 | 93.2 | 92.1 | 90.775 | 90.2
-17.5 | 93.9 | 93.2667 | 92.1833 | 90.9 | 90.25
-17.4 | 93.9286 | 93.4286 | 92.35 | 91.0333 | 90.3
-17.3 | 94.0833 | 93.5 | 92.5286 | 91.175 | 90.38
-17.2 | 94.1429 | 93.65 | 92.7571 | 91.3 | 90.56
-17.1 | 94.3 | 93.8143 | 92.8167 | 91.425 | 90.7333
-17 | 94.3833 | 93.9 | 93 | 91.55 | 90.9143
-16.9 | 94.45 | 94.06 | 93.06 | 91.7 | 91.05
-16.8 | 94.6167 | 94.1286 | 93.3 | 91.85 | 91.175
-16.7 | 94.7 | 94.2833 | 93.4125 | 91.9 | 91.2889
-16.6 | 94.8429 | 94.3571 | 93.5375 | 92.0429 | 91.4125
-16.5 | 94.9 | 94.5 | 93.6 | 92.2167 | 91.5429
-16.4 | 95.05 | 94.5167 | 93.7375 | 92.3833 | 91.6
-16.3 | 95.2 | 94.6 | 93.8 | 92.4167 | 91.76
-16.2 | 95.2833 | 94.7667 | 93.9 | 92.5625 | 91.9
-16.1 | 95.3 | 94.9 | 94.0143 | 92.7 | 91.9429
-16 | 95.3833 | 94.9429 | 94.1 | 92.8 | 92.0857
-15.9 | 95.48 | 95.0857 | 94.2333 | 92.9 | 92.225
-15.8 | 95.5167 | 95.1 | 94.3 | 92.9833 | 92.3
-15.7 | 95.6 | 95.2333 | 94.4 | 93.15 | 92.4
-15.6 | 95.75 | 95.3 | 94.5375 | 93.2 | 92.5286
-15.5 | 95.8286 | 95.4 | 94.6 | 93.3143 | 92.6833
-15.4 | 95.9 | 95.5 | 94.7 | 93.48 | 92.8143
-15.3 | 95.9333 | 95.5 | 94.8 | 93.5375 | 92.98
-15.2 | 96 | 95.6 | 94.9 | 93.6625 | 93.1143
-15.1 | 96.1667 | 95.7 | 95 | 93.7 | 93.2571
-15 | 96.2286 | 95.78 | 95.1 | 93.8167 | 93.3
-14.9 | 96.3 | 95.8571 | 95.2 | 93.9556 | 93.4
-14.8 | 96.3 | 95.9333 | 95.3 | 94 | 93.5
-14.7 | 96.3 | 96 | 95.3857 | 94.1 | 93.5667
-14.6 | 96.4 | 96.1 | 95.4167 | 94.2167 | 93.7286
-14.5 | 96.5 | 96.1857 | 95.5833 | 94.3833 | 93.8714
-14.4 | 96.5875 | 96.3 | 95.6 | 94.5333 | 94.0167
-14.3 | 96.6 | 96.3667 | 95.7 | 94.6 | 94.1
-14.2 | 96.74 | 96.4 | 95.7 | 94.7 | 94.1
-14.1 | 96.8 | 96.4857 | 95.8 | 94.7667 | 94.2429
-14 | 96.8 | 96.5 | 95.9 | 94.8875 | 94.4
-13.9 | 96.9 | 96.5125 | 95.9833 | 95 | 94.4375
-13.8 | 97 | 96.65 | 96 | 95.1 | 94.5625
-13.7 | 97 | 96.7 | 96.1 | 95.2 | 94.6875
-13.6 | 97.05 | 96.7 | 96.1 | 95.2833 | 94.8
-13.5 | 97.1167 | 96.8 | 96.2 | 95.4 | 94.8429
-13.4 | 97.2 | 96.85 | 96.3429 | 95.4857 | 94.9
-13.3 | 97.2333 | 97 | 96.4 | 95.5 | 94.9833
-13.2 | 97.32 | 97.0571 | 96.5 | 95.6 | 95.1
-13.1 | 97.4 | 97.1 | 96.65 | 95.7 | 95.2
-13 | 97.4 | 97.2 | 96.7333 | 95.8 | 95.2714
-12.9 | 97.5 | 97.2333 | 96.8 | 95.88 | 95.4
-12.8 | 97.5857 | 97.3167 | 96.8375 | 95.9167 | 95.4
-12.7 | 97.6 | 97.4 | 96.9 | 96 | 95.4875
-12.6 | 97.7 | 97.5 | 97 | 96.1143 | 95.6
-12.5 | 97.8 | 97.6 | 97.0429 | 96.2571 | 95.6833
-12.4 | 97.8857 | 97.6 | 97.175 | 96.3 | 95.8571
-12.3 | 97.9 | 97.6167 | 97.2 | 96.45 | 95.9
-12.2 | 97.9 | 97.7 | 97.3 | 96.5 | 96
-12.1 | 97.9333 | 97.7333 | 97.3286 | 96.5 | 96
-12 | 98.0857 | 97.9 | 97.4 | 96.6 | 96.0571
-11.9 | 98.1 | 97.9833 | 97.5 | 96.6 | 96.1
-11.8 | 98.2 | 98 | 97.6 | 96.7333 | 96.2
-11.7 | 98.2375 | 98.06 | 97.6714 | 96.8 | 96.275
-11.6 | 98.4 | 98.04 | 97.7 | 96.9 | 96.3
-11.5 | 98.45 | 98.1 | 97.7667 | 96.92 | 96.4125
-11.4 | 98.5 | 98.2 | 97.8333 | 97 | 96.5375
-11.3 | 98.6 | 98.2 | 97.9 | 97.1 | 96.6625
-11.2 | 98.6 | 98.3 | 97.9429 | 97.1 | 96.7333
-11.1 | 98.7 | 98.4 | 98 | 97.25 | 96.8
-11 | 98.7833 | 98.5 | 98.0714 | 97.3 | 96.8
-10.9 | 98.8 | 98.5 | 98.1 | 97.4143 | 96.8714
-10.8 | 98.8333 | 98.6 | 98.1857 | 97.5571 | 97
-10.7 | 98.9 | 98.6143 | 98.2 | 97.6111 | 97
-10.6 | 98.9 | 98.7 | 98.2 | 97.7 | 97.075
-10.5 | 98.9286 | 98.72 | 98.3 | 97.75 | 97.2
-10.4 | 99 | 98.8 | 98.3833 | 97.8 | 97.3
-10.3 | 99.0667 | 98.8 | 98.46 | 97.8167 | 97.3375
-10.2 | 99.1 | 98.8 | 98.5 | 97.9 | 97.4
-10.1 | 99.1 | 98.85 | 98.5857 | 97.9286 | 97.525
-10 | 99.2 | 98.9 | 98.6667 | 98 | 97.6
-9.9 | 99.2 | 99 | 98.7 | 98.05 | 97.7
-9.8 | 99.2375 | 99.075 | 98.8 | 98.1 | 97.7
-9.7 | 99.3 | 99.125 | 98.8 | 98.2571 | 97.8
-9.6 | 99.3 | 99.2 | 98.9 | 98.3 | 97.8
-9.5 | 99.3 | 99.2 | 98.9 | 98.3667 | 97.8
-9.4 | 99.4 | 99.2333 | 98.9 | 98.4 | 97.9
-9.3 | 99.42 | 99.3 | 99 | 98.4 | 97.9
-9.2 | 99.5 | 99.4 | 99 | 98.45 | 98
-9.1 | 99.6 | 99.4 | 99.1 | 98.5 | 98.1
-9 | 99.6667 | 99.4167 | 99.1 | 98.6429 | 98.1833
-8.9 | 99.74 | 99.5 | 99.2 | 98.7857 | 98.2
-8.8 | 99.8 | 99.6 | 99.3 | 98.8 | 98.2857
-8.7 | 99.8 | 99.6 | 99.3 | 98.9 | 98.3
-8.6 | 99.9 | 99.7667 | 99.4 | 98.9 | 98.3125
-8.5 | 99.9 | 99.7429 | 99.4 | 99 | 98.4375
-8.4 | 99.9 | 99.8 | 99.45 | 99 | 98.5
-8.3 | 100 | 99.8 | 99.5 | 99 | 98.5286
-8.2 | 100 | 99.8 | 99.5143 | 99 | 98.6
-8.1 | 100 | 99.9 | 99.6 | 99.075 | 98.6556
-8 | 100.1 | 99.9 | 99.6 | 99.1 | 98.7
-7.9 | 100.2 | 99.9 | 99.6 | 99.2 | 98.7429
-7.8 | 100.2 | 99.9571 | 99.6714 | 99.2143 | 98.8
-7.7 | 100.2 | 100 | 99.7 | 99.3 | 98.8667
-7.6 | 100.2 | 100.1 | 99.7 | 99.3 | 98.9286
-7.5 | 100.2667 | 100.1 | 99.8 | 99.4 | 99
-7.4 | 100.3 | 100.1 | 99.8 | 99.425 | 99.1
-7.3 | 100.3 | 100.1 | 99.8429 | 99.5 | 99.1
-7.2 | 100.3 | 100.2 | 99.9 | 99.5333 | 99.1
-7.1 | 100.3 | 100.2 | 99.9 | 99.6 | 99.2
-7 | 100.3429 | 100.2 | 99.9 | 99.6 | 99.2714
-6.9 | 100.4 | 100.2 | 100 | 99.6 | 99.3
-6.8 | 100.4 | 100.2667 | 100 | 99.6 | 99.3
-6.7 | 100.4 | 100.3 | 100.06 | 99.7 | 99.3667
-6.6 | 100.5 | 100.3 | 100.2 | 99.7 | 99.4875
-6.5 | 100.5 | 100.4 | 100.2 | 99.7143 | 99.5
-6.4 | 100.52 | 100.4 | 100.2167 | 99.8 | 99.5
-6.3 | 100.6 | 100.4 | 100.3 | 99.875 | 99.6
-6.2 | 100.6 | 100.5 | 100.3 | 99.9 | 99.6
-6.1 | 100.7 | 100.5 | 100.3 | 99.9571 | 99.6
-6 | 100.7 | 100.5 | 100.3 | 100 | 99.6
-5.9 | 100.7 | 100.5 | 100.3 | 100 | 99.5
-5.8 | 100.7 | 100.5 | 100.3 | 100 | 99.5
-5.7 | 100.7 | 100.5 | 100.3 | 100 | 99.5
-5.6 | 100.7 | 100.5 | 100.3 | 100 | 99.525
-5.5 | 100.7167 | 100.5 | 100.3571 | 100.0571 | 99.6
-5.4 | 100.8 | 100.58 | 100.375 | 100.1 | 99.6
-5.3 | 100.82 | 100.5571 | 100.3833 | 100.1 | 99.6667
-5.2 | 100.9 | 100.6 | 100.3 | 100.1 | 99.7
-5.1 | 100.9 | 100.6 | 100.3 | 100.02 | 99.7143
-5 | 100.9 | 100.6 | 100.3 | 100 | 99.8
-4.9 | 100.8 | 100.6 | 100.3 | 100 | 99.7
-4.8 | 100.8 | 100.6 | 100.3 | 100 | 99.7
-4.7 | 100.75 | 100.5667 | 100.3667 | 100 | 99.7143
-4.6 | 100.7 | 100.5 | 100.4 | 100 | 99.8
-4.5 | 100.7 | 100.5 | 100.4 | 100 | 99.8
-4.4 | 100.7 | 100.5 | 100.3333 | 100 | 99.8
-4.3 | 100.6857 | 100.5 | 100.3 | 100 | 99.8
-4.2 | 100.6 | 100.5 | 100.2167 | 100 | 99.8
-4.1 | 100.5571 | 100.4143 | 100.2 | 100 | 99.8
-4 | 100.5 | 100.4 | 100.1 | 99.9143 | 99.9
-3.9 | 100.5 | 100.4 | 100.1 | 99.9167 | 99.9
-3.8 | 100.5 | 100.4 | 100.1714 | 100 | 99.9
-3.7 | 100.5 | 100.4 | 100.2 | 100.1 | 99.9571
-3.6 | 100.4 | 100.4 | 100.2 | 100.175 | 100
-3.5 | 100.4 | 100.4 | 100.2 | 100.2 | 100
-3.4 | 100.4 | 100.3857 | 100.2 | 100.2 | 100.0875
-3.3 | 100.4333 | 100.3 | 100.2714 | 100.2 | 100.1
-3.2 | 100.5 | 100.34 | 100.3 | 100.1 | 100.1
-3.1 | 100.5 | 100.3 | 100.3 | 100.15 | 100.1
-3 | 100.5 | 100.3 | 100.3 | 100.2 | 100.1
-2.9 | 100.5 | 100.3 | 100.3 | 100.2429 | 100.1
-2.8 | 100.5 | 100.34 | 100.3 | 100.3 | 100.1
-2.7 | 100.4 | 100.4 | 100.3 | 100.25 | 100.05
-2.6 | 100.4 | 100.4 | 100.3 | 100.2 | 100
-2.5 | 100.45 | 100.4 | 100.3 | 100.3 | 100
-2.4 | 100.4167 | 100.4 | 100.3 | 100.2 | 100.1
-2.3 | 100.5 | 100.4 | 100.3 | 100.2 | 100.1
-2.2 | 100.5 | 100.3 | 100.3 | 100.2 | 100.1
-2.1 | 100.5 | 100.3167 | 100.3 | 100.3 | 100.1
-2 | 100.5 | 100.4 | 100.3 | 100.3 | 100.1
-1.9 | 100.4 | 100.4 | 100.2 | 100.3 | 100.1
-1.8 | 100.4 | 100.4 | 100.2 | 100.3 | 100.1
-1.7 | 100.4 | 100.4 | 100.2 | 100.2 | 100.1
-1.6 | 100.4 | 100.3 | 100.2 | 100.2 | 100.1
-1.5 | 100.4 | 100.3833 | 100.2 | 100.2875 | 100.05
-1.4 | 100.32 | 100.3 | 100.2 | 100.2 | 100
-1.3 | 100.3 | 100.2 | 100.1667 | 100.2 | 100
-1.2 | 100.3 | 100.2 | 100.1 | 100.2 | 100
-1.1 | 100.3 | 100.1 | 100.1 | 100.2 | 100
-1 | 100.2833 | 100.1 | 100.1 | 100.15 | 100
-0.9 | 100.2 | 100.05 | 100 | 100.1 | 100
-0.8 | 100.16 | 100.1 | 100 | 100.15 | 100
-0.7 | 100.1 | 100 | 100 | 100.2 | 100
-0.6 | 100.1 | 100.1 | 100 | 100.2 | 100
-0.5 | 100 | 100 | 100 | 100.1 | 100
-0.4 | 100 | 100 | 99.9 | 100.2 | 100
-0.3 | 100 | 100 | 99.9 | 100.1 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100.0143 | 100
0.3 | 100 | 100 | 99.95 | 100.1 | 100
0.4 | 100 | 100 | 99.9 | 100.1 | 100
0.5 | 100 | 100 | 100 | 100.1 | 100
0.6 | 100.0333 | 100.0667 | 100 | 100.1714 | 100
0.7 | 100.1 | 100 | 100 | 100.1556 | 100
0.8 | 100.18 | 100.1 | 100 | 100.125 | 100
0.9 | 100.2 | 100.1 | 100.0571 | 100.1 | 100
1 | 100.3 | 100.0667 | 100.1 | 100.1667 | 100
1.1 | 100.3 | 100.1 | 100.1 | 100.2 | 100
1.2 | 100.3 | 100.1833 | 100.1 | 100.2 | 100
1.3 | 100.3 | 100.2333 | 100.1167 | 100.2 | 100
1.4 | 100.3 | 100.3 | 100.2 | 100.2125 | 100
1.5 | 100.4 | 100.3 | 100.2 | 100.2625 | 100.05
1.6 | 100.4 | 100.3 | 100.2 | 100.2 | 100.1
1.7 | 100.4 | 100.4 | 100.2 | 100.22 | 100.1
1.8 | 100.4 | 100.4 | 100.2 | 100.3 | 100.1
1.9 | 100.4833 | 100.4 | 100.2 | 100.3 | 100.1
2 | 100.5 | 100.4 | 100.2667 | 100.3 | 100.1
2.1 | 100.5 | 100.3 | 100.3 | 100.3 | 100.1
2.2 | 100.5 | 100.3143 | 100.3 | 100.2714 | 100.1
2.3 | 100.5 | 100.4 | 100.3 | 100.2 | 100.1
2.4 | 100.4167 | 100.4 | 100.3 | 100.2 | 100.1
2.5 | 100.4286 | 100.4 | 100.3 | 100.2667 | 100
2.6 | 100.4 | 100.4 | 100.3 | 100.2 | 100
2.7 | 100.4167 | 100.4 | 100.3 | 100.2833 | 100.05
2.8 | 100.5 | 100.3167 | 100.3 | 100.3 | 100.1
2.9 | 100.5 | 100.3286 | 100.3 | 100.3 | 100.1
3 | 100.5 | 100.3 | 100.3 | 100.2 | 100.1
3.1 | 100.5 | 100.3333 | 100.3 | 100.1333 | 100.1
3.2 | 100.5 | 100.4 | 100.3 | 100.1286 | 100.1
3.3 | 100.4167 | 100.3 | 100.2167 | 100.2 | 100.1
3.4 | 100.4 | 100.3429 | 100.2 | 100.2 | 100.0875
3.5 | 100.4 | 100.4 | 100.2 | 100.2 | 100
3.6 | 100.4 | 100.4 | 100.2 | 100.15 | 100
3.7 | 100.5 | 100.4 | 100.2 | 100.1 | 99.9571
3.8 | 100.5 | 100.4 | 100.2 | 100 | 99.9
3.9 | 100.5 | 100.4 | 100.1 | 99.9875 | 99.9
4 | 100.5 | 100.4 | 100.1 | 99.9 | 99.9
4.1 | 100.5857 | 100.4 | 100.18 | 99.9857 | 99.8
4.2 | 100.6 | 100.5 | 100.26 | 99.9125 | 99.8
4.3 | 100.6143 | 100.5 | 100.3 | 100 | 99.8
4.4 | 100.7 | 100.5 | 100.38 | 100 | 99.8
4.5 | 100.7 | 100.5 | 100.4 | 100 | 99.8
4.6 | 100.7 | 100.5 | 100.4 | 100 | 99.8
4.7 | 100.78 | 100.6 | 100.32 | 100 | 99.7143
4.8 | 100.8 | 100.6 | 100.3 | 100 | 99.7
4.9 | 100.8 | 100.6 | 100.3 | 100 | 99.7
5 | 100.9 | 100.6 | 100.3 | 100 | 99.8
5.1 | 100.9 | 100.6 | 100.3 | 100.04 | 99.7143
5.2 | 100.9 | 100.56 | 100.3 | 100.1 | 99.7
5.3 | 100.8 | 100.6 | 100.3333 | 100.1 | 99.6667
5.4 | 100.8 | 100.52 | 100.4 | 100.1 | 99.6
5.5 | 100.7 | 100.5 | 100.4 | 100.0333 | 99.6
5.6 | 100.7 | 100.5 | 100.35 | 100 | 99.525
5.7 | 100.7 | 100.5 | 100.3 | 100 | 99.5
5.8 | 100.7 | 100.5 | 100.3 | 100 | 99.5
5.9 | 100.7 | 100.5 | 100.3 | 100 | 99.5
6 | 100.7 | 100.4167 | 100.3 | 100 | 99.6
6.1 | 100.675 | 100.5 | 100.3 | 99.9333 | 99.6
6.2 | 100.6 | 100.4833 | 100.3 | 99.9 | 99.6
6.3 | 100.6 | 100.4 | 100.3 | 99.85 | 99.6
6.4 | 100.5833 | 100.4 | 100.2667 | 99.8 | 99.5
6.5 | 100.5 | 100.38 | 100.2 | 99.7 | 99.5
6.6 | 100.5 | 100.3 | 100.2 | 99.7 | 99.4875
6.7 | 100.4571 | 100.3 | 100.1 | 99.7 | 99.3667
6.8 | 100.4 | 100.3 | 100 | 99.6 | 99.3
6.9 | 100.4 | 100.2 | 100 | 99.6 | 99.3
7 | 100.4 | 100.2 | 99.96 | 99.6 | 99.2714
7.1 | 100.3 | 100.2 | 99.9 | 99.6 | 99.2
7.2 | 100.3 | 100.2 | 99.9 | 99.5167 | 99.1
7.3 | 100.3 | 100.1 | 99.8857 | 99.5 | 99.1
7.4 | 100.2333 | 100.1 | 99.8 | 99.4 | 99.1
7.5 | 100.24 | 100.1 | 99.7667 | 99.4 | 99
7.6 | 100.2 | 100.0857 | 99.7 | 99.3 | 98.9286
7.7 | 100.2 | 100 | 99.7 | 99.3 | 98.8667
7.8 | 100.2 | 99.9125 | 99.7 | 99.2 | 98.8
7.9 | 100.2 | 99.9 | 99.6 | 99.2 | 98.7429
8 | 100.1 | 99.9 | 99.6 | 99.1 | 98.7
8.1 | 100 | 99.9 | 99.6 | 99.1 | 98.6556
8.2 | 100 | 99.8 | 99.5 | 99.025 | 98.6
8.3 | 100 | 99.8 | 99.5 | 99 | 98.5286
8.4 | 99.9 | 99.8 | 99.4 | 99 | 98.5
8.5 | 99.9 | 99.8 | 99.4 | 99 | 98.4375
8.6 | 99.9 | 99.7167 | 99.3714 | 98.9 | 98.3125
8.7 | 99.8 | 99.6571 | 99.3 | 98.9 | 98.3
8.8 | 99.8 | 99.6 | 99.3 | 98.8 | 98.2857
8.9 | 99.72 | 99.5 | 99.22 | 98.7625 | 98.2
9 | 99.65 | 99.4667 | 99.1333 | 98.7 | 98.1833
9.1 | 99.6 | 99.4 | 99.0714 | 98.5 | 98.1
9.2 | 99.5 | 99.4 | 99 | 98.42 | 98
9.3 | 99.4 | 99.3429 | 98.9714 | 98.4 | 97.9
9.4 | 99.4 | 99.28 | 98.9 | 98.4 | 97.9
9.5 | 99.3667 | 99.2 | 98.9 | 98.3429 | 97.8
9.6 | 99.3 | 99.2 | 98.9 | 98.3 | 97.8
9.7 | 99.3 | 99.1 | 98.8 | 98.2286 | 97.8
9.8 | 99.2125 | 99.1 | 98.8 | 98.1714 | 97.7
9.9 | 99.2 | 98.9857 | 98.7 | 98.0286 | 97.7
10 | 99.2 | 98.9 | 98.6143 | 98 | 97.6
10.1 | 99.1571 | 98.9 | 98.5429 | 97.9 | 97.525
10.2 | 99.1 | 98.8 | 98.5 | 97.9 | 97.4
10.3 | 99.0429 | 98.8 | 98.4167 | 97.8 | 97.3375
10.4 | 99 | 98.8 | 98.3429 | 97.8 | 97.3
10.5 | 99 | 98.7 | 98.3 | 97.8 | 97.2
10.6 | 98.9 | 98.7 | 98.2 | 97.7 | 97.075
10.7 | 98.9 | 98.6 | 98.2 | 97.675 | 97
10.8 | 98.8 | 98.6 | 98.2 | 97.5333 | 97
10.9 | 98.8 | 98.54 | 98.1 | 97.4 | 96.8714
11 | 98.7667 | 98.5 | 98.1 | 97.3625 | 96.8
11.1 | 98.7 | 98.3667 | 98 | 97.2167 | 96.8
11.2 | 98.6 | 98.3 | 97.9 | 97.1 | 96.7333
11.3 | 98.6 | 98.2167 | 97.9 | 97.1 | 96.6625
11.4 | 98.5167 | 98.16 | 97.8833 | 97 | 96.5375
11.5 | 98.4333 | 98.1 | 97.7167 | 96.9 | 96.4125
11.6 | 98.3667 | 98.1 | 97.7 | 96.9 | 96.3
11.7 | 98.2143 | 98.1 | 97.6167 | 96.8 | 96.275
11.8 | 98.1857 | 98 | 97.6 | 96.7 | 96.2
11.9 | 98.1 | 98 | 97.5 | 96.64 | 96.1
12 | 98.0667 | 97.85 | 97.4 | 96.6 | 96.0571
12.1 | 98 | 97.78 | 97.3714 | 96.52 | 96
12.2 | 97.9 | 97.7 | 97.3 | 96.5 | 96
12.3 | 97.9 | 97.6667 | 97.2 | 96.4286 | 95.9
12.4 | 97.8571 | 97.6 | 97.1286 | 96.3 | 95.8571
12.5 | 97.8 | 97.6 | 97 | 96.2286 | 95.6833
12.6 | 97.7 | 97.5167 | 96.9571 | 96.0875 | 95.6
12.7 | 97.65 | 97.4 | 96.9 | 96 | 95.4875
12.8 | 97.5714 | 97.36 | 96.8875 | 95.9875 | 95.4
12.9 | 97.5 | 97.2857 | 96.8 | 95.85 | 95.4
13 | 97.4 | 97.2 | 96.7 | 95.8 | 95.2714
13.1 | 97.4 | 97.1 | 96.6286 | 95.7 | 95.2
13.2 | 97.3 | 97 | 96.5 | 95.6333 | 95.1
13.3 | 97.2 | 97 | 96.4 | 95.5 | 94.9833
13.4 | 97.2 | 96.9 | 96.3 | 95.5 | 94.9
13.5 | 97.1 | 96.8 | 96.2333 | 95.4167 | 94.8429
13.6 | 97.0333 | 96.7125 | 96.14 | 95.2571 | 94.8
13.7 | 97 | 96.7 | 96.1 | 95.2 | 94.6875
13.8 | 97 | 96.7 | 96.0143 | 95.1 | 94.5625
13.9 | 96.9 | 96.5571 | 96 | 95.0143 | 94.4375
14 | 96.8 | 96.5 | 95.8833 | 94.9429 | 94.4
14.1 | 96.8 | 96.5 | 95.8 | 94.7444 | 94.2429
14.2 | 96.7167 | 96.4 | 95.7286 | 94.7 | 94.1
14.3 | 96.6333 | 96.3143 | 95.6667 | 94.6 | 94.1
14.4 | 96.5625 | 96.275 | 95.6 | 94.5 | 94.0167
14.5 | 96.5 | 96.2 | 95.5286 | 94.35 | 93.8714
14.6 | 96.4 | 96.1 | 95.4 | 94.2 | 93.7286
14.7 | 96.3 | 96 | 95.3333 | 94.1333 | 93.5667
14.8 | 96.3 | 95.9833 | 95.3 | 94.0714 | 93.5
14.9 | 96.3 | 95.9 | 95.25 | 93.9286 | 93.4
15 | 96.2 | 95.8 | 95.1 | 93.9 | 93.3
15.1 | 96.1333 | 95.6714 | 95.0167 | 93.7625 | 93.2571
15.2 | 96.0667 | 95.6 | 94.9 | 93.6375 | 93.1143
15.3 | 95.9143 | 95.55 | 94.8 | 93.5125 | 92.98
15.4 | 95.9 | 95.48 | 94.6857 | 93.45 | 92.8143
15.5 | 95.9 | 95.4 | 94.6 | 93.3 | 92.6833
15.6 | 95.7333 | 95.3429 | 94.5 | 93.2667 | 92.5286
15.7 | 95.6 | 95.2 | 94.42 | 93.125 | 92.4
15.8 | 95.5 | 95.12 | 94.3 | 93 | 92.3
15.9 | 95.44 | 95.0333 | 94.2 | 92.9 | 92.225
16 | 95.3667 | 94.9 | 94.1167 | 92.8 | 92.0857
16.1 | 95.3 | 94.85 | 94 | 92.6714 | 91.9429
16.2 | 95.2571 | 94.8 | 93.9 | 92.5375 | 91.9
16.3 | 95.2 | 94.65 | 93.8333 | 92.4 | 91.76
16.4 | 95.02 | 94.5 | 93.7875 | 92.35 | 91.6
16.5 | 94.9 | 94.48 | 93.65 | 92.2 | 91.5429
16.6 | 94.8143 | 94.4 | 93.5 | 92.1 | 91.4125
16.7 | 94.7667 | 94.24 | 93.4571 | 91.9714 | 91.2889
16.8 | 94.6 | 94.1714 | 93.3 | 91.8167 | 91.175
16.9 | 94.5 | 94.1 | 93.12 | 91.6833 | 91.05
17 | 94.4 | 93.92 | 92.94 | 91.5286 | 90.9143
17.1 | 94.28 | 93.8 | 92.8667 | 91.4 | 90.7333
17.2 | 94.1143 | 93.7 | 92.7143 | 91.275 | 90.56
17.3 | 94.06 | 93.5333 | 92.5667 | 91.15 | 90.38
17.4 | 94 | 93.3833 | 92.4 | 91.0143 | 90.3
17.5 | 93.8667 | 93.3 | 92.2333 | 90.875 | 90.25
17.6 | 93.8 | 93.15 | 92.1 | 90.75 | 90.2
17.7 | 93.7 | 93.0833 | 92 | 90.7 | 90.05
17.8 | 93.5833 | 92.9286 | 91.8143 | 90.55 | 89.8429
17.9 | 93.4167 | 92.7667 | 91.675 | 90.3333 | 89.7
18 | 93.3571 | 92.6 | 91.55 | 90.2 | 89.5333
18.1 | 93.225 | 92.4333 | 91.4143 | 90.0286 | 89.3429
18.2 | 93.1667 | 92.3625 | 91.2714 | 89.8571 | 89.1375
18.3 | 93 | 92.2375 | 91.1167 | 89.7 | 89.0125
18.4 | 92.8375 | 92.0714 | 90.96 | 89.6429 | 88.8875
18.5 | 92.7667 | 91.9 | 90.8125 | 89.4 | 88.7333
18.6 | 92.6 | 91.7 | 90.5571 | 89.24 | 88.55
18.7 | 92.4333 | 91.6167 | 90.4143 | 89.16 | 88.425
18.8 | 92.26 | 91.45 | 90.2875 | 88.9 | 88.2571
18.9 | 92.1667 | 91.3 | 90.15 | 88.64 | 88.1
19 | 91.9 | 91.1167 | 89.9833 | 88.55 | 88.0167
19.1 | 91.6667 | 90.95 | 89.7333 | 88.3833 | 87.9333
19.2 | 91.375 | 90.7333 | 89.5667 | 88.2375 | 87.7625
19.3 | 91.1 | 90.5714 | 89.4 | 88.0667 | 87.575
19.4 | 90.6143 | 90.3333 | 89.1714 | 87.9143 | 87.3857
19.5 | 89.75 | 89.95 | 89.0167 | 87.6667 | 87.22
19.6 | 88.7 | 89.7 | 88.7333 | 87.4714 | 87.0333
19.7 | 87.2 | 89.4 | 88.5143 | 87.275 | 86.8857
19.8 | 84.6375 | 89.075 | 88.175 | 87.1 | 86.7429
19.9 | 80.45 | 88.72 | 88 | 86.8667 | 86.5667
20 | 74.4667 | 88.1143 | 87.65 | 86.6 | 86.3375
20.1 | 67.2 | 87.5 | 87.28 | 86.3778 | 86.2
20.2 | 58.8 | 86.6714 | 87.1 | 86.225 | 86.075
20.3 | 49.5571 | 85.4 | 86.9333 | 86.0375 | 85.9
20.4 | 40.9143 | 83.55 | 86.62 | 85.825 | 85.6286
20.5 | 33.175 | 80.8714 | 86.2667 | 85.675 | 85.4556
20.6 | 26.4 | 77 | 86.1 | 85.425 | 85.2625
20.7 | 21.4333 | 71.3571 | 85.5857 | 85.2167 | 85.075
20.8 | 17.4 | 64.1667 | 85.2 | 85.0333 | 84.9125
20.9 | 15.1667 | 56.0375 | 84.55 | 84.75 | 84.6875
21 | 13.7286 | 48.0375 | 84.02 | 84.375 | 84.525
21.1 | 12.85 | 40.4375 | 83.2286 | 84.25 | 84.2333
21.2 | 12.25 | 33.6 | 82.4 | 84.05 | 84.05
21.3 | 11.8333 | 27.8667 | 81.0571 | 83.7 | 83.9
21.4 | 11.54 | 23.2 | 79.2 | 83.375 | 83.7286
21.5 | 11.2333 | 19.75 | 76.6 | 83.1571 | 83.5
21.6 | 10.9833 | 17.2857 | 72.6143 | 82.9667 | 83.25
21.7 | 10.8 | 15.8 | 67.2667 | 82.7 | 82.94
21.8 | 10.675 | 14.7571 | 61.4 | 82.3143 | 82.7667
21.9 | 10.4667 | 14.1 | 53.9 | 81.9714 | 82.625
22 | 10.3 | 13.5571 | 46.8857 | 81.6857 | 82.4
22.1 | 10.2 | 13.16 | 40.3143 | 81.5 | 82.2143
22.2 | 10.0286 | 12.7833 | 34.2 | 81.1286 | 81.9286
22.3 | 9.8833 | 12.5571 | 29.275 | 80.7429 | 81.6857
22.4 | 9.7375 | 12.26 | 25.0875 | 80.4333 | 81.4857
22.5 | 9.5875 | 12.0667 | 22.0714 | 80.0714 | 81.2
22.6 | 9.5 | 11.86 | 20.0714 | 79.7 | 81.05
22.7 | 9.3833 | 11.6 | 18.7 | 79.2 | 80.7714
22.8 | 9.3 | 11.4333 | 17.7857 | 78.7 | 80.55
22.9 | 9.1333 | 11.22 | 17 | 78.125 | 80.3286
23 | 9.0833 | 11 | 16.45 | 77.5 | 80.1143
23.1 | 9 | 10.8571 | 16.0125 | 76.8286 | 79.8286
23.2 | 8.85 | 10.7143 | 15.65 | 75.8167 | 79.45
23.3 | 8.8 | 10.5667 | 15.2333 | 74.4 | 79.2556
23.4 | 8.7 | 10.4 | 14.95 | 72.5 | 79.0571
23.5 | 8.6 | 10.2286 | 14.6333 | 69.8857 | 78.7333
23.6 | 8.5833 | 10.16 | 14.3286 | 66.3857 | 78.3875
23.7 | 8.5 | 9.9667 | 14.12 | 62.15 | 78.075
23.8 | 8.4 | 9.8 | 13.8833 | 57.3167 | 77.7
23.9 | 8.4 | 9.65 | 13.675 | 51.8 | 77.3857
24 | 8.3 | 9.5833 | 13.4 | 46.2333 | 77.1667
24.1 | 8.3 | 9.4167 | 13.225 | 40.75 | 76.9
24.2 | 8.2 | 9.3 | 12.9889 | 35.9 | 76.54
24.3 | 8.1857 | 9.2286 | 12.8143 | 31.78 | 76.12
24.4 | 8.1 | 9.1 | 12.5667 | 28.3 | 75.7333
24.5 | 8 | 9.0571 | 12.4143 | 25.7 | 75.5
24.6 | 7.9 | 8.9143 | 12.1667 | 23.9875 | 74.9714
24.7 | 7.85 | 8.8 | 12.025 | 22.675 | 74.56
24.8 | 7.8 | 8.7167 | 11.86 | 21.6875 | 74.1
24.9 | 7.7375 | 8.6625 | 11.6833 | 21.0375 | 73.525
25 | 7.7 | 8.6 | 11.5167 | 20.45 | 72.8125
25.1 | 7.6429 | 8.5 | 11.2833 | 19.95 | 72
25.2 | 7.5714 | 8.45 | 11.1375 | 19.55 | 70.9625
25.3 |  | 8.4 | 10.9857 | 19.25 | 69.5
25.4 |  | 8.32 | 10.9 | 18.8125 | 67.75
25.5 |  | 8.2571 | 10.8 | 18.4 | 65.4571
25.6 |  | 8.2 | 10.6714 | 18.1333 | 62.2571
25.7 |  | 8.1 | 10.5167 | 17.9143 | 58.3167
25.8 |  | 8.025 | 10.3571 | 17.5571 | 54.0143
25.9 |  | 7.9571 | 10.3 | 17.2857 | 49.2
26 |  |  | 10.175 | 17.1 | 44.5571
26.1 |  |  | 10.0429 | 16.8 | 40.2
26.2 |  |  | 9.9 | 16.54 | 36.3
26.3 |  |  | 9.8 | 16.3333 | 32.6333
26.4 |  |  | 9.7125 | 16.1 | 29.9
26.5 |  |  | 9.6778 | 15.8556 | 27.7
26.6 |  |  | 9.5625 | 15.6571 | 26.2
26.7 |  |  | 9.4286 | 15.5 | 25.1
26.8 |  |  | 9.3 | 15.2429 | 24.1857
26.9 |  |  | 9.2143 | 15 | 23.4
27 |  |  |  | 14.825 | 23
27.1 |  |  |  | 14.6375 | 22.5
27.2 |  |  |  | 14.4 | 22.1
27.3 |  |  |  | 14.2286 | 21.825
27.4 |  |  |  | 14.0571 | 21.45
27.5 |  |  |  | 13.85 | 21.075
27.6 |  |  |  | 13.66 | 20.8222
27.7 |  |  |  | 13.4857 | 20.475
27.8 |  |  |  | 13.35 | 20.1875
27.9 |  |  |  | 13.15 | 19.875
28 |  |  |  | 13 | 19.6
28.1 |  |  |  | 12.8571 | 19.3143
28.2 |  |  |  | 12.65 | 19.1143
28.3 |  |  |  | 12.5111 | 18.9
28.4 |  |  |  | 12.3875 | 18.7333
28.5 |  |  |  | 12.2625 | 18.4667
28.6 |  |  |  | 12.1375 | 18.2286
28.7 |  |  |  | 12.0125 | 18.0667
28.8 |  |  |  | 11.8875 | 17.925
28.9 |  |  |  | 11.7625 | 17.6667
29 |  |  |  |  | 17.375
29.1 |  |  |  |  | 17.2
29.2 |  |  |  |  | 17.02
29.3 |  |  |  |  | 16.8
29.4 |  |  |  |  | 16.5857
29.5 |  |  |  |  | 16.45
29.6 |  |  |  |  | 16.25
29.7 |  |  |  |  | 16
29.8 |  |  |  |  | 15.775
29.9 |  |  |  |  | 15.65
30 |  |  |  |  | 15.525
30.1 |  |  |  |  | 15.4
30.2 |  |  |  |  | 15.1778
30.3 |  |  |  |  | 15.025
30.4 |  |  |  |  | 14.8375
30.5 |  |  |  |  | 14.7125
30.6 |  |  |  |  | 14.575
30.7 |  |  |  |  | 14.3625
30.8 |  |  |  |  | 14.2375
30.9 |  |  |  |  | 14.0833

## Wedge 45 Output Factors
Output factors
Wedge 45 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20
| 3 | 0.881 | 0.893 | 0.9 | 0.909 | 0.916 | 0.922 | 0.927
| 4 | 0.893 | 0.906 | 0.915 | 0.927 | 0.936 | 0.945 | 0.95
| 5 | 0.901 | 0.916 | 0.927 | 0.939 | 0.951 | 0.962 | 0.969
| 7 | 0.912 | 0.931 | 0.944 | 0.961 | 0.975 | 0.99 | 1
| 10 | 0.923 | 0.944 | 0.96 | 0.981 | 1 | 1.019 | 1.032
| 15 | 0.932 | 0.955 | 0.973 | 0.999 | 1.023 | 1.048 | 1.067
| 20 | 0.937 | 0.962 | 0.982 | 1.01 | 1.038 | 1.068 | 1.091
| 30 | 0.944 | 0.971 | 0.992 | 1.024 | 1.057 | 1.095 | 1.124
| 40 | 0.948 | 0.976 | 0.999 | 1.033 | 1.068 | 1.11 | 1.143

## Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 39.2 | 45 | 50 | 56.4
0.1 | 44.3 | 49.8 | 54.4667 | 61.2
0.2 | 52.575 | 57.2 | 62.0333 | 67.625
0.3 | 62.98 | 67.4667 | 70.4 | 76.1
0.4 | 72.94 | 75.8667 | 78.6 | 82.3
0.5 | 79.98 | 82.6333 | 84.75 | 87.45
0.6 | 85.4167 | 87.2 | 89.3 | 91.4
0.7 | 89.5 | 91.2 | 92.54 | 94.2
0.8 | 92.9 | 93.775 | 94.725 | 96.2
0.9 | 95.34 | 95.975 | 96.575 | 97.6
1 | 97.1 | 97.45 | 98.15 | 98.65
1.1 | 98.4 | 98.7 | 98.975 | 99.3
1.2 | 99.0667 | 99.42 | 99.55 | 99.8333
1.3 | 99.7 | 99.7333 | 99.9 | 99.975
1.4 | 99.925 | 99.9 | 100 | 100
1.5 | 100 | 100 | 100 | 100
1.6 | 99.9667 | 100 | 100 | 99.9
1.7 | 99.9 | 99.9 | 99.9 | 99.7
1.8 | 99.7 | 99.775 | 99.7667 | 99.5
1.9 | 99.4 | 99.525 | 99.525 | 99.2
2 | 99.12 | 99.175 | 99.175 | 98.8
2.1 | 98.8 | 98.875 | 98.85 | 98.5
2.2 | 98.4667 | 98.6 | 98.35 | 98.175
2.3 | 97.9 | 98.1333 | 98.1 | 97.925
2.4 | 97.3333 | 97.6 | 97.74 | 97.55
2.5 | 96.9 | 97.3 | 97.3 | 97.1
2.6 | 96.425 | 96.9 | 96.9 | 96.75
2.7 | 95.96 | 96.5 | 96.55 | 96.48
2.8 | 95.48 | 96.06 | 96.1 | 96
2.9 | 94.82 | 95.7 | 95.8 | 95.65
3 | 94.4 | 95.35 | 95.4 | 95.3
3.1 | 94.05 | 94.925 | 94.975 | 95
3.2 | 93.5 | 94.475 | 94.625 | 94.5667
3.3 | 93.05 | 94.1 | 94.25 | 94.15
3.4 | 92.6333 | 93.675 | 93.925 | 93.85
3.5 | 92.1 | 93.225 | 93.575 | 93.3667
3.6 | 91.7 | 92.78 | 93.15 | 93.075
3.7 | 91.08 | 92.38 | 92.76 | 92.65
3.8 | 90.7333 | 92.1 | 92.3 | 92.15
3.9 | 90.2571 | 91.6 | 91.9 | 91.9
4 | 89.76 | 91.26 | 91.55 | 91.6
4.1 | 89.3667 | 90.775 | 91.1 | 91.25
4.2 | 88.8667 | 90.375 | 90.8 | 90.8
4.3 | 88.35 | 90.025 | 90.42 | 90.45
4.4 | 87.85 | 89.65 | 90.0667 | 90.1
4.5 | 87.35 | 89.2 | 89.66 | 89.86
4.6 | 86.8 | 88.775 | 89.275 | 89.525
4.7 | 86.45 | 88.45 | 88.9333 | 89.1
4.8 | 86 | 88 | 88.4 | 88.7
4.9 | 85.58 | 87.6 | 88.2 | 88.15
5 | 85.18 | 87.275 | 87.76 | 87.9
5.1 | 84.55 | 86.8 | 87.325 | 87.6
5.2 | 84.1 | 86.5 | 86.975 | 87.25
5.3 | 83.7 | 86.0333 | 86.68 | 86.975
5.4 | 83.2 | 85.65 | 86.4 | 86.65
5.5 | 82.7 | 85.325 | 86 | 86.34
5.6 | 82.2 | 84.8 | 85.7 | 85.85
5.7 | 81.8667 | 84.5 | 85.3 | 85.3
5.8 | 81.45 | 84.04 | 85 | 85.05
5.9 | 81.05 | 83.8 | 84.54 | 84.7667
6 | 80.5 | 83.45 | 84.2667 | 84.4
6.1 | 80.14 | 82.9 | 83.75 | 84.125
6.2 | 79.46 | 82.52 | 83.3667 | 83.7
6.3 | 79.0667 | 82.2333 | 83 | 83.45
6.4 | 78.82 | 81.9 | 82.7 | 83
6.5 | 78.35 | 81.4333 | 82.32 | 82.65
6.6 | 77.98 | 81.05 | 81.95 | 82.375
6.7 | 77.5 | 80.75 | 81.625 | 82.05
6.8 | 77.08 | 80.2333 | 81.4 | 81.7
6.9 | 76.6667 | 79.78 | 81 | 81.4
7 | 76.1 | 79.45 | 80.5 | 80.8333
7.1 | 75.8 | 79.1 | 80.2 | 80.4667
7.2 | 75.25 | 78.875 | 79.7 | 80.2333
7.3 | 74.82 | 78.375 | 79.3 | 79.875
7.4 | 74.36 | 77.925 | 79.1 | 79.625
7.5 | 73.9 | 77.575 | 78.725 | 79.325
7.6 | 73.7 | 77.2 | 78.3 | 78.925
7.7 | 73.275 | 76.8 | 78.05 | 78.6333
7.8 | 72.82 | 76.46 | 77.725 | 78.175
7.9 | 72.45 | 75.95 | 77.325 | 77.9
8 | 72.1167 | 75.725 | 77 | 77.55
8.1 | 71.68 | 75.3 | 76.6 | 77.225
8.2 | 71.25 | 74.9 | 76.25 | 76.85
8.3 | 70.72 | 74.6333 | 75.94 | 76.54
8.4 | 70.22 | 74.2667 | 75.6 | 76.2
8.5 | 69.82 | 73.8667 | 75.2 | 75.8
8.6 | 69.575 | 73.52 | 74.875 | 75.5
8.7 | 69.15 | 73.25 | 74.45 | 75.1
8.8 | 68.7333 | 72.8667 | 74.15 | 74.725
8.9 | 68.4 | 72.5 | 73.9333 | 74.4667
9 | 67.9333 | 72.1 | 73.7 | 74.34
9.1 | 67.575 | 71.75 | 73.35 | 73.7667
9.2 | 67.12 | 71.425 | 72.9 | 73.55
9.3 | 66.725 | 70.95 | 72.55 | 73.1
9.4 | 66.3 | 70.575 | 72.15 | 72.85
9.5 | 66 | 70.225 | 71.8 | 72.5
9.6 | 65.6 | 69.925 | 71.5 | 72.275
9.7 | 65.1 | 69.6 | 71.1 | 71.85
9.8 | 64.95 | 69.15 | 70.8 | 71.575
9.9 | 64.625 | 68.9 | 70.55 | 71.4
10 | 64.2 | 68.65 | 70.2 | 71
10.1 | 63.8 | 68.2 | 69.875 | 70.7
10.2 | 63.3 | 67.85 | 69.6667 | 70.3
10.3 | 63.0333 | 67.45 | 69.4 | 70.025
10.4 | 62.7 | 67.2 | 69.05 | 69.8333
10.5 | 62.24 | 66.925 | 68.8 | 69.3
10.6 | 61.8 | 66.6 | 68.4667 | 69.05
10.7 | 61.7 | 66.1 | 68.1 | 68.725
10.8 | 61.35 | 65.86 | 67.8 | 68.55
10.9 | 60.94 | 65.525 | 67.5667 | 68.2667
11 | 60.5 | 65.25 | 67.1 | 67.8333
11.1 | 60.2 | 64.95 | 66.8 | 67.55
11.2 | 59.8 | 64.6333 | 66.5667 | 67.225
11.3 | 59.45 | 64.175 | 66.22 | 66.96
11.4 | 59.15 | 63.9 | 65.925 | 66.64
11.5 | 58.8714 | 63.64 | 65.55 | 66.375
11.6 | 58.5667 | 63.3 | 65.125 | 66.1333
11.7 | 58.1429 | 62.9 | 64.8667 | 65.7333
11.8 | 57.8333 | 62.475 | 64.5 | 65.5
11.9 | 57.44 | 62.3 | 64.2667 | 65.2
12 | 57.12 | 62 | 63.9 | 64.8333
12.1 | 56.825 | 61.65 | 63.6 | 64.5667
12.2 | 56.5667 | 61.35 | 63.3667 | 64.38
12.3 | 56.22 | 60.9 | 63.1 | 64.05
12.4 | 55.85 | 60.7 | 62.8 | 63.7
12.5 | 55.55 | 60.3 | 62.46 | 63.34
12.6 | 55.25 | 60.1 | 62.125 | 63
12.7 | 54.9 | 59.875 | 61.8 | 62.85
12.8 | 54.65 | 59.5 | 61.5333 | 62.625
12.9 | 54.2429 | 59.2667 | 61.2667 | 62.275
13 | 54 | 58.9 | 61.1 | 62.04
13.1 | 53.65 | 58.6 | 60.7333 | 61.7333
13.2 | 53.36 | 58.3 | 60.4 | 61.325
13.3 | 53.16 | 57.9667 | 60.075 | 61.1
13.4 | 52.8333 | 57.7 | 59.825 | 60.9
13.5 | 52.3 | 57.45 | 59.475 | 60.6
13.6 | 52 | 57.125 | 59.25 | 60.2
13.7 | 51.8 | 56.775 | 59 | 59.925
13.8 | 51.5 | 56.54 | 58.78 | 59.775
13.9 | 51.2429 | 56.2 | 58.45 | 59.525
14 | 50.88 | 56.0333 | 58.1 | 59.18
14.1 | 50.58 | 55.65 | 57.85 | 58.9
14.2 | 50.36 | 55.35 | 57.6 | 58.6
14.3 | 50.075 | 55.05 | 57.275 | 58.325
14.4 | 49.76 | 54.8 | 57.025 | 58.075
14.5 | 49.5 | 54.5333 | 56.775 | 57.8333
14.6 | 49.2333 | 54.2 | 56.4667 | 57.55
14.7 | 48.9667 | 54 | 56.2 | 57.3
14.8 | 48.6 | 53.5667 | 55.9333 | 57.05
14.9 | 48.2857 | 53.35 | 55.6667 | 56.82
15 | 48 | 53.1 | 55.3 | 56.55
15.1 | 47.74 | 52.7 | 55.1 | 56.2
15.2 | 47.5 | 52.5 | 54.9 | 55.9
15.3 | 47.2667 | 52.25 | 54.5333 | 55.74
15.4 | 47.0286 | 52.02 | 54.35 | 55.5
15.5 | 46.725 | 51.725 | 54.05 | 55.2333
15.6 | 46.4667 | 51.4333 | 53.8 | 54.9667
15.7 | 46.22 | 51.2 | 53.62 | 54.8
15.8 | 45.85 | 50.95 | 53.375 | 54.525
15.9 | 45.6 | 50.675 | 53.05 | 54.2
16 | 45.325 | 50.425 | 52.8 | 53.875
16.1 | 45 | 50.25 | 52.5 | 53.725
16.2 | 44.7667 | 49.84 | 52.275 | 53.5
16.3 | 44.6 | 49.6 | 52.025 | 53.35
16.4 | 44.3333 | 49.4 | 51.7 | 53
16.5 | 44.0667 | 49.2333 | 51.55 | 52.75
16.6 | 43.8 | 48.9667 | 51.3 | 52.6
16.7 | 43.56 | 48.7 | 51.025 | 52.3
16.8 | 43.26 | 48.375 | 50.85 | 52.1
16.9 | 43.05 | 48.125 | 50.575 | 51.7667
17 | 42.85 | 47.88 | 50.4 | 51.5
17.1 | 42.62 | 47.6333 | 50.15 | 51.25
17.2 | 42.3143 | 47.34 | 49.78 | 51.05
17.3 | 42.05 | 47.1 | 49.6 | 50.8
17.4 | 41.8 | 46.7333 | 49.3 | 50.6
17.5 | 41.6 | 46.6 | 49.075 | 50.3
17.6 | 41.3 | 46.375 | 48.75 | 50.05
17.7 | 41.12 | 46.08 | 48.6 | 49.8
17.8 | 40.825 | 45.9333 | 48.3333 | 49.56
17.9 | 40.6833 | 45.7667 | 48.1 | 49.325
18 | 40.5286 | 45.5 | 47.9333 | 49.175
18.1 | 40.225 | 45.1333 | 47.65 | 48.925
18.2 | 39.975 | 44.9 | 47.4 | 48.7
18.3 | 39.8 | 44.75 | 47.275 | 48.5
18.4 | 39.55 | 44.5 | 47.05 | 48.325
18.5 | 39.325 | 44.275 | 46.8 | 48.1
18.6 | 39.0333 | 44.15 | 46.6 | 47.82
18.7 | 38.82 | 43.9 | 46.4 | 47.55
18.8 | 38.6667 | 43.6 | 46.1 | 47.32
18.9 | 38.4333 | 43.4 | 45.8 | 47.1333
19 | 38.16 | 43.1 | 45.55 | 46.9
19.1 | 37.9667 | 42.825 | 45.35 | 46.6667
19.2 | 37.75 | 42.675 | 45.1 | 46.4
19.3 | 37.5833 | 42.325 | 44.8667 | 46.1333
19.4 | 37.4167 | 42.2 | 44.6667 | 46
19.5 | 37.25 | 41.9 | 44.5 | 45.85
19.6 | 36.9667 | 41.66 | 44.1667 | 45.5
19.7 | 36.7167 | 41.525 | 44 | 45.375
19.8 | 36.4667 | 41.275 | 43.8 | 45.2
19.9 | 36.3 | 41.025 | 43.6333 | 44.9
20 | 36.1333 | 40.88 | 43.375 | 44.8
20.1 | 35.85 | 40.7 | 43.125 | 44.525
20.2 | 35.65 | 40.5 | 42.9 | 44.3
20.3 | 35.45 | 40.2667 | 42.7 | 44.1
20.4 | 35.2667 | 40.1 | 42.5 | 44
20.5 | 35.1 | 39.8 | 42.3 | 43.7
20.6 | 34.94 | 39.5 | 42.1 | 43.55
20.7 | 34.725 | 39.3667 | 41.9 | 43.3333
20.8 | 34.5333 | 39.2 | 41.72 | 43.075
20.9 | 34.2714 | 39.04 | 41.575 | 42.925
21 | 34.1167 | 38.8 | 41.34 | 42.775
21.1 | 34.02 | 38.55 | 41.1 | 42.6
21.2 | 33.7 | 38.4 | 40.9333 | 42.4
21.3 | 33.4857 | 38.175 | 40.675 | 42.2
21.4 | 33.2667 | 38.025 | 40.5 | 41.95
21.5 | 33.1 | 37.825 | 40.2 | 41.725
21.6 | 32.9667 | 37.6 | 40.0667 | 41.575
21.7 | 32.75 | 37.4 | 39.8 | 41.28
21.8 | 32.62 | 37.2333 | 39.6 | 41.1
21.9 | 32.45 | 36.9667 | 39.5 | 40.98
22 | 32.26 | 36.7667 | 39.35 | 40.75
22.1 | 31.96 | 36.6 | 39.2 | 40.5333
22.2 | 31.84 | 36.425 | 38.95 | 40.3333
22.3 | 31.7 | 36.3 | 38.725 | 40.175
22.4 | 31.5429 | 36.16 | 38.475 | 39.9667
22.5 | 31.3167 | 35.925 | 38.325 | 39.8333
22.6 | 31.15 | 35.68 | 38.2 | 39.575
22.7 | 31 | 35.55 | 37.95 | 39.5
22.8 | 30.85 | 35.4 | 37.7 | 39.2
22.9 | 30.6667 | 35.175 | 37.575 | 39.1
23 | 30.4286 | 35.025 | 37.3 | 38.9
23.1 | 30.2667 | 34.8 | 37.2 | 38.7
23.2 | 30.1 | 34.6 | 37.0667 | 38.5
23.3 | 30 | 34.44 | 36.9 | 38.3
23.4 | 29.8 | 34.2 | 36.75 | 38.1
23.5 | 29.68 | 34.05 | 36.52 | 37.9333
23.6 | 29.5 | 33.9 | 36.275 | 37.75
23.7 | 29.22 | 33.7 | 36.1 | 37.5
23.8 | 29.1 | 33.5 | 35.9 | 37.4
23.9 | 28.9167 | 33.3 | 35.7 | 37.3
24 | 28.8 | 33.1667 | 35.6 | 37.1
24.1 | 28.6286 | 32.9333 | 35.375 | 36.8333
24.2 | 28.45 | 32.8 | 35.2 | 36.65
24.3 | 28.3833 | 32.65 | 35 | 36.55
24.4 | 28.2167 | 32.5 | 34.9 | 36.3333
24.5 | 28.05 | 32.3 | 34.7 | 36.15
24.6 | 27.8333 | 32.2 | 34.6 | 36.025
24.7 | 27.7 | 31.975 | 34.45 | 35.875
24.8 | 27.55 | 31.725 | 34.2 | 35.6667
24.9 | 27.3857 | 31.6 | 34 | 35.5
25 | 27.2333 | 31.475 | 33.8 | 35.4
25.1 | 27.1429 | 31.3 | 33.7 | 35.2
25.2 | 26.9833 | 31.2 | 33.5333 | 34.95
25.3 | 26.78 | 31 | 33.375 | 34.8
25.4 | 26.68 | 30.75 | 33.2 | 34.66
25.5 | 26.48 | 30.625 | 33.1 | 34.5
25.6 | 26.28 | 30.475 | 32.95 | 34.4
25.7 | 26.15 | 30.3 | 32.8 | 34.2333
25.8 | 26.04 | 30.2 | 32.55 | 34.1
25.9 | 25.9 | 30.0333 | 32.4 | 33.85
26 | 25.7 | 29.9 | 32.2333 | 33.7
26.1 | 25.6 | 29.68 | 32 | 33.45
26.2 | 25.45 | 29.5 | 31.9 | 33.3333
26.3 | 25.3 | 29.4 | 31.74 | 33.2
26.4 | 25.2 | 29.2667 | 31.6 | 33
26.5 | 25.0167 | 29.1 | 31.5 | 32.9
26.6 | 24.95 | 29.04 | 31.325 | 32.8
26.7 | 24.8 | 28.8667 | 31.2 | 32.65
26.8 | 24.6333 | 28.6 | 31 | 32.5
26.9 | 24.4714 | 28.5 | 30.8 | 32.375
27 | 24.4 | 28.4 | 30.7 | 32.125
27.1 | 24.2 | 28.275 | 30.6 | 31.9
27.2 | 24.15 | 28.025 | 30.45 | 31.7667
27.3 | 24 | 27.9 | 30.3 | 31.6
27.4 | 23.8 | 27.75 | 30.05 | 31.4
27.5 | 23.6833 | 27.7 | 29.925 | 31.3
27.6 | 23.5286 | 27.55 | 29.775 | 31.1333
27.7 | 23.4 | 27.3 | 29.68 | 30.95
27.8 | 23.3 | 27.25 | 29.5 | 30.8
27.9 | 23.2 | 27.1 | 29.3 | 30.6333
28 | 23.1 | 27 | 29.2 | 30.5
28.1 | 23 | 26.8 | 29.1 | 30.4
28.2 | 22.8571 | 26.68 | 28.9667 | 30.3
28.3 | 22.7 | 26.5 | 28.775 | 30.175
28.4 | 22.5333 | 26.4 | 28.625 | 30
28.5 | 22.425 | 26.3 | 28.5 | 29.8
28.6 | 22.32 | 26.15 | 28.35 | 29.7
28.7 | 22.2333 | 26 | 28.3 | 29.525
28.8 | 22.1 | 25.8 | 28.2 | 29.4
28.9 | 21.95 | 25.65 | 27.975 | 29.2
29 | 21.8 | 25.58 | 27.725 | 29.1
29.1 | 21.7333 | 25.45 | 27.68 | 28.9
29.2 | 21.6667 | 25.325 | 27.55 | 28.75
29.3 | 21.5 | 25.1667 | 27.4 | 28.6667
29.4 | 21.32 | 25.1 | 27.2667 | 28.6
29.5 | 21.25 | 24.9 | 27.1 | 28.5
29.6 | 21.1857 | 24.8 | 27 | 28.3
29.7 | 21.1 | 24.7 | 26.825 | 28.14
29.8 | 20.9 | 24.575 | 26.7667 | 28
29.9 | 20.8 | 24.4 | 26.6 | 27.9
30 | 20.6 | 24.275 | 26.5 | 27.8

## SSD80 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 38 | 44.9 | 51.4 | 60.9
0.1 | 42.65 | 49.3 | 55.66 | 64.9
0.2 | 48.5 | 55.2 | 61.7 | 70.1
0.3 | 58.85 | 63.95 | 69.225 | 76.3875
0.4 | 68.4 | 73.2 | 77.075 | 82.925
0.5 | 77.1 | 80.35 | 83.875 | 88.2375
0.6 | 83.04 | 86.05 | 88.925 | 92.15
0.7 | 87.975 | 90.3 | 92.4667 | 94.8333
0.8 | 91.725 | 93.46 | 95 | 96.75
0.9 | 94.45 | 95.9 | 97.02 | 98.175
1 | 96.35 | 97.3 | 98.36 | 99.025
1.1 | 97.925 | 98.575 | 99.275 | 99.64
1.2 | 98.925 | 99.35 | 99.7 | 99.9
1.3 | 99.55 | 99.7714 | 99.9 | 100
1.4 | 99.875 | 99.9571 | 100 | 100
1.5 | 100 | 100 | 100 | 99.9
1.6 | 100 | 99.975 | 99.925 | 99.7
1.7 | 99.8 | 99.75 | 99.8 | 99.45
1.8 | 99.55 | 99.6125 | 99.4333 | 99.1
1.9 | 99.3333 | 99.2 | 99.12 | 98.6857
2 | 98.7625 | 98.9 | 98.68 | 98.3
2.1 | 98.35 | 98.3 | 98.2 | 98
2.2 | 97.9286 | 97.95 | 97.9 | 97.55
2.3 | 97.5 | 97.4 | 97.4 | 97.1
2.4 | 96.9 | 96.95 | 96.925 | 96.65
2.5 | 96.45 | 96.55 | 96.62 | 96.2
2.6 | 95.8 | 96.05 | 96.025 | 95.75
2.7 | 95.35 | 95.55 | 95.64 | 95.4
2.8 | 94.7 | 95.22 | 95.2 | 94.9
2.9 | 94.2333 | 94.65 | 94.7857 | 94.4
3 | 93.74 | 94.2333 | 94.2429 | 93.96
3.1 | 93.175 | 93.75 | 93.85 | 93.6
3.2 | 92.65 | 93.3571 | 93.35 | 93.25
3.3 | 92.0667 | 92.8571 | 92.875 | 92.75
3.4 | 91.525 | 92.5 | 92.65 | 92.2667
3.5 | 91.08 | 91.9 | 92 | 91.8
3.6 | 90.46 | 91.5 | 91.6 | 91.5
3.7 | 89.95 | 91 | 91.2 | 91
3.8 | 89.5 | 90.5667 | 90.75 | 90.6
3.9 | 88.9 | 90 | 90.3 | 90.14
4 | 88.3 | 89.6 | 89.88 | 89.65
4.1 | 87.9667 | 89.1 | 89.4 | 89.26
4.2 | 87.3 | 88.6 | 88.88 | 88.95
4.3 | 86.8 | 88.2 | 88.525 | 88.4
4.4 | 86.3 | 87.75 | 88.05 | 87.9
4.5 | 85.8 | 87.3 | 87.65 | 87.6
4.6 | 85.36 | 86.85 | 87.28 | 87.075
4.7 | 84.8 | 86.48 | 86.8 | 86.75
4.8 | 84.28 | 86.0667 | 86.4 | 86.35
4.9 | 83.8 | 85.5333 | 86 | 85.9333
5 | 83.275 | 84.975 | 85.46 | 85.4333
5.1 | 82.85 | 84.65 | 85.05 | 85
5.2 | 82.225 | 84.175 | 84.6667 | 84.625
5.3 | 81.68 | 83.85 | 84.15 | 84.15
5.4 | 81.075 | 83.275 | 83.925 | 83.825
5.5 | 80.65 | 82.8125 | 83.34 | 83.45
5.6 | 80.1 | 82.2857 | 82.875 | 83.0333
5.7 | 79.75 | 81.9 | 82.5 | 82.7
5.8 | 79.2667 | 81.5 | 81.975 | 82.25
5.9 | 78.6 | 81.1 | 81.7 | 81.8
6 | 78.16 | 80.5 | 81.3 | 81.425
6.1 | 77.7 | 80.125 | 80.9 | 81.1
6.2 | 77.25 | 79.825 | 80.4 | 80.6
6.3 | 76.7 | 79.375 | 80.1 | 80.2
6.4 | 76.4 | 79.025 | 79.7 | 79.8
6.5 | 75.8 | 78.5 | 79.25 | 79.6
6.6 | 75.3 | 78.05 | 78.85 | 79.04
6.7 | 74.8333 | 77.6 | 78.5 | 78.66
6.8 | 74.4667 | 77.35 | 78.15 | 78.2667
6.9 | 74 | 76.8 | 77.7 | 77.875
7 | 73.54 | 76.4 | 77.2 | 77.5
7.1 | 73.15 | 76.025 | 76.925 | 77.1
7.2 | 72.75 | 75.5667 | 76.5 | 76.7429
7.3 | 72.25 | 75.1667 | 76.1 | 76.4
7.4 | 71.8 | 74.6667 | 75.7 | 75.92
7.5 | 71.4 | 74.2 | 75.2625 | 75.6
7.6 | 70.9 | 73.8333 | 74.85 | 75.2
7.7 | 70.48 | 73.48 | 74.54 | 74.8
7.8 | 70 | 73.05 | 74.2 | 74.475
7.9 | 69.5429 | 72.74 | 73.8 | 74.0667
8 | 69.1 | 72.3 | 73.3 | 73.7
8.1 | 68.775 | 71.95 | 72.95 | 73.35
8.2 | 68.4333 | 71.5 | 72.6 | 73.1
8.3 | 67.95 | 71.12 | 72.3333 | 72.65
8.4 | 67.45 | 70.725 | 71.8 | 72.275
8.5 | 67.15 | 70.35 | 71.5 | 71.9
8.6 | 66.6 | 69.88 | 71.15 | 71.5
8.7 | 66.2 | 69.4 | 70.7 | 71.175
8.8 | 65.7667 | 69.15 | 70.4 | 70.9
8.9 | 65.4 | 68.75 | 70.1 | 70.4
9 | 65.05 | 68.425 | 69.6 | 70.12
9.1 | 64.5 | 67.9333 | 69.3 | 69.8
9.2 | 64.125 | 67.5667 | 68.975 | 69.4333
9.3 | 63.8333 | 67.175 | 68.55 | 69.05
9.4 | 63.35 | 66.9 | 68.175 | 68.65
9.5 | 63 | 66.5 | 67.85 | 68.25
9.6 | 62.55 | 66.0667 | 67.4857 | 68.0222
9.7 | 62.1 | 65.75 | 67.125 | 67.72
9.8 | 61.725 | 65.325 | 66.75 | 67.4
9.9 | 61.3 | 64.95 | 66.4333 | 67
10 | 60.95 | 64.6 | 66.05 | 66.675
10.1 | 60.58 | 64.2 | 65.7 | 66.25
10.2 | 60.25 | 63.9 | 65.4 | 65.9
10.3 | 59.8 | 63.5333 | 64.98 | 65.55
10.4 | 59.4667 | 63.12 | 64.7 | 65.4
10.5 | 58.95 | 62.8 | 64.3333 | 64.95
10.6 | 58.75 | 62.35 | 63.9 | 64.7
10.7 | 58.45 | 62.02 | 63.62 | 64.3
10.8 | 57.875 | 61.62 | 63.25 | 63.95
10.9 | 57.55 | 61.325 | 63 | 63.575
11 | 57.3 | 60.975 | 62.55 | 63.28
11.1 | 56.78 | 60.625 | 62.2 | 63.05
11.2 | 56.5 | 60.35 | 61.8667 | 62.6
11.3 | 56.2 | 59.95 | 61.52 | 62.3
11.4 | 55.8 | 59.575 | 61.225 | 62
11.5 | 55.55 | 59.4 | 60.925 | 61.7
11.6 | 55.1 | 59.1 | 60.6333 | 61.35
11.7 | 54.8 | 58.6 | 60.35 | 61.1667
11.8 | 54.45 | 58.3 | 59.95 | 60.825
11.9 | 54.1 | 57.9333 | 59.65 | 60.55
12 | 53.65 | 57.64 | 59.3667 | 60.225
12.1 | 53.425 | 57.25 | 59.075 | 59.9
12.2 | 53.0333 | 57 | 58.8 | 59.5667
12.3 | 52.775 | 56.6 | 58.4 | 59.275
12.4 | 52.44 | 56.325 | 58 | 58.925
12.5 | 52.175 | 55.95 | 57.7333 | 58.6
12.6 | 51.78 | 55.74 | 57.375 | 58.325
12.7 | 51.5 | 55.3 | 57.14 | 58
12.8 | 51.16 | 55.1 | 56.8 | 57.65
12.9 | 50.8 | 54.85 | 56.5 | 57.28
13 | 50.5 | 54.425 | 56.3 | 57.0333
13.1 | 50.2667 | 54.2167 | 55.95 | 56.875
13.2 | 49.8 | 53.7778 | 55.6 | 56.5667
13.3 | 49.6 | 53.55 | 55.3 | 56.25
13.4 | 49.3 | 53.25 | 55 | 55.95
13.5 | 48.975 | 52.95 | 54.72 | 55.6
13.6 | 48.7 | 52.6 | 54.45 | 55.375
13.7 | 48.3667 | 52.36 | 54.125 | 55.08
13.8 | 48.05 | 52 | 53.875 | 54.8
13.9 | 47.7571 | 51.7 | 53.6 | 54.55
14 | 47.3857 | 51.525 | 53.375 | 54.2
14.1 | 47.14 | 51.15 | 53.025 | 53.95
14.2 | 46.9 | 50.825 | 52.775 | 53.7
14.3 | 46.6 | 50.6 | 52.45 | 53.4
14.4 | 46.35 | 50.2 | 52.2333 | 53.125
14.5 | 46.0429 | 49.9 | 51.9 | 52.9
14.6 | 45.8 | 49.65 | 51.54 | 52.625
14.7 | 45.5 | 49.3 | 51.24 | 52.275
14.8 | 45.25 | 49.15 | 51.05 | 52.0333
14.9 | 44.9 | 48.8 | 50.7 | 51.74
15 | 44.7625 | 48.5 | 50.4 | 51.4
15.1 | 44.425 | 48.175 | 50.1 | 51.14
15.2 | 44.05 | 47.925 | 49.875 | 50.9
15.3 | 43.68 | 47.75 | 49.625 | 50.7667
15.4 | 43.5 | 47.5 | 49.38 | 50.375
15.5 | 43.2333 | 47.2 | 49.08 | 50.2
15.6 | 43 | 46.9333 | 48.85 | 49.975
15.7 | 42.7 | 46.5857 | 48.6 | 49.65
15.8 | 42.45 | 46.35 | 48.325 | 49.4
15.9 | 42.175 | 46.1 | 48.05 | 49.05
16 | 42 | 45.9 | 47.82 | 48.9
16.1 | 41.66 | 45.6667 | 47.575 | 48.6
16.2 | 41.3875 | 45.325 | 47.2667 | 48.4
16.3 | 41.24 | 45.075 | 47 | 48.05
16.4 | 40.9 | 44.88 | 46.7625 | 47.9
16.5 | 40.6 | 44.54 | 46.4625 | 47.65
16.6 | 40.425 | 44.3 | 46.3286 | 47.34
16.7 | 40.2 | 44.05 | 46.05 | 47.075
16.8 | 39.925 | 43.8 | 45.8 | 46.85
16.9 | 39.775 | 43.475 | 45.55 | 46.6
17 | 39.4667 | 43.225 | 45.3 | 46.35
17.1 | 39.2167 | 42.975 | 45.0571 | 46.1
17.2 | 38.925 | 42.7333 | 44.725 | 45.925
17.3 | 38.6333 | 42.54 | 44.475 | 45.65
17.4 | 38.48 | 42.2857 | 44.225 | 45.4
17.5 | 38.25 | 42.05 | 44.075 | 45.1
17.6 | 38.1 | 41.7857 | 43.78 | 44.975
17.7 | 37.75 | 41.6222 | 43.6333 | 44.8
17.8 | 37.6 | 41.3667 | 43.4 | 44.475
17.9 | 37.3 | 41.15 | 43.15 | 44.3
18 | 37.1 | 40.9 | 42.9 | 44.08
18.1 | 36.8 | 40.66 | 42.65 | 43.85
18.2 | 36.625 | 40.425 | 42.4 | 43.6
18.3 | 36.475 | 40.2 | 42.275 | 43.35
18.4 | 36.225 | 40.025 | 42.025 | 43.2
18.5 | 36.075 | 39.8 | 41.7167 | 42.96
18.6 | 35.74 | 39.5 | 41.475 | 42.675
18.7 | 35.6 | 39.2 | 41.32 | 42.425
18.8 | 35.35 | 39.05 | 41 | 42.2429
18.9 | 35.1 | 38.8 | 40.86 | 41.95
19 | 34.95 | 38.65 | 40.625 | 41.775
19.1 | 34.7 | 38.42 | 40.38 | 41.54
19.2 | 34.525 | 38.1 | 40.225 | 41.4
19.3 | 34.3 | 37.96 | 39.96 | 41.2333
19.4 | 34.05 | 37.74 | 39.8 | 41
19.5 | 33.8571 | 37.5429 | 39.575 | 40.75
19.6 | 33.66 | 37.3625 | 39.3 | 40.5
19.7 | 33.44 | 37.2 | 39.1 | 40.25
19.8 | 33.2833 | 36.9429 | 38.9 | 40.08
19.9 | 33.1 | 36.6571 | 38.7 | 39.825
20 | 32.94 | 36.45 | 38.5 | 39.6571
20.1 | 32.6 | 36.2333 | 38.24 | 39.425
20.2 | 32.45 | 36.05 | 38 | 39.275
20.3 | 32.3 | 35.8 | 37.85 | 39.06
20.4 | 32 | 35.575 | 37.68 | 38.8333
20.5 | 31.9 | 35.4 | 37.45 | 38.64
20.6 | 31.625 | 35.2 | 37.2667 | 38.5
20.7 | 31.45 | 35.025 | 37 | 38.2333
20.8 | 31.3 | 34.8 | 36.85 | 38.05
20.9 | 31.1 | 34.6 | 36.5667 | 37.84
21 | 30.9 | 34.425 | 36.4 | 37.66
21.1 | 30.75 | 34.2 | 36.25 | 37.4143
21.2 | 30.5 | 34.05 | 36 | 37.3167
21.3 | 30.32 | 33.8 | 35.86 | 37.15
21.4 | 30.275 | 33.6 | 35.625 | 36.95
21.5 | 30.04 | 33.4 | 35.45 | 36.7
21.6 | 29.82 | 33.2556 | 35.2 | 36.5167
21.7 | 29.6375 | 33.0571 | 35.075 | 36.3
21.8 | 29.42 | 32.8667 | 34.9 | 36.1
21.9 | 29.3 | 32.6833 | 34.7 | 35.9
22 | 29.1625 | 32.5375 | 34.525 | 35.725
22.1 | 29.0167 | 32.325 | 34.3 | 35.575
22.2 | 28.775 | 32.2 | 34.1 | 35.325
22.3 | 28.54 | 31.9333 | 33.9 | 35.175
22.4 | 28.375 | 31.775 | 33.8 | 35
22.5 | 28.2333 | 31.65 | 33.6 | 34.9
22.6 | 28.1 | 31.5 | 33.4 | 34.65
22.7 | 27.875 | 31.25 | 33.3 | 34.5
22.8 | 27.7 | 31.1 | 33.075 | 34.26
22.9 | 27.5333 | 30.9 | 32.92 | 34.1
23 | 27.4 | 30.725 | 32.675 | 33.9
23.1 | 27.3 | 30.575 | 32.5 | 33.775
23.2 | 27.125 | 30.325 | 32.3 | 33.6
23.3 | 26.94 | 30.18 | 32.1333 | 33.3857
23.4 | 26.8 | 30.05 | 32 | 33.26
23.5 | 26.6 | 29.875 | 31.85 | 33.025
23.6 | 26.4333 | 29.7 | 31.6 | 32.8875
23.7 | 26.3 | 29.525 | 31.5 | 32.74
23.8 | 26.15 | 29.4 | 31.3 | 32.6
23.9 | 26 | 29.25 | 31.1 | 32.4
24 | 25.86 | 29 | 31 | 32.22
24.1 | 25.7 | 28.85 | 30.85 | 32.075
24.2 | 25.5 | 28.7 | 30.6 | 31.825
24.3 | 25.4 | 28.5667 | 30.4667 | 31.675
24.4 | 25.3 | 28.5 | 30.32 | 31.525
24.5 | 25.1167 | 28.275 | 30.175 | 31.4
24.6 | 24.9875 | 28.05 | 30.025 | 31.225
24.7 | 24.825 | 27.9 | 29.775 | 31.075
24.8 | 24.675 | 27.775 | 29.625 | 30.95
24.9 | 24.4667 | 27.625 | 29.4833 | 30.7
25 | 24.3 | 27.48 | 29.3167 | 30.55
25.1 | 24.2 | 27.3 | 29.18 | 30.4
25.2 | 24 | 27.1833 | 29 | 30.26
25.3 | 23.95 | 27 | 28.8 | 30.05
25.4 | 23.8 | 26.8 | 28.65 | 29.925
25.5 | 23.6 | 26.68 | 28.5 | 29.85
25.6 | 23.5 | 26.5 | 28.3875 | 29.7
25.7 | 23.35 | 26.3667 | 28.2625 | 29.4375
25.8 | 23.2 | 26.2143 | 28.075 | 29.3125
25.9 | 23.05 | 26.1 | 27.9 | 29.1875
26 | 22.9 | 25.95 | 27.8 | 29.0571
26.1 | 22.8 | 25.8 | 27.6 | 28.9333
26.2 | 22.7 | 25.6667 | 27.5 | 28.775
26.3 | 22.5625 | 25.5 | 27.35 | 28.525
26.4 | 22.5 | 25.35 | 27.2 | 28.475
26.5 | 22.3 | 25.2 | 27 | 28.325
26.6 | 22.125 | 25.1 | 26.92 | 28.175
26.7 | 21.975 | 24.975 | 26.775 | 27.98
26.8 | 21.9 | 24.8 | 26.625 | 27.8333
26.9 | 21.8 | 24.6571 | 26.4167 | 27.7
27 | 21.625 | 24.5 | 26.3 | 27.54
27.1 | 21.6 | 24.3667 | 26.1714 | 27.4
27.2 | 21.425 | 24.24 | 26.0444 | 27.3
27.3 | 21.2333 | 24.1 | 25.925 | 27.1333
27.4 | 21.175 | 23.9375 | 25.7667 | 27
27.5 | 21 | 23.8125 | 25.6 | 26.9
27.6 | 20.94 | 23.725 | 25.45 | 26.775
27.7 | 20.775 | 23.6 | 25.4 | 26.525
27.8 | 20.7 | 23.525 | 25.2 | 26.4333
27.9 | 20.5333 | 23.3 | 25.1 | 26.3
28 | 20.4 | 23.2 | 24.95 | 26.15
28.1 | 20.3 | 23.075 | 24.8 | 26.05
28.2 | 20.2 | 22.9 | 24.7 | 25.9
28.3 | 20.1 | 22.875 | 24.55 | 25.8
28.4 | 19.96 | 22.625 | 24.4 | 25.625
28.5 | 19.8 | 22.6 | 24.3 | 25.5
28.6 | 19.7 | 22.425 | 24.2 | 25.4
28.7 | 19.6 | 22.3 | 24 | 25.2375
28.8 | 19.5 | 22.2 | 23.9 | 25.1125
28.9 | 19.4 | 22.05 | 23.8 | 24.9875
29 | 19.275 | 21.92 | 23.6 | 24.8571
29.1 | 19.15 | 21.8 | 23.5 | 24.7
29.2 | 19 | 21.7 | 23.375 | 24.675
29.3 | 18.9 | 21.6 | 23.3 | 24.5
29.4 | 18.8 | 21.5 | 23.2 | 24.3286
29.5 | 18.6333 | 21.3833 | 23.04 | 24.3
29.6 | 18.6 | 21.275 | 22.9 | 24.1
29.7 | 18.55 | 21.0875 | 22.7444 | 24
29.8 | 18.4778 | 21 | 22.6 | 23.8875
29.9 | 18.35 | 20.9 | 22.55 | 23.75
30 | 18.175 | 20.8125 | 22.4714 | 23.625

## SSD120 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 39.3 | 44.8 | 49.4 | 55.1
0.1 | 44.7 | 49.85 | 54.275 | 60.3
0.2 | 52.9 | 57.7 | 61.5 | 66.5
0.3 | 64.2 | 68.34 | 70.975 | 75.5
0.4 | 73.4 | 75.875 | 78.5 | 81.775
0.5 | 80.25 | 82.82 | 84.45 | 87.16
0.6 | 85.8 | 87.6 | 88.875 | 90.85
0.7 | 89.25 | 91.14 | 92.05 | 93.65
0.8 | 92.5333 | 93.925 | 94.75 | 95.8
0.9 | 94.7667 | 95.5 | 96.48 | 97.2
1 | 96.725 | 97.125 | 97.675 | 98.42
1.1 | 98.1 | 98.55 | 98.775 | 99.2286
1.2 | 99.025 | 99.15 | 99.45 | 99.6625
1.3 | 99.4667 | 99.65 | 99.7 | 99.9
1.4 | 99.775 | 99.875 | 99.95 | 100
1.5 | 100 | 100 | 100 | 100
1.6 | 100 | 100 | 100 | 100
1.7 | 100 | 99.9 | 99.9 | 99.8
1.8 | 99.85 | 99.7 | 99.775 | 99.72
1.9 | 99.6 | 99.5 | 99.45 | 99.55
2 | 99.3 | 99.25 | 99.2 | 99.2
2.1 | 99 | 98.9 | 99.0286 | 99.05
2.2 | 98.6 | 98.55 | 98.8 | 98.75
2.3 | 98.3 | 98.3 | 98.3125 | 98.35
2.4 | 97.85 | 97.875 | 98.05 | 97.95
2.5 | 97.25 | 97.6 | 97.7 | 97.7
2.6 | 96.8 | 97.325 | 97.4 | 97.4
2.7 | 96.5 | 96.96 | 97.08 | 96.98
2.8 | 96.14 | 96.525 | 96.7 | 96.7
2.9 | 95.65 | 96.1667 | 96.4 | 96.4
3 | 95.25 | 95.85 | 96.02 | 96
3.1 | 94.65 | 95.38 | 95.7143 | 95.7
3.2 | 94.25 | 95 | 95.4286 | 95.35
3.3 | 93.85 | 94.825 | 95.05 | 94.95
3.4 | 93.575 | 94.3 | 94.775 | 94.7
3.5 | 93.225 | 93.975 | 94.425 | 94.3
3.6 | 92.85 | 93.65 | 93.925 | 94.1
3.7 | 92.15 | 93.1 | 93.525 | 93.7
3.8 | 91.7 | 92.8 | 93.275 | 93.4
3.9 | 91.4 | 92.5 | 92.95 | 93.1667
4 | 90.8 | 92.05 | 92.6 | 92.8
4.1 | 90.35 | 91.7667 | 92.2 | 92.4
4.2 | 89.9 | 91.46 | 91.95 | 92
4.3 | 89.5 | 91.05 | 91.5 | 91.775
4.4 | 89 | 90.6 | 91.2 | 91.525
4.5 | 88.75 | 90.26 | 90.8 | 91.15
4.6 | 88.3 | 89.9333 | 90.475 | 90.75
4.7 | 87.95 | 89.6571 | 90.1833 | 90.425
4.8 | 87.52 | 89.2667 | 89.8 | 90.1
4.9 | 87.125 | 89 | 89.4 | 89.6667
5 | 86.45 | 88.56 | 89.1571 | 89.3667
5.1 | 86.14 | 88.05 | 88.8333 | 89
5.2 | 85.65 | 87.725 | 88.55 | 88.7333
5.3 | 85.4 | 87.35 | 88.125 | 88.45
5.4 | 84.7 | 86.95 | 87.8125 | 88.125
5.5 | 84.3 | 86.6 | 87.3111 | 87.7
5.6 | 83.95 | 86.3 | 87 | 87.4
5.7 | 83.5 | 86 | 86.55 | 87.0667
5.8 | 83.1 | 85.5 | 86.3 | 86.8714
5.9 | 82.5625 | 85.1 | 86.05 | 86.58
6 | 82.1875 | 84.95 | 85.6667 | 86.275
6.1 | 81.875 | 84.54 | 85.3 | 85.85
6.2 | 81.4111 | 84.25 | 85.02 | 85.5
6.3 | 81.1 | 83.875 | 84.75 | 85.15
6.4 | 80.5857 | 83.625 | 84.425 | 84.8
6.5 | 80.1 | 83.15 | 84.075 | 84.48
6.6 | 79.75 | 82.78 | 83.825 | 84.2
6.7 | 79.25 | 82.4 | 83.475 | 83.8333
6.8 | 78.875 | 82.15 | 83.125 | 83.56
6.9 | 78.6 | 81.75 | 82.8 | 83.32
7 | 78.2 | 81.35 | 82.45 | 82.9333
7.1 | 77.75 | 81.0333 | 82.2 | 82.64
7.2 | 77.1 | 80.6 | 81.85 | 82.35
7.3 | 76.95 | 80.3 | 81.45 | 81.975
7.4 | 76.72 | 80.05 | 80.975 | 81.68
7.5 | 76.25 | 79.4667 | 80.55 | 81.3
7.6 | 75.75 | 79.2 | 80.4 | 80.9286
7.7 | 75.35 | 79 | 80.125 | 80.6667
7.8 | 74.94 | 78.725 | 79.8 | 80.3667
7.9 | 74.7 | 78.05 | 79.55 | 80.1
8 | 74.35 | 77.775 | 79.22 | 79.8
8.1 | 73.85 | 77.55 | 78.875 | 79.45
8.2 | 73.5333 | 77.3 | 78.4 | 79.1
8.3 | 73.05 | 76.9 | 78.1 | 78.8
8.4 | 72.5333 | 76.5 | 77.825 | 78.625
8.5 | 72.1667 | 76.2 | 77.575 | 78.25
8.6 | 71.9 | 75.7 | 77.15 | 77.94
8.7 | 71.45 | 75.375 | 76.9 | 77.55
8.8 | 71.1333 | 75.05 | 76.5 | 77.25
8.9 | 70.825 | 74.9 | 76.2 | 76.82
9 | 70.35 | 74.55 | 76 | 76.55
9.1 | 69.9667 | 74.2 | 75.6 | 76.225
9.2 | 69.6 | 73.925 | 75.25 | 75.975
9.3 | 69.1 | 73.48 | 74.85 | 75.625
9.4 | 68.85 | 73.1 | 74.675 | 75.275
9.5 | 68.6 | 72.78 | 74.38 | 74.9375
9.6 | 68.12 | 72.55 | 74 | 74.68
9.7 | 67.82 | 72.1 | 73.525 | 74.45
9.8 | 67.35 | 71.75 | 73.3 | 74.0667
9.9 | 67.1 | 71.5 | 72.9 | 73.725
10 | 66.75 | 71.25 | 72.7 | 73.35
10.1 | 66.2 | 70.9 | 72.45 | 73.05
10.2 | 65.8 | 70.5 | 72.025 | 72.8
10.3 | 65.5 | 70.15 | 71.8333 | 72.5333
10.4 | 65.3 | 69.8 | 71.45 | 72.225
10.5 | 64.8857 | 69.45 | 71.225 | 72.0667
10.6 | 64.6 | 69.2 | 71.14 | 71.7
10.7 | 64.16 | 68.85 | 70.75 | 71.35
10.8 | 63.8286 | 68.6 | 70.4 | 71.1
10.9 | 63.55 | 68.34 | 70.15 | 70.85
11 | 63.225 | 68 | 69.825 | 70.6
11.1 | 62.85 | 67.7 | 69.45 | 70.3
11.2 | 62.5 | 67.3 | 69.025 | 69.9714
11.3 | 62.15 | 66.8 | 68.725 | 69.775
11.4 | 61.8 | 66.4 | 68.475 | 69.4
11.5 | 61.475 | 66.15 | 68.2444 | 69.125
11.6 | 61.15 | 65.9 | 68 | 68.875
11.7 | 60.75 | 65.475 | 67.725 | 68.6667
11.8 | 60.325 | 65.25 | 67.3667 | 68.375
11.9 | 60.075 | 64.975 | 67.2 | 68.1
12 | 59.75 | 64.725 | 66.925 | 67.85
12.1 | 59.3 | 64.48 | 66.5 | 67.64
12.2 | 59.1 | 64.1 | 66.175 | 67.2
12.3 | 58.8333 | 63.7 | 66 | 66.95
12.4 | 58.4667 | 63.35 | 65.7 | 66.625
12.5 | 58.1 | 63.2 | 65.425 | 66.475
12.6 | 57.8 | 62.85 | 65.1 | 66.125
12.7 | 57.5 | 62.6 | 64.84 | 65.775
12.8 | 57.35 | 62.2667 | 64.55 | 65.525
12.9 | 57 | 62 | 64.3 | 65.275
13 | 56.7 | 61.65 | 64 | 65.1
13.1 | 56.35 | 61.4 | 63.7 | 64.8
13.2 | 55.9571 | 61.12 | 63.46 | 64.45
13.3 | 55.8143 | 60.9 | 63.225 | 64.1
13.4 | 55.5 | 60.65 | 62.975 | 63.85
13.5 | 55.1875 | 60.225 | 62.74 | 63.7
13.6 | 54.95 | 59.95 | 62.5 | 63.35
13.7 | 54.6625 | 59.7 | 62.05 | 63.1
13.8 | 54.325 | 59.4 | 61.7 | 62.875
13.9 | 54.0125 | 59.15 | 61.5 | 62.55
14 | 53.7 | 58.8 | 61.2286 | 62.35
14.1 | 53.4 | 58.45 | 60.9857 | 62.1
14.2 | 53.2 | 58.2 | 60.675 | 61.8
14.3 | 52.85 | 57.9 | 60.425 | 61.5
14.4 | 52.55 | 57.8 | 60.225 | 61.16
14.5 | 52.3 | 57.475 | 59.9 | 61
14.6 | 52.075 | 57.125 | 59.64 | 60.775
14.7 | 51.725 | 56.875 | 59.375 | 60.525
14.8 | 51.475 | 56.625 | 59.025 | 60.1625
14.9 | 51.225 | 56.4 | 58.78 | 59.825
15 | 50.875 | 56 | 58.55 | 59.575
15.1 | 50.6 | 55.7 | 58.3 | 59.3444
15.2 | 50.38 | 55.55 | 58.075 | 59.2
15.3 | 49.9 | 55.2 | 57.725 | 58.925
15.4 | 49.75 | 54.86 | 57.475 | 58.6556
15.5 | 49.4 | 54.6 | 57.225 | 58.5167
15.6 | 49.2667 | 54.375 | 57 | 58.225
15.7 | 49.05 | 54.125 | 56.8333 | 57.975
15.8 | 48.825 | 54.05 | 56.5 | 57.8125
15.9 | 48.45 | 53.625 | 56.3 | 57.4625
16 | 48.125 | 53.4 | 56 | 57.125
16.1 | 47.975 | 53.25 | 55.8 | 56.8889
16.2 | 47.68 | 53 | 55.6 | 56.65
16.3 | 47.4286 | 52.65 | 55.35 | 56.4
16.4 | 47.2 | 52.4 | 55.2 | 56.2
16.5 | 46.95 | 52.15 | 54.775 | 55.975
16.6 | 46.7 | 52 | 54.45 | 55.8
16.7 | 46.4667 | 51.85 | 54.175 | 55.6
16.8 | 46.2 | 51.5 | 54.025 | 55.375
16.9 | 45.975 | 51.25 | 53.88 | 55.125
17 | 45.725 | 51 | 53.65 | 54.875
17.1 | 45.475 | 50.72 | 53.4 | 54.525
17.2 | 45.225 | 50.425 | 53.15 | 54.225
17.3 | 44.875 | 50.275 | 53 | 54.1
17.4 | 44.7333 | 49.95 | 52.725 | 53.95
17.5 | 44.475 | 49.7 | 52.4111 | 53.7667
17.6 | 44.125 | 49.525 | 52.2286 | 53.52
17.7 | 43.9 | 49.4 | 51.9667 | 53.25
17.8 | 43.7 | 48.9 | 51.575 | 53.025
17.9 | 43.475 | 48.625 | 51.325 | 52.75
18 | 43.24 | 48.5 | 51.2 | 52.425
18.1 | 43 | 48.25 | 51.1 | 52.275
18.2 | 42.8 | 48.1 | 50.8 | 52.1
18.3 | 42.55 | 47.875 | 50.5667 | 51.9
18.4 | 42.4 | 47.575 | 50.375 | 51.65
18.5 | 42.15 | 47.325 | 50.125 | 51.38
18.6 | 41.925 | 47.075 | 49.8 | 51.15
18.7 | 41.675 | 46.8 | 49.625 | 50.9
18.8 | 41.425 | 46.575 | 49.375 | 50.7143
18.9 | 41.1 | 46.425 | 49.125 | 50.5778
19 | 40.925 | 46.1 | 49 | 50.325
19.1 | 40.7 | 45.95 | 48.825 | 50.02
19.2 | 40.4667 | 45.7 | 48.5 | 49.8333
19.3 | 40.2714 | 45.5 | 48.3 | 49.6
19.4 | 40.0429 | 45.24 | 48 | 49.3
19.5 | 39.88 | 45.1 | 47.7333 | 49.1
19.6 | 39.6333 | 44.9 | 47.4667 | 48.9
19.7 | 39.3875 | 44.7 | 47.34 | 48.725
19.8 | 39.2625 | 44.475 | 47.1222 | 48.525
19.9 | 39.1 | 44.225 | 46.96 | 48.3333
20 | 38.84 | 44 | 46.75 | 48.2
20.1 | 38.5 | 43.825 | 46.4875 | 47.85
20.2 | 38.275 | 43.6 | 46.3143 | 47.7
20.3 | 38.15 | 43.375 | 46.05 | 47.4
20.4 | 37.95 | 43.125 | 45.825 | 47.2
20.5 | 37.8 | 42.8667 | 45.7 | 47
20.6 | 37.575 | 42.725 | 45.5 | 46.8
20.7 | 37.4111 | 42.5 | 45.325 | 46.55
20.8 | 37.175 | 42.375 | 45.1 | 46.325
20.9 | 36.925 | 42.125 | 44.975 | 46.2
21 | 36.675 | 41.875 | 44.8 | 46.025
21.1 | 36.5125 | 41.5833 | 44.55 | 45.875
21.2 | 36.28 | 41.4286 | 44.25 | 45.7
21.3 | 36.1 | 41.2 | 44.1 | 45.5
21.4 | 36 | 41.05 | 43.9 | 45.32
21.5 | 35.85 | 40.9 | 43.8 | 45.05
21.6 | 35.7 | 40.65 | 43.65 | 44.8
21.7 | 35.45 | 40.5 | 43.3 | 44.66
21.8 | 35.3 | 40.25 | 43.175 | 44.425
21.9 | 35.175 | 40.02 | 42.925 | 44.2875
22 | 34.925 | 39.875 | 42.675 | 44.125
22.1 | 34.7 | 39.7 | 42.525 | 43.9375
22.2 | 34.525 | 39.5 | 42.375 | 43.8125
22.3 | 34.4 | 39.375 | 42.2 | 43.4875
22.4 | 34.3 | 39.2 | 41.9 | 43.3625
22.5 | 34 | 39 | 41.7 | 43.1333
22.6 | 33.75 | 38.75 | 41.5 | 42.925
22.7 | 33.5 | 38.5 | 41.25 | 42.75
22.8 | 33.35 | 38.35 | 41.05 | 42.6
22.9 | 33.22 | 38.2 | 40.9 | 42.275
23 | 33 | 37.9 | 40.8 | 42.125
23.1 | 32.9 | 37.8 | 40.575 | 41.975
23.2 | 32.775 | 37.55 | 40.4 | 41.8
23.3 | 32.525 | 37.425 | 40.2 | 41.675
23.4 | 32.3 | 37.275 | 39.95 | 41.425
23.5 | 32.2 | 37.125 | 39.775 | 41.2
23.6 | 32.0286 | 36.875 | 39.6 | 41.04
23.7 | 31.9 | 36.625 | 39.4 | 40.9
23.8 | 31.775 | 36.5 | 39.25 | 40.75
23.9 | 31.5625 | 36.3333 | 39.1 | 40.6
24 | 31.4375 | 36.18 | 38.9 | 40.35
24.1 | 31.225 | 36 | 38.75 | 40.3
24.2 | 31.08 | 35.775 | 38.6 | 40.1
24.3 | 30.9 | 35.6125 | 38.4 | 39.825
24.4 | 30.7 | 35.5 | 38.2 | 39.6
24.5 | 30.55 | 35.3 | 38 | 39.5
24.6 | 30.4 | 35.15 | 37.85 | 39.275
24.7 | 30.1667 | 34.84 | 37.7 | 39.1
24.8 | 30.1 | 34.675 | 37.45 | 38.9667
24.9 | 30 | 34.525 | 37.3 | 38.9
25 | 29.8 | 34.4 | 37.1 | 38.65
25.1 | 29.625 | 34.3 | 36.95 | 38.5
25.2 | 29.4 | 34.075 | 36.7429 | 38.25
25.3 | 29.225 | 34 | 36.6 | 38.125
25.4 | 29.1 | 33.775 | 36.4429 | 37.8833
25.5 | 29 | 33.6 | 36.2714 | 37.675
25.6 | 28.8 | 33.5 | 36.05 | 37.5125
25.7 | 28.625 | 33.3 | 35.975 | 37.46
25.8 | 28.5 | 33.15 | 35.825 | 37.3222
25.9 | 28.35 | 32.925 | 35.6 | 37.1
26 | 28.125 | 32.8 | 35.5 | 36.875
26.1 | 28 | 32.7 | 35.25 | 36.7
26.2 | 27.8833 | 32.475 | 35.1 | 36.525
26.3 | 27.675 | 32.2 | 34.9 | 36.4
26.4 | 27.5125 | 32.1333 | 34.8 | 36.275
26.5 | 27.5 | 32 | 34.6 | 36.1
26.6 | 27.3429 | 31.9 | 34.5 | 35.9
26.7 | 27.1375 | 31.675 | 34.3 | 35.8
26.8 | 27 | 31.525 | 34.125 | 35.625
26.9 | 26.9 | 31.4 | 33.9667 | 35.475
27 | 26.7 | 31.1875 | 33.85 | 35.4
27.1 | 26.6 | 31.0625 | 33.6889 | 35.2
27.2 | 26.4 | 30.875 | 33.55 | 35
27.3 | 26.3 | 30.7125 | 33.35 | 34.8
27.4 | 26.25 | 30.575 | 33.225 | 34.6667
27.5 | 26.1 | 30.4 | 33.1 | 34.5
27.6 | 25.9 | 30.3 | 32.875 | 34.425
27.7 | 25.8 | 30.15 | 32.75 | 34.2
27.8 | 25.6 | 30 | 32.55 | 34.1
27.9 | 25.42 | 29.85 | 32.45 | 33.9
28 | 25.4 | 29.7 | 32.3 | 33.7875
28.1 | 25.275 | 29.6 | 32.2 | 33.7
28.2 | 25.1333 | 29.45 | 32.025 | 33.475
28.3 | 24.9333 | 29.2667 | 31.9 | 33.3
28.4 | 24.8 | 29.15 | 31.75 | 33.25
28.5 | 24.7 | 29.025 | 31.5 | 33.125
28.6 | 24.6 | 28.875 | 31.4 | 32.95
28.7 | 24.5 | 28.7 | 31.3 | 32.8333
28.8 | 24.4 | 28.575 | 31.1 | 32.7
28.9 | 24.2333 | 28.4 | 31 | 32.55
29 | 24.1 | 28.3 | 30.875 | 32.4
29.1 | 24 | 28.1667 | 30.7 | 32.25
29.2 | 23.8 | 28 | 30.6 | 32.1
29.3 | 23.7 | 28 | 30.4 | 31.975
29.4 | 23.6 | 27.8 | 30.3 | 31.8
29.5 | 23.475 | 27.7 | 30.15 | 31.6
29.6 | 23.3 | 27.5 | 30 | 31.45
29.7 | 23.2 | 27.45 | 29.9 | 31.3
29.8 | 23.1 | 27.3 | 29.725 | 31.175
29.9 | 22.9667 | 27.1429 | 29.5833 | 31.025
30 | 22.8 | 26.975 | 29.475 | 30.88

## Wedge 60 Lateral Profiles
Lateral Profiles
Wedge 60
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.5 cm FieldSize:4x4 cm2 | Depth:1.5 cm FieldSize:10x10 cm2 | Depth:1.5 cm FieldSize:15x15 cm2 | Depth:1.5 cm FieldSize:40x15 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:15x15 cm2 | Depth:5 cm FieldSize:40x15 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:15x15 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:15x15 cm2 | Depth:20 cm FieldSize:40x15 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:15x15 cm2 | Depth:30 cm FieldSize:40x15 cm2
-14.7 | [18 empty] | 15.4333 | 21.275
-14.6 | [18 empty] | 15.75 | 21.575
-14.5 | [18 empty] | 16.0333 | 21.75
-14.4 | [18 empty] | 16.2333 | 22
-14.3 | [18 empty] | 16.4 | 22.3
-14.2 | [18 empty] | 16.62 | 22.55
-14.1 | [18 empty] | 16.8667 | 22.9
-14 | [18 empty] | 17.1 | 23.1667
-13.9 | [14 empty] | 12.4 | 17.5192 |  |  | 17.4667 | 23.4
-13.8 | [14 empty] | 12.5667 | 17.7191 |  |  | 17.8 | 23.7
-13.7 | [14 empty] | 12.75 | 17.894 |  |  | 18.08 | 24
-13.6 | [14 empty] | 12.975 | 18.2139 |  |  | 18.3 | 24.3667
-13.5 | [14 empty] | 13.125 | 18.3939 |  |  | 18.5667 | 24.7
-13.4 | [14 empty] | 13.375 | 18.5938 |  |  | 19 | 24.8667
-13.3 | [14 empty] | 13.6 | 18.8937 |  |  | 19.1 | 25.1
-13.2 | [10 empty] | 8.6 | 12.9 |  |  | 13.9 | 19.1436 |  |  | 19.4 | 25.5
-13.1 | [10 empty] | 8.7 | 13.15 |  |  | 14.15 | 19.2936 |  |  | 19.5333 | 25.8
-13 | [10 empty] | 8.925 | 13.325 |  |  | 14.35 | 19.4935 |  |  | 19.75 | 26.05
-12.9 | [10 empty] | 9.1 | 13.475 |  |  | 14.6 | 19.7934 |  |  | 20.075 | 26.46
-12.8 | [6 empty] | 6.9 | 11 |  |  | 9.225 | 13.6 |  |  | 14.725 | 20.0733 |  |  | 20.54 | 26.8333
-12.7 | [6 empty] | 7.025 | 11.1 |  |  | 9.3333 | 13.8 |  |  | 15.1 | 20.4182 |  |  | 20.95 | 27.175
-12.6 |  |  | 6.4 |  |  |  | 7.15 | 11.2667 |  |  | 9.55 | 14 |  |  | 15.24 | 20.6681 |  |  | 21.4 | 27.4667
-12.5 |  |  | 6.425 | 10.425 |  |  | 7.25 | 11.4 |  |  | 9.75 | 14.15 |  |  | 15.55 | 20.918 |  |  | 21.6 | 27.7333
-12.4 |  |  | 6.5667 | 10.6 |  |  | 7.4 | 11.5 |  |  | 9.95 | 14.3 |  |  | 15.825 | 21.1429 |  |  | 21.9667 | 28.4
-12.3 |  |  | 6.6 | 10.7 |  |  | 7.5 | 11.7 |  |  | 10.1 | 14.6 |  |  | 16.15 | 21.4928 |  |  | 22.35 | 28.675
-12.2 |  |  | 6.72 | 10.8 |  |  | 7.6 | 11.8 |  |  | 10.3 | 14.8 |  |  | 16.5 | 21.9127 |  |  | 22.7 | 29.0333
-12.1 |  |  | 6.82 | 10.9 |  |  | 7.8 | 11.95 |  |  | 10.55 | 14.95 |  |  | 16.775 | 22.1426 |  |  | 23.1333 | 29.4
-12 |  |  | 6.95 | 11.1 |  |  | 7.9 | 12.1 |  |  | 10.7333 | 15.3 |  |  | 17.05 | 22.4925 |  |  | 23.5 | 29.725
-11.9 |  |  | 7.025 | 11.2 |  |  | 8.05 | 12.2333 |  |  | 10.9 | 15.4667 |  |  | 17.375 | 22.8924 |  |  | 23.75 | 30.1
-11.8 |  |  | 7.1 | 11.225 |  |  | 8.2333 | 12.38 |  |  | 11.1 | 15.725 |  |  | 17.8 | 23.1923 |  |  | 24.2333 | 30.6
-11.7 |  |  | 7.3 | 11.375 |  |  | 8.3 | 12.625 |  |  | 11.35 | 15.925 |  |  | 18.05 | 23.5588 |  |  | 24.7667 | 30.975
-11.6 |  |  | 7.4 | 11.5333 |  |  | 8.4 | 12.8 |  |  | 11.6 | 16.175 |  |  | 18.4333 | 23.992 |  |  | 25.2333 | 31.3667
-11.5 |  |  | 7.5 | 11.6667 |  |  | 8.625 | 12.975 |  |  | 11.825 | 16.525 |  |  | 18.9 | 24.3919 |  |  | 25.85 | 31.8
-11.4 |  |  | 7.58 | 11.82 |  |  | 8.8 | 13.1 |  |  | 12 | 16.7667 |  |  | 19.25 | 24.6718 |  | 10.4 | 26.2667 | 32.3
-11.3 |  |  | 7.7 | 12.05 |  |  | 9 | 13.3 |  |  | 12.375 | 17 |  |  | 19.75 | 25.0416 |  | 10.5 | 26.775 | 32.8
-11.2 |  |  | 7.825 | 12.2 |  |  | 9.2 | 13.4667 |  |  | 12.6 | 17.2667 |  |  | 20.0333 | 25.4915 |  | 10.7 | 27.1667 | 33.4
-11.1 |  |  | 7.975 | 12.3 |  |  | 9.35 | 13.7 |  |  | 12.84 | 17.5 |  |  | 20.5 | 25.8664 |  | 10.9 | 27.65 | 34
-11 |  |  | 8.1 | 12.45 |  |  | 9.5 | 13.95 |  |  | 13.15 | 17.7 |  |  | 20.85 | 26.3246 |  | 11.175 | 28.3 | 34.5
-10.9 |  |  | 8.3 | 12.7 |  |  | 9.7 | 14.2 |  |  | 13.5 | 18.1 |  | 7.9 | 21.36 | 26.6911 |  | 11.3 | 29 | 35.225
-10.8 |  |  | 8.4 | 12.8 |  |  | 9.9 | 14.4 |  |  | 13.68 | 18.5 |  | 7.9333 | 21.75 | 27.141 |  | 11.425 | 29.7667 | 35.8
-10.7 |  |  | 8.575 | 12.95 |  |  | 10.125 | 14.6 |  |  | 14.1333 | 18.8 |  | 8.1 | 22.175 | 27.5908 |  | 11.6 | 30.7 | 36.55
-10.6 |  |  | 8.7 | 13.15 |  |  | 10.3 | 14.7667 |  |  | 14.4667 | 19.08 |  | 8.3 | 22.675 | 28.1906 |  | 11.8 | 31.7 | 37.6
-10.5 |  |  | 8.8667 | 13.3333 |  |  | 10.55 | 15 |  |  | 14.7333 | 19.5 |  | 8.45 | 23.35 | 28.5905 |  | 11.9 | 33.2333 | 38.8
-10.4 |  |  | 9.1 | 13.5 |  |  | 10.8 | 15.275 |  | 5.3 | 15.125 | 19.8333 |  | 8.6 | 23.775 | 29.1653 |  | 12.2 | 35 | 40.5
-10.3 |  |  | 9.2333 | 13.7 |  |  | 11.05 | 15.5 |  | 5.4 | 15.45 | 20.15 |  | 8.75 | 24.35 | 29.7901 |  | 12.4 | 37.7333 | 42.85
-10.2 |  |  | 9.4 | 13.875 |  | 4.2 | 11.3 | 15.8 |  | 5.5 | 15.8 | 20.65 |  | 8.9 | 24.98 | 30.2566 |  | 12.5667 | 41.5 | 46.3
-10.1 |  |  | 9.5 | 14.025 |  | 4.2 | 11.6 | 16.05 |  | 5.6 | 16.3 | 21.0333 |  | 9.1 | 25.7 | 31.0896 |  | 12.85 | 47.6 | 51.8
-10 |  | 3.9 | 9.7 | 14.2333 |  | 4.3 | 11.8333 | 16.3 |  | 5.7 | 16.7 | 21.35 |  | 9.2 | 26.55 | 31.8227 |  | 13.1 | 55.975 | 59.1667
-9.9 |  | 3.9 | 10 | 14.45 |  | 4.3 | 12.175 | 16.6667 |  | 5.8 | 17.1 | 21.925 |  | 9.3 | 27.6 | 32.7391 |  | 13.25 | 67 | 69
-9.8 |  | 4 | 10.14 | 14.6667 |  | 4.4 | 12.45 | 16.9667 |  | 5.9 | 17.6 | 22.3333 |  | 9.45 | 28.75 | 34.0886 |  | 13.5 | 79.925 | 81.1
-9.7 |  | 4 | 10.3667 | 14.9 |  | 4.5 | 12.85 | 17.3 |  | 6 | 18.15 | 22.8333 |  | 9.6667 | 30.4 | 35.3383 |  | 13.75 | 93.82 | 92.575
-9.6 |  | 4.1 | 10.525 | 15.1 |  | 4.6 | 13.125 | 17.6 |  | 6.1 | 18.7 | 23.2333 |  | 9.9 | 32.58 | 37.4375 |  | 13.85 | 106.3 | 104.9
-9.5 |  | 4.1 | 10.8 | 15.325 |  | 4.6 | 13.55 | 18 |  | 6.2 | 19.3 | 23.8667 |  | 10.05 | 35.725 | 40.6364 |  | 14.075 | 118.25 | 115.04
-9.4 |  | 4.1 | 11.05 | 15.6 |  | 4.7 | 13.95 | 18.45 |  | 6.275 | 19.82 | 24.6 |  | 10.2 | 40.975 | 45.5848 |  | 14.4333 | 127.5333 | 124
-9.3 |  | 4.2 | 11.3 | 15.85 |  | 4.8 | 14.425 | 18.8333 |  | 6.4 | 20.675 | 25.25 |  | 10.35 | 49.175 | 53.0823 |  | 14.6667 | 134.9 | 129.9
-9.2 |  | 4.2 | 11.6 | 16.175 |  | 4.9 | 14.85 | 19.3 |  | 6.5 | 21.525 | 26.025 |  | 10.625 | 61.25 | 63.1989 |  | 14.95 | 138.7 | 133.7
-9.1 |  | 4.3 | 11.875 | 16.4333 |  | 5 | 15.32 | 19.8 |  | 6.6 | 22.575 | 27.125 |  | 10.875 | 75.225 | 75.9746 |  | 15.325 | 141.05 | 135.85
-9 |  | 4.3333 | 12.1667 | 16.8 |  | 5.1 | 15.8 | 20.3333 |  | 6.8 | 24.02 | 28.475 |  | 11.15 | 90.5 | 90.1033 |  | 15.525 | 142.3 | 137.1
-8.9 |  | 4.4 | 12.45 | 17.1 |  | 5.1 | 16.55 | 20.925 |  | 6.975 | 25.975 | 30.3667 |  | 11.3667 | 105 | 104.065 |  | 15.8 | 143.1 | 137.75
-8.8 |  | 4.525 | 12.8667 | 17.5 |  | 5.2 | 17.4 | 21.9 |  | 7.1 | 28.8 | 33.175 |  | 11.6333 | 119.3 | 117.061 |  | 16.16 | 143.4 | 138.1333
-8.7 |  | 4.625 | 13.4333 | 17.9333 |  | 5.3 | 18.55 | 23 |  | 7.3 | 33.3 | 37.4 |  | 11.8 | 131.8 | 128.2075 |  | 16.52 | 143.5333 | 138.2
-8.6 |  | 4.7 | 13.8667 | 18.5 |  | 5.4 | 20.1 | 24.35 |  | 7.525 | 41.1 | 44.2 |  | 12.075 | 140.3 | 136.155 |  | 16.8 | 143.6 | 138.2
-8.5 |  | 4.7 | 14.6667 | 19.2 |  | 5.55 | 22.35 | 26.475 |  | 7.7 | 52.8 | 55.5 |  | 12.4 | 146.1 | 140.753 |  | 17.2 | 143.525 | 138.2
-8.4 |  | 4.8 | 15.7 | 20.1 |  | 5.6 | 25.58 | 29.8 |  | 7.8 | 68.94 | 70.5 |  | 12.5 | 148.92 | 143.352 |  | 17.5 | 143.3 | 138
-8.3 |  | 4.9 | 17.2333 | 21.6 |  | 5.8 | 30.85 | 35.3 |  | 7.92 | 86.56 | 87.5 |  | 12.9 | 150.2 | 145.052 |  | 18 | 142.925 | 137.8
-8.2 |  | 5 | 19.6 | 24 |  | 5.9 | 40.6 | 44.1333 |  | 8.15 | 105.05 | 104.9 |  | 13.4 | 150.875 | 145.626 |  | 18.3 | 142.7 | 137.6
-8.1 |  | 5.025 | 23.6667 | 27.6 |  | 6.05 | 56.35 | 57.95 |  | 8.4 | 122.7 | 121.6 |  | 13.5667 | 151.1 | 145.851 |  | 18.7 | 142.5333 | 137.4
-8 |  | 5.1 | 30.7333 | 34.25 |  | 6.2 | 75.675 | 75.6 |  | 8.6 | 138.5 | 135.54 |  | 13.95 | 151.1 | 145.851 |  | 18.9667 | 142.3 | 137.3
-7.9 |  | 5.2 | 44.3 | 46.125 |  | 6.3 | 94.9667 | 95.2 |  | 8.8 | 150.4 | 145.825 |  | 14.3 | 151.1 | 145.8177 |  | 19.45 | 141.7333 | 136.85
-7.8 |  | 5.3 | 61.525 | 63.65 |  | 6.475 | 115.3 | 114.6 |  | 9.0667 | 157.75 | 152.825 |  | 14.5333 | 150.75 | 145.6177 |  | 20 | 141.3 | 136.45
-7.7 |  | 5.425 | 84.675 | 83.9 |  | 6.7 | 134.7 | 132.2 |  | 9.35 | 161 | 156.45 |  | 14.95 | 150.4 | 145.1853 |  | 20.5667 | 141 | 136.1667
-7.6 |  | 5.575 | 106.0333 | 105.45 |  | 6.8 | 150.4 | 147.5 |  | 9.575 | 162.5 | 157.75 |  | 15.2667 | 150.0667 | 145.052 |  | 21.05 | 140.6333 | 135.8333
-7.5 |  | 5.7 | 128.6667 | 127.325 |  | 7 | 162.0333 | 157.65 |  | 9.86 | 163.0333 | 158.2667 |  | 15.7333 | 149.7 | 144.6853 | 4.2 | 21.9 | 140.175 | 135.4667
-7.4 |  | 5.8 | 148.7667 | 145.725 |  | 7.175 | 168.4 | 163.8 |  | 10.3 | 163.1 | 158.3 | 2.901 | 16.25 | 149.2 | 144.152 | 4.3 | 22.6 | 139.8 | 135.1
-7.3 |  | 6 | 164.2 | 159.2 |  | 7.35 | 170.8 | 166.2667 |  | 10.55 | 163.05 | 158.3 | 3.001 | 16.7 | 148.45 | 143.552 | 4.3 | 23.4 | 139.175 | 134.6
-7.2 |  | 6.1 | 173.8667 | 168.6 |  | 7.6 | 171.3 | 167.125 |  | 10.8 | 162.7 | 158.1 | 3.101 | 17.25 | 147.8 | 143.052 | 4.4 | 24.6333 | 138.65 | 134.1667
-7.1 |  | 6.22 | 177.6 | 172.45 | 1.5008 | 7.85 | 171.4 | 167.2 | 1.9 | 11.1 | 162 | 157.55 | 3.2011 | 17.7 | 147.2 | 142.702 | 4.5 | 26.8 | 138.3 | 133.7667
-7 | 1.4 | 6.35 | 178.7 | 173.3 | 1.5008 | 8.0333 | 171.375 | 167.2 | 1.9 | 11.5333 | 161.44 | 156.7 | 3.2011 | 18.35 | 146.75 | 142.203 | 4.6 | 29.4667 | 137.9 | 133.375
-6.9 | 1.3 | 6.54 | 178.7 | 173.3 | 1.5008 | 8.4 | 170.9 | 166.75 | 1.9 | 11.9333 | 160.45 | 156.02 | 3.2011 | 19.2 | 146.15 | 141.653 | 4.75 | 33 | 137.5 | 133.125
-6.8 | 1.3 | 6.78 | 178.54 | 173.22 | 1.5008 | 8.65 | 170.2 | 165.8 | 2 | 12.3667 | 159.6 | 155.3 | 3.2011 | 20.25 | 145.5 | 141.053 | 4.9 | 38.8 | 136.9 | 132.55
-6.7 | 1.3 | 6.925 | 177.7 | 172.5 | 1.5008 | 8.9333 | 169.1 | 165 | 2 | 12.7 | 159 | 154.45 | 3.2011 | 21.7 | 144.8 | 140.503 | 4.9 | 47.5667 | 136.22 | 132.15
-6.6 | 1.3 | 7.1 | 176.25 | 171.2 | 1.5008 | 9.2 | 168 | 164 | 2 | 13.35 | 158 | 153.725 | 3.3011 | 23.7 | 144.0333 | 140.028 | 4.9 | 58.4 | 135.75 | 131.7
-6.5 | 1.3 | 7.3 | 175 | 170.0667 | 1.5008 | 9.6 | 166.95 | 162.85 | 2.1 | 13.95 | 157.3 | 153.05 | 3.4011 | 26.3 | 143.4 | 139.4867 | 4.9 | 69.7 | 135.0667 | 131.3
-6.4 | 1.3 | 7.6 | 173.66 | 168.75 | 1.5008 | 9.9667 | 165.6 | 161.9 | 2.1 | 14.775 | 156.3 | 151.825 | 3.5012 | 30.75 | 142.8 | 138.8207 | 5 | 81.4 | 134.65 | 130.9
-6.3 | 1.4 | 7.86 | 172.0667 | 167.425 | 1.6008 | 10.5 | 164.55 | 160.7 | 2.1 | 15.8 | 154.8 | 151.1 | 3.6012 | 37.55 | 142 | 138.2873 | 5 | 93.8333 | 134.15 | 130.25
-6.2 | 1.4 | 8.1333 | 170.56 | 166.3333 | 1.6008 | 11.15 | 163.4 | 159.3 | 2.1 | 17.0667 | 154.15 | 150.3 | 3.7012 | 47.3 | 141.55 | 137.454 | 5.2 | 103.6 | 133.725 | 129.65
-6.1 | 1.4 | 8.5 | 169.3 | 164.5667 | 1.6008 | 12 | 162.1 | 158.25 | 2.2 | 19.025 | 153.2 | 149.4 | 3.8013 | 59.425 | 140.7 | 136.979 | 5.2 | 111.6 | 132.925 | 129.2
-6 | 1.4 | 9.04 | 167.55 | 163.15 | 1.6008 | 13.15 | 161.025 | 157.225 | 2.2 | 21.6 | 152.2 | 148.475 | 3.9013 | 72.825 | 139.95 | 136.155 | 5.3 | 116.4667 | 132.32 | 128.7333
-5.9 | 1.4333 | 9.7 | 166.1 | 161.7333 | 1.7008 | 14.8 | 159.7 | 156.1 | 2.3 | 26.3333 | 151.1333 | 147.45 | 4.0013 | 85.925 | 139.34 | 135.455 | 5.3 | 119.3 | 131.775 | 128.275
-5.8 | 1.5 | 10.7 | 164.8 | 160.225 | 1.7008 | 16.9667 | 158.3 | 154.825 | 2.4 | 34.1 | 150.3 | 146.6 | 4.0013 | 98.35 | 138.625 | 135.055 | 5.5 | 121.3 | 131.1 | 127.85
-5.7 | 1.5 | 12.3 | 163 | 158.95 | 1.7008 | 21.0667 | 157.15 | 153.3 | 2.4 | 45.2 | 149.4333 | 145.7333 | 4.2014 | 109 | 137.875 | 134.235 | 5.7 | 122.1 | 130.525 | 127.3
-5.6 | 1.5 | 14.7 | 161.675 | 157.325 | 1.7008 | 27.3667 | 155.525 | 152.4 | 2.5 | 59.1 | 148.3 | 144.6 | 4.2347 | 116.7 | 137 | 133.4557 | 5.8 | 122.4667 | 129.85 | 126.75
-5.5 | 1.5 | 18.9333 | 160.05 | 156.2 | 1.8009 | 37.8 | 154.75 | 151.1 | 2.5 | 73.95 | 147 | 143.8 | 4.3014 | 120.75 | 136.325 | 132.856 | 5.95 | 122.6 | 129.4 | 126.15
-5.4 | 1.6 | 26.7 | 158.575 | 154.8 | 1.8009 | 52.0667 | 153.45 | 149.975 | 2.6 | 89.3 | 145.8333 | 142.6 | 4.4015 | 123.6 | 135.6 | 132.106 | 6.15 | 122.55 | 128.725 | 125.65
-5.3 | 1.6 | 38.575 | 156.925 | 153.4 | 1.8009 | 68.55 | 151.9 | 148.8 | 2.65 | 102.64 | 145.05 | 141.9 | 4.5015 | 125.02 | 134.7333 | 131.456 | 6.3 | 122.4 | 128 | 125.325
-5.2 | 1.6 | 54.15 | 155.8 | 151.95 | 1.9009 | 85.325 | 150.3 | 147.55 | 2.75 | 113.875 | 143.8 | 140.7 | 4.5015 | 125.55 | 133.9 | 130.756 | 6.4 | 122.2 | 127.425 | 124.75
-5.1 | 1.6 | 71.95 | 154.25 | 150.575 | 2.001 | 100.9 | 149.1 | 146.3667 | 2.8 | 122.425 | 143.0333 | 139.8 | 4.5015 | 125.7 | 133.3 | 130.157 | 6.35 | 122.075 | 126.8 | 124.0667
-5 | 1.6 | 90 | 152.6 | 149.325 | 2.001 | 113.7667 | 148.04 | 145.2 | 2.9 | 127.325 | 142.025 | 138.95 | 4.7016 | 125.7 | 132.5 | 129.307 | 6.5 | 121.8 | 126.3667 | 123.5
-4.9 | 1.7 | 106.1667 | 151.2 | 147.9333 | 2.101 | 124.65 | 146.8 | 143.9 | 3 | 129.6 | 140.85 | 138.075 | 4.8016 | 125.675 | 131.675 | 128.657 | 6.6 | 121.5 | 125.7 | 123
-4.8 | 1.75 | 121.425 | 149.7 | 146.45 | 2.101 | 130.55 | 145.4 | 142.8 | 3.1 | 130.425 | 139.6 | 137 | 4.9016 | 125.325 | 130.8 | 127.657 | 6.7 | 121.2 | 125 | 122.6
-4.7 | 1.8 | 130.8 | 148.3 | 145.2 | 2.2011 | 133.7333 | 144.3 | 141.7 | 3.2 | 130.625 | 138.7 | 135.975 | 5.1017 | 125.0333 | 130 | 127.0913 | 6.9 | 120.8 | 124.375 | 122
-4.6 | 1.8 | 136.08 | 146.8 | 143.975 | 2.3011 | 134.7333 | 142.9 | 140.5 | 3.3 | 130.7 | 137.4 | 134.84 | 5.3018 | 124.5667 | 129.3 | 126.4247 | 7.05 | 120.2 | 123.9 | 121.42
-4.5 | 1.8 | 138.3 | 145.4333 | 142.5 | 2.3011 | 134.9 | 141.75 | 139.225 | 3.4 | 130.5667 | 136.4 | 133.9 | 5.4351 | 124 | 128.5 | 125.858 | 7.3 | 119.9 | 123.4 | 120.775
-4.4 | 1.9 | 138.5 | 144.2 | 141.0333 | 2.4012 | 134.85 | 140.25 | 137.875 | 3.6 | 130.1667 | 135.4 | 132.75 | 5.6019 | 123.6 | 127.58 | 125.258 | 7.5 | 119.65 | 122.75 | 120.475
-4.3 | 1.9 | 138.5 | 142.7 | 139.8 | 2.5013 | 134.625 | 139.225 | 136.68 | 3.7 | 129.7 | 134.35 | 132.1 | 5.7019 | 123 | 126.96 | 124.5585 | 7.7 | 119.2 | 122.3 | 120.05
-4.2 | 1.9 | 138.2 | 141.3 | 138.55 | 2.6013 | 133.9 | 137.72 | 135.5333 | 3.9 | 129.1 | 133.1 | 130.95 | 5.9353 | 122.65 | 126.325 | 123.759 | 7.9 | 118.7667 | 121.65 | 119.675
-4.1 | 2 | 137.3333 | 140.075 | 137.2 | 2.7014 | 133.28 | 136.55 | 134.35 | 4 | 128.2 | 132.1 | 130.025 | 6.102 | 122.2 | 125.575 | 123.309 | 7.9 | 118.5 | 121.3 | 119.175
-4 | 2.1 | 136.6 | 138.475 | 135.9 | 2.8014 | 132.5 | 135.5667 | 133.35 | 4.1 | 127.6 | 131.3 | 129.175 | 6.4021 | 121.8 | 124.8 | 122.584 | 8.2 | 118.175 | 120.75 | 118.625
-3.9 | 2.1 | 135.4 | 137.1 | 134.65 | 2.9015 | 131.5 | 134.1 | 132.15 | 4.3 | 126.88 | 130.2 | 128.2 | 6.6022 | 121.1 | 124.1 | 121.659 | 8.4667 | 117.825 | 120.125 | 118.2
-3.8 | 2.2 | 134 | 135.6333 | 133.45 | 3.0015 | 130.3 | 132.8 | 131.05 | 4.45 | 126.225 | 129.2333 | 127.325 | 6.9023 | 120.5333 | 123.4 | 121.185 | 8.8 | 117.45 | 119.575 | 117.75
-3.7 | 2.3 | 133.26 | 134.5 | 132.2 | 3.2016 | 129.5 | 131.6333 | 129.9333 | 4.7 | 125.575 | 128.3 | 126.475 | 7.1024 | 120.025 | 122.7667 | 120.66 | 9.15 | 117.125 | 119.05 | 117.16
-3.6 | 2.4 | 131.9 | 133 | 130.975 | 3.3016 | 128.5 | 130.56 | 128.9 | 4.85 | 124.75 | 127.4 | 125.55 | 7.4525 | 119.65 | 122.1667 | 120.06 | 9.65 | 116.775 | 118.475 | 116.7
-3.5 | 2.5 | 130.8333 | 131.75 | 129.9 | 3.4017 | 127.9 | 129.4 | 128.1 | 5.1 | 124.2 | 126.48 | 124.8 | 7.8526 | 119.2667 | 121.6 | 119.5933 | 10.2 | 116.4 | 118.075 | 116.4
-3.4 | 2.5 | 129.775 | 130.5 | 128.6 | 3.6018 | 127.1 | 128.45 | 127.1 | 5.3 | 123.4 | 125.8 | 123.925 | 8.3028 | 118.85 | 120.9667 | 118.66 | 11.25 | 116.1 | 117.7 | 115.95
-3.3 | 2.6 | 128.6 | 129.6667 | 127.7 | 3.9019 | 126.225 | 127.48 | 126.04 | 5.65 | 122.7333 | 124.825 | 123.25 | 8.8029 | 118.325 | 120.3 | 118.361 | 12.05 | 115.55 | 117.25 | 115.48
-3.2 | 2.8 | 127.8333 | 128.3667 | 126.5 | 4.1021 | 125.475 | 126.6 | 124.9 | 6 | 122.35 | 124.075 | 122.5 | 9.4532 | 117.675 | 119.7 | 117.761 | 13.5 | 115.15 | 116.8 | 115.1333
-3.1 | 3 | 126.9 | 127.45 | 125.7 | 4.5523 | 124.55 | 125.65 | 124.075 | 6.5 | 121.8 | 123.4 | 121.775 | 10.5535 | 117.44 | 119.1 | 117.286 | 15.5667 | 114.8667 | 116.3 | 114.75
-3 | 3.3 | 125.9 | 126.5 | 124.75 | 4.9024 | 123.86 | 124.75 | 123.325 | 7 | 120.8 | 122.6 | 120.9667 | 11.954 | 117.125 | 118.7 | 116.761 | 18.5 | 114.4 | 115.7 | 114.2333
-2.9 | 3.65 | 125.1 | 125.4 | 123.975 | 5.6528 | 123.2 | 123.75 | 122.425 | 8.2 | 120.2667 | 121.7 | 120.3 | 13.9046 | 116.5667 | 117.9 | 116.111 | 23.8 | 114 | 115.2667 | 113.65
-2.8 | 4.15 | 124.2 | 124.5667 | 123.0667 | 6.5533 | 122.3 | 122.96 | 121.8 | 9.6 | 119.675 | 120.95 | 119.4 | 17.7059 | 116.0333 | 117.45 | 115.761 | 30.9 | 113.6333 | 114.56 | 113.175
-2.7 | 5 | 123.2 | 123.7333 | 122.1333 | 8.004 | 121.6 | 122 | 120.95 | 11.6 | 119 | 120.1 | 118.8 | 22.6075 | 115.55 | 116.85 | 115.1953 | 40.4 | 113.25 | 113.9 | 112.85
-2.6 | 6.2333 | 122.3667 | 122.55 | 121.24 | 10.1051 | 120.6 | 121.3 | 120 | 15.2 | 118.4667 | 119.3 | 118.08 | 30.8103 | 115.05 | 116.2 | 114.462 | 52 | 112.8667 | 113.35 | 112.25
-2.5 | 8.7 | 121.4 | 121.8 | 120.25 | 13.6569 | 120.025 | 120.4 | 119.2667 | 20.3 | 117.85 | 118.625 | 117.375 | 41.4138 | 114.5 | 115.425 | 113.912 | 63.8 | 112.4333 | 113 | 111.875
-2.4 | 12.6 | 120.55 | 120.66 | 119.575 | 19.8599 | 118.9 | 119.64 | 118.5 | 28.95 | 117.1 | 117.8667 | 116.725 | 53.6179 | 114 | 114.78 | 113.412 | 74.3 | 112.1 | 112.5 | 111.55
-2.3 | 19.3 | 119.7333 | 119.825 | 118.5 | 29.4147 | 118.175 | 118.675 | 117.825 | 40.8 | 116.65 | 117.0333 | 115.9 | 65.922 | 113.4333 | 114.3 | 112.962 | 85.2 | 111.8 | 112.15 | 111
-2.2 | 30 | 118.7 | 119.3 | 117.575 | 41.4207 | 117.525 | 117.725 | 116.975 | 53.05 | 115.9 | 116.4 | 115.275 | 77.7259 | 112.9333 | 113.55 | 112.288 | 93.8 | 111.05 | 111.575 | 110.6
-2.1 | 43.2 | 117.9 | 118 | 116.85 | 55.1943 | 116.7 | 117.05 | 116.1 | 66.2 | 115.1 | 115.6667 | 114.6 | 88.3294 | 112.25 | 112.7667 | 111.763 | 99.15 | 110.55 | 111.25 | 110.1
-2 | 58.2 | 116.85 | 117.2 | 116 | 67.834 | 115.9667 | 116.2 | 115.2 | 78.75 | 114.325 | 114.7333 | 113.775 | 95.5819 | 111.675 | 112.3 | 111.313 | 102.6 | 109.95 | 110.7 | 109.55
-1.9 | 72 | 116.05 | 116.2 | 115.2 | 80.8404 | 115.02 | 115.32 | 114.5 | 89.7 | 113.575 | 113.9 | 113.225 | 101.034 | 111.125 | 111.7333 | 110.563 | 104.6 | 109.55 | 110.1667 | 109.05
-1.8 | 85 | 115.275 | 115.4 | 114.3 | 92.046 | 114.275 | 114.24 | 113.5667 | 97.1 | 112.9 | 113.2333 | 112.3 | 104.135 | 110.48 | 111.0333 | 110.0963 | 105.8333 | 109.1667 | 109.45 | 108.75
-1.7 | 95.65 | 114.42 | 114.3 | 113.4 | 99.7498 | 113.5667 | 113.4 | 112.7 | 102.65 | 111.9667 | 112.4 | 111.8333 | 105.835 | 110 | 110.55 | 109.563 | 106.25 | 108.6667 | 109 | 108.25
-1.6 | 103 | 113.5 | 113.45 | 112.6333 | 104.302 | 112.6 | 112.5 | 111.9 | 105.2 | 111.4 | 111.7 | 110.95 | 106.335 | 109.4 | 110.0667 | 108.964 | 106.5 | 108.1333 | 108.7 | 107.9
-1.5 | 106.7 | 112.5 | 112.6 | 111.7667 | 106.453 | 111.825 | 111.9 | 111.25 | 106.45 | 110.7 | 110.9333 | 110.2333 | 106.636 | 108.8 | 109.35 | 108.464 | 106.7 | 107.6 | 107.95 | 107.225
-1.4 | 108.1 | 111.7 | 111.9 | 111 | 107.354 | 111.1 | 111.1 | 110.4333 | 106.8333 | 110.1333 | 110.2 | 109.5667 | 106.736 | 108.2 | 108.6 | 108.064 | 106.4333 | 107.15 | 107.5 | 106.7667
-1.3 | 108.2 | 110.9 | 110.8 | 110.3667 | 107.554 | 110.2 | 110.0333 | 109.7 | 107 | 109.2333 | 109.4 | 108.8 | 106.736 | 107.5 | 108 | 107.364 | 106 | 106.45 | 106.9667 | 106.325
-1.2 | 108.2 | 109.9 | 110.05 | 109.3 | 107.554 | 109.46 | 109.5 | 108.8 | 107 | 108.625 | 108.6 | 108.3 | 106.536 | 106.95 | 107.4333 | 106.564 | 105.7 | 105.9 | 106.3667 | 105.8
-1.1 | 107.95 | 109.2 | 109.15 | 108.5 | 107.554 | 108.35 | 108.4 | 108.1667 | 106.8 | 107.8333 | 107.8 | 107.45 | 106.035 | 106.625 | 106.8 | 106.19 | 105.4 | 105.5 | 105.75 | 105.3
-1 | 107.6 | 108.2 | 108.2 | 107.9 | 107.154 | 107.825 | 107.7333 | 107.4 | 106.4 | 107.1 | 106.98 | 106.68 | 105.635 | 105.9 | 106.1667 | 105.615 | 105.1 | 104.9 | 105.2333 | 104.8
-0.9 | 106.8667 | 107.15 | 107.425 | 106.8 | 106.453 | 107 | 106.95 | 106.55 | 105.8 | 106.5 | 106.55 | 105.95 | 105.135 | 105.3 | 105.5 | 104.94 | 104.75 | 104.5 | 104.4 | 104.35
-0.8 | 106.3 | 106.4 | 106.6 | 105.95 | 105.953 | 105.925 | 106.1 | 105.8 | 105.2 | 105.8 | 105.725 | 105.425 | 104.435 | 104.8 | 104.875 | 104.24 | 104.2 | 104.1 | 104.1 | 104.1
-0.7 | 105.9 | 105.65 | 105.725 | 105.3 | 105.303 | 105.275 | 105.3 | 105.15 | 104.6 | 104.9667 | 104.95 | 104.875 | 104.035 | 104.08 | 104.2667 | 103.715 | 103.9 | 103.7 | 103.725 | 103.6
-0.6 | 105.1 | 104.9 | 104.85 | 104.5 | 104.452 | 104.475 | 104.5 | 104.4 | 104.05 | 104.2 | 104.16 | 104.1333 | 103.585 | 103.575 | 103.64 | 103.216 | 103.4 | 103 | 103.15 | 103.2
-0.5 | 104.1 | 104 | 104.3 | 103.5 | 103.852 | 103.5 | 103.875 | 103.6 | 103.5 | 103.55 | 103.5 | 103.4 | 102.984 | 102.75 | 102.9 | 102.666 | 102.6 | 102.35 | 102.625 | 102.6
-0.4 | 103.1 | 103.0667 | 103.3 | 102.95 | 102.951 | 102.9667 | 102.96 | 102.9 | 102.75 | 102.95 | 102.9 | 102.6 | 102.284 | 102.24 | 102.5667 | 101.966 | 102.2 | 101.9 | 102.16 | 101.95
-0.3 | 102.3 | 102.2667 | 102.5 | 102.3 | 102.351 | 102 | 102.2 | 102.1 | 101.95 | 102.14 | 102.2 | 101.9 | 101.784 | 101.75 | 102.15 | 101.566 | 101.9 | 101.35 | 101.6 | 101.52
-0.2 | 101.65 | 101.4 | 101.8 | 101.525 | 101.451 | 101.4 | 101.4 | 101.5 | 101.45 | 101.5 | 101.1 | 101.4 | 101.3007 | 101.4 | 101.2667 | 101.116 | 101.15 | 100.9 | 101 | 101.05
-0.1 | 100.7 | 100.6 | 100.95 | 100.775 | 100.8 | 100.85 | 100.5 | 100.7 | 100.8 | 100.7 | 100.7 | 100.5 | 100.584 | 100.65 | 100.55 | 100.566 | 100.6 | 100.4667 | 100.6 | 100.5
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9998 | 100 | 100 | 100.0001 | 100 | 100 | 100 | 100
0.1 | 99.2 | 99.3 | 99.2333 | 99.425 | 99.4164 | 99.1 | 99.2 | 99.2 | 99.4 | 99.5 | 99.3 | 99.3 | 99.4331 | 99.45 | 99.4667 | 99.5168 | 99.6333 | 99.7 | 99.45 | 99.5
0.2 | 98.45 | 98.2667 | 98.575 | 98.7333 | 99.1496 | 98.5 | 98.4 | 98.675 | 98.7 | 98.7 | 98.75 | 98.6333 | 98.8329 | 98.76 | 99.1333 | 98.917 | 99 | 99.25 | 99.1 | 99.1
0.3 | 97.65 | 97.6 | 97.9 | 97.86 | 98.1491 | 97.6 | 97.8 | 97.975 | 97.9 | 98.04 | 97.9 | 98.0667 | 98.2327 | 98.32 | 98.5333 | 98.3422 | 98.5 | 98.625 | 98.6 | 98.72
0.4 | 96.9 | 96.76 | 97.075 | 97.25 | 97.3487 | 97.1 | 97.1 | 97.325 | 97.2 | 97.25 | 97.4 | 97.6 | 97.6325 | 97.675 | 97.9 | 98.0006 | 98 | 97.9333 | 98.0667 | 98.35
0.5 | 95.9667 | 96.175 | 96.5 | 96.7 | 96.6483 | 96.3 | 96.4667 | 96.7 | 96.7 | 96.7667 | 96.7333 | 97 | 97.0323 | 97.25 | 97.3667 | 97.3925 | 97.4333 | 97.45 | 97.6 | 97.95
0.6 | 95.3 | 95.4 | 95.7667 | 95.9667 | 95.6478 | 95.6 | 95.7 | 96.0333 | 95.9 | 96.1 | 96.2 | 96.2 | 96.4321 | 96.65 | 96.5667 | 96.8177 | 96.75 | 97 | 97.1 | 97.45
0.7 | 94.6 | 94.7 | 94.9 | 95.3 | 94.6473 | 95.04 | 95.25 | 95.4333 | 95.3 | 95.5 | 95.5 | 95.6333 | 95.8319 | 96.15 | 96.2 | 96.3179 | 96.4 | 96.5 | 96.6667 | 97.1
0.8 | 93.6 | 94.1 | 94.2667 | 94.6333 | 94.2471 | 94.225 | 94.4333 | 94.7667 | 94.4 | 94.9 | 94.8 | 95.1333 | 95.0317 | 95.4 | 95.75 | 95.893 | 95.9 | 96 | 96.1667 | 96.6
0.9 | 92.9 | 93.35 | 93.7 | 94 | 93.3467 | 93.675 | 93.8667 | 94.1 | 93.8 | 94.4 | 94.175 | 94.6 | 94.3815 | 94.85 | 95.25 | 95.3282 | 94.85 | 95.6 | 95.6667 | 96
1 | 92.15 | 92.7 | 93.1 | 93.2 | 92.3462 | 93.2 | 93.1667 | 93.65 | 92.8 | 93.6667 | 93.75 | 93.9667 | 93.6812 | 94.2333 | 94.6 | 94.8017 | 94.15 | 95.1333 | 95.175 | 95.6
1.1 | 91.1 | 91.95 | 92.4 | 92.5333 | 91.6458 | 92.4667 | 92.5 | 93 | 92 | 93.075 | 93.1667 | 93.4667 | 92.9309 | 93.6667 | 94.05 | 94.2353 | 93.8 | 94.75 | 94.65 | 95.2
1.2 | 89.8 | 91.24 | 91.8 | 91.95 | 90.6453 | 91.6667 | 91.975 | 92.4 | 91.05 | 92.425 | 92.6 | 92.9 | 92.0807 | 93.3 | 93.6 | 93.7354 | 93.1 | 94.1 | 94.3 | 94.8
1.3 | 88.8 | 90.5667 | 91.2 | 91.5 | 89.2446 | 91.1 | 91.2333 | 91.75 | 89.95 | 91.7 | 91.9 | 92.25 | 91.3638 | 92.7667 | 92.95 | 93.3689 | 92.3 | 93.7 | 93.775 | 94.44
1.4 | 87.3 | 90.1 | 90.5 | 90.8 | 87.6438 | 90.3 | 90.5333 | 91.2 | 88.85 | 91.15 | 91.3667 | 91.7333 | 90.8303 | 92.05 | 92.3 | 92.7191 | 91.7 | 93.2667 | 93.25 | 93.95
1.5 | 84.85 | 89.3333 | 89.8333 | 90.1667 | 85.5761 | 89.6 | 89.9333 | 90.6 | 87.45 | 90.6 | 90.775 | 91.15 | 89.6299 | 91.5 | 91.7 | 92.2692 | 90.95 | 92.625 | 92.825 | 93.45
1.6 | 81.2 | 88.7 | 89.0667 | 89.55 | 83.0415 | 88.8 | 89.5 | 89.9 | 85.4 | 89.8 | 90.1 | 90.5333 | 88.2294 | 91.1 | 91.325 | 91.6944 | 90.2 | 91.875 | 92.3 | 92.85
1.7 | 75.7 | 87.9 | 88.44 | 88.96 | 79.0395 | 88.3333 | 88.7 | 89.45 | 82.7 | 89.2667 | 89.4333 | 90 | 86.8289 | 90.6667 | 90.7 | 91.1196 | 89.4 | 91.3 | 91.75 | 92.4667
1.8 | 67.4 | 87.3 | 87.9 | 88.4 | 72.1361 | 87.8 | 88.12 | 88.9 | 78.9 | 88.6 | 88.85 | 89.3 | 84.9784 | 90.05 | 90.35 | 90.6698 | 88 | 90.9 | 91.2 | 92.0333
1.9 | 57.8 | 86.7 | 87.1 | 87.7667 | 63.9987 | 86.95 | 87.56 | 87.975 | 72.6 | 87.95 | 88.3 | 88.8667 | 82.0273 | 89.55 | 89.675 | 90.1199 | 86.6 | 90.4 | 90.85 | 91.6
2 | 46.7 | 85.975 | 86.45 | 87.0667 | 54.1271 | 86.5 | 86.925 | 87.35 | 65 | 87.2667 | 87.8 | 88.35 | 78.0593 | 88.8 | 89.15 | 89.6035 | 84.7 | 90 | 90.275 | 91.1
2.1 | 35.9667 | 85.15 | 85.9 | 86.3 | 42.9215 | 85.8 | 86.275 | 86.94 | 54.9 | 86.76 | 87 | 87.8 | 72.4741 | 87.875 | 88.7 | 89.1036 | 82.2 | 89.3 | 89.75 | 90.625
2.2 | 25.4 | 84.55 | 85.3 | 85.75 | 33.1166 | 85.1667 | 85.55 | 86.36 | 45.6 | 86.2 | 86.45 | 87.1 | 64.1714 | 87.4 | 88.2 | 88.5455 | 78.55 | 88.85 | 89.425 | 90.2667
2.3 | 16.7333 | 83.94 | 84.6667 | 85.15 | 23.912 | 84.5 | 85.05 | 85.85 | 35.3 | 85.5 | 85.8333 | 86.5667 | 55.4685 | 87 | 87.4 | 88.0707 | 72.8 | 88.35 | 89.05 | 90
2.4 | 10.6 | 83.2667 | 84.0333 | 84.6667 | 16.4582 | 83.775 | 84.45 | 85.3 | 26.1 | 84.7 | 85.3667 | 86.1 | 45.7486 | 86.4 | 87 | 87.5708 | 65.45 | 87.9 | 88.525 | 89.5
2.5 | 7.3 | 82.6 | 83.3667 | 84.0333 | 11.3056 | 83.15 | 83.8 | 84.5 | 18.9 | 84 | 84.9 | 85.4 | 36.1454 | 85.75 | 86.5 | 87.171 | 56.9 | 87.5333 | 88.05 | 88.9
2.6 | 5.45 | 82.1 | 82.9 | 83.4 | 8.3041 | 82.525 | 83.2 | 84 | 13.4 | 83.4 | 84 | 84.8667 | 27.9093 | 85.2 | 85.8 | 86.6711 | 48.4 | 86.7 | 87.45 | 88.3667
2.7 | 4.35 | 81.3 | 82.35 | 82.9 | 6.5032 | 81.9 | 82.6667 | 83.4333 | 10.6 | 82.8333 | 83.6 | 84.3667 | 20.907 | 84.6667 | 85.3 | 86.0913 | 39.2 | 86.3 | 87 | 87.9
2.8 | 3.6 | 80.7 | 81.65 | 82.35 | 5.4027 | 81.4 | 82.25 | 83.05 | 8.6 | 82.35 | 83.125 | 83.8667 | 16.1054 | 84.1 | 84.95 | 85.6714 | 31.3 | 86 | 86.5667 | 87.5
2.9 | 3.1 | 80.14 | 81.24 | 81.825 | 4.669 | 80.8 | 81.65 | 82.6 | 7.3 | 81.6 | 82.5 | 83.35 | 12.9377 | 83.575 | 84.4667 | 85.1466 | 24.2 | 85.4 | 86.2 | 87.0667
3 | 2.7 | 79.575 | 80.75 | 81.35 | 4.3022 | 80.0667 | 81.1 | 82 | 6.3 | 81.125 | 82.15 | 82.7 | 10.9036 | 82.7667 | 84.05 | 84.5218 | 18.9333 | 84.7 | 85.7 | 86.55
3.1 | 2.45 | 78.825 | 80.1667 | 80.8 | 3.7018 | 79.4333 | 80.625 | 81.35 | 5.7 | 80.575 | 81.56 | 82.2 | 9.4532 | 82.4333 | 83.62 | 84.047 | 15.1 | 84.25 | 85.5 | 86.22
3.2 | 2.3 | 78.3 | 79.55 | 80.28 | 3.4017 | 78.9 | 80.0333 | 80.925 | 5.3 | 79.96 | 81 | 81.7333 | 8.4028 | 81.7667 | 83.05 | 83.6971 | 13.1 | 83.7 | 84.7 | 85.65
3.3 | 2.2 | 77.8 | 79.075 | 79.775 | 3.3016 | 78.4667 | 79.5667 | 80.4 | 4.85 | 79.275 | 80.55 | 81.275 | 7.8026 | 81.1 | 82.375 | 83.3722 | 11.3667 | 83.1 | 84 | 85.15
3.4 | 2.1 | 77.2 | 78.46 | 79.275 | 3.1016 | 77.75 | 79 | 79.875 | 4.6 | 78.6333 | 79.95 | 80.9 | 7.3024 | 80.4 | 81.9667 | 82.7224 | 10.6 | 82.6 | 83.5667 | 84.85
3.5 | 2 | 76.7 | 78.05 | 78.8 | 2.9015 | 77.1 | 78.5 | 79.4667 | 4.4 | 77.95 | 79.5 | 80.325 | 6.9023 | 80 | 81.45 | 82.3725 | 9.75 | 82.2 | 83.3 | 84.4
3.6 | 1.9 | 76 | 77.7 | 78.2667 | 2.7014 | 76.6 | 78.1333 | 79 | 4.1 | 77.45 | 79.1 | 79.825 | 6.6522 | 79.5667 | 80.95 | 82.0726 | 9.15 | 81.6 | 83.05 | 83.95
3.7 | 1.9 | 75.5333 | 77.1 | 78 | 2.6013 | 75.95 | 77.5667 | 78.6 | 3.9333 | 76.8667 | 78.5333 | 79.425 | 6.3021 | 79.1 | 80.4333 | 81.4728 | 8.8 | 81.05 | 82.45 | 83.4
3.8 | 1.8 | 75.1 | 76.5333 | 77.4333 | 2.5013 | 75.42 | 77.1333 | 77.95 | 3.75 | 76.25 | 78.1 | 78.9 | 6.052 | 78.5 | 79.95 | 81.023 | 8.5 | 80.5 | 82.05 | 83.0667
3.9 | 1.7 | 74.4 | 76.0667 | 77.1 | 2.3512 | 74.875 | 76.6333 | 77.6 | 3.6 | 75.82 | 77.6 | 78.4667 | 5.902 | 77.9 | 79.6 | 80.6231 | 8.2 | 80 | 81.7 | 82.7333
4 | 1.6 | 73.95 | 75.7 | 76.5 | 2.2011 | 74.25 | 76.2 | 77.1 | 3.4 | 75.18 | 77.0667 | 78.1 | 5.7019 | 77.3 | 79.15 | 80.2732 | 7.8 | 79.35 | 81.08 | 82.375
4.1 | 1.6 | 73.15 | 75.3 | 76.0667 | 2.1511 | 73.575 | 75.8 | 76.5667 | 3.2 | 74.55 | 76.6333 | 77.5 | 5.5018 | 76.8 | 78.64 | 79.6234 | 7.6 | 78.6 | 80.6 | 81.9667
4.2 | 1.55 | 72.5 | 74.8 | 75.7 | 2.101 | 72.8333 | 75.3333 | 76.0667 | 3.2 | 73.9333 | 76.0333 | 76.95 | 5.4018 | 76.35 | 78.15 | 79.2736 | 7.3 | 78.325 | 80.2333 | 81.5
4.3 | 1.5 | 71.6 | 74.2667 | 75.0667 | 2.001 | 72.1 | 74.7 | 75.8 | 3.1 | 73.0667 | 75.6667 | 76.55 | 5.2017 | 75.4667 | 77.675 | 78.8237 | 7.2 | 77.825 | 79.8 | 81.075
4.4 | 1.4 | 70.5 | 73.8 | 74.7 | 1.9009 | 71.4 | 74.25 | 75.275 | 2.9667 | 72.375 | 75.175 | 76.1333 | 4.9016 | 74.925 | 77.2 | 78.4139 | 7 | 77.275 | 79.4333 | 80.7333
4.5 | 1.4 | 69.2 | 73.3333 | 74.3 | 1.8009 | 70.2 | 73.9 | 74.875 | 2.9 | 71.6 | 74.675 | 75.75 | 4.8016 | 74.375 | 76.85 | 77.924 | 6.8 | 76.95 | 78.9 | 80.3667
4.6 | 1.4 | 67.1 | 72.8 | 73.95 | 1.8009 | 69.15 | 73.5 | 74.275 | 2.8 | 71 | 74.4 | 75.3 | 4.6015 | 73.7 | 76.425 | 77.5991 | 6.7 | 76.18 | 78.5 | 80.15
4.7 | 1.3 | 63.95 | 72.35 | 73.375 | 1.7008 | 67.5 | 73.025 | 73.925 | 2.7 | 70 | 73.8 | 74.8 | 4.4015 | 73.3 | 75.85 | 77.1243 | 6.5 | 75.7667 | 78.05 | 79.5
4.8 | 1.3 | 58.6 | 71.95 | 72.95 | 1.7008 | 64.95 | 72.475 | 73.46 | 2.6 | 68.9333 | 73.35 | 74.45 | 4.4015 | 72.6 | 75.475 | 76.6344 | 6.3 | 75.2 | 77.82 | 78.95
4.9 | 1.3 | 52.2 | 71.575 | 72.5333 | 1.7008 | 61.18 | 71.925 | 73.1 | 2.5 | 67.4667 | 73 | 74 | 4.3014 | 71.8 | 75.05 | 76.1746 | 6.2 | 74.8 | 77.575 | 78.6333
5 | 1.2 | 44.0333 | 70.975 | 71.9333 | 1.6008 | 56.2 | 71.38 | 72.6333 | 2.4 | 65.8 | 72.45 | 73.45 | 4.2014 | 71.2 | 74.7 | 75.7748 | 5.9 | 74.34 | 77 | 78.3
5.1 | 1.2 | 35.7 | 70.56 | 71.55 | 1.5008 | 49.25 | 71.08 | 72.15 | 2.3333 | 63 | 71.925 | 72.95 | 4.0013 | 70.5 | 74.4 | 75.3999 | 5.8 | 73.4 | 76.45 | 78
5.2 | 1.2 | 28.1667 | 69.9667 | 71 | 1.5008 | 41.425 | 70.5 | 71.7 | 2.3 | 58.9667 | 71.5333 | 72.625 | 3.9013 | 69.5 | 73.8 | 74.975 | 5.7 | 72.9667 | 76.15 | 77.5
5.3 | 1.2 | 20.65 | 69.6 | 70.5 | 1.4007 | 33.9667 | 70 | 71.1 | 2.2 | 53.2 | 70.98 | 71.9667 | 3.9013 | 68.4 | 73.5 | 74.6751 | 5.6 | 72.3667 | 75.65 | 77.05
5.4 | 1.1 | 14.8333 | 69.02 | 70.1 | 1.4007 | 26.8 | 69.5 | 70.625 | 2.2 | 46.475 | 70.5667 | 71.6 | 3.8013 | 67.22 | 72.9 | 74.1086 | 5.6 | 71.6667 | 75.4 | 76.5
5.5 | 1.1 | 10.9 | 68.6 | 69.6 | 1.4007 | 20.45 | 69.1 | 70.2 | 2.1 | 39.725 | 70.1 | 71.15 | 3.7012 | 65.2 | 72.4667 | 73.5255 | 5.4 | 71 | 75.1 | 76.2
5.6 | 1.1 | 8.7333 | 68 | 69.02 | 1.4007 | 15.1 | 68.5667 | 69.7667 | 2.1 | 32 | 69.55 | 70.7 | 3.6012 | 62.4 | 71.875 | 73.2006 | 5.2 | 70.0333 | 74.6 | 75.8333
5.7 | 1.1 | 7.5 | 67.65 | 68.56 | 1.4007 | 12 | 68.25 | 69.2333 | 2 | 25.775 | 69 | 70.15 | 3.5012 | 58.55 | 71.55 | 72.8007 | 5.1 | 69.5333 | 74.2 | 75.45
5.8 | 1.1 | 6.6 | 67.1 | 68.15 | 1.3007 | 10.1 | 67.5667 | 68.75 | 1.9 | 20.325 | 68.6333 | 69.7 | 3.4011 | 53.375 | 70.9667 | 72.2759 | 5 | 68.7 | 73.75 | 74.875
5.9 | 1.1 | 6.15 | 66.6333 | 67.75 | 1.3007 | 8.8667 | 67.15 | 68.35 | 1.9 | 16.3 | 68.1 | 69.3 | 3.4011 | 48 | 70.575 | 71.7261 | 4.9 | 67.1333 | 73.3 | 74.6333
6 | 1.1 | 5.8 | 66.25 | 67.225 | 1.3007 | 7.95 | 66.7 | 67.75 | 1.8 | 13.65 | 67.55 | 68.825 | 3.2011 | 41.1 | 70.325 | 71.2762 | 4.9 | 65.3 | 72.8667 | 74.2
6.1 | 1.1 | 5.475 | 65.5 | 66.7 | 1.3007 | 7.46 | 66.02 | 67.25 | 1.8 | 12.1 | 67.15 | 68.1 | 3.1511 | 35 | 69.8 | 70.9097 | 4.8 | 62.9 | 72.3 | 73.7
6.2 | 1 | 5.25 | 65.0667 | 66.05 | 1.2006 | 6.9667 | 65.6 | 66.8 | 1.8 | 10.9667 | 66.575 | 67.7667 | 3.001 | 29 | 69.4 | 70.4515 | 4.7 | 59.4 | 71.85 | 73.3667
6.3 | 1 | 5.1 | 64.55 | 65.625 | 1.2006 | 6.6 | 65.15 | 66.15 | 1.7 | 10.2 | 66.1 | 67.25 | 3.001 | 24 | 68.75 | 70.0267 | 4.6 | 54.9 | 71.3333 | 72.9
6.4 | 1 | 4.9667 | 64.025 | 65.15 | 1.2006 | 6.3667 | 64.475 | 65.8 | 1.7 | 9.65 | 65.575 | 66.7 | 2.901 | 20.1 | 68.3 | 69.4102 | 4.6 | 50 | 70.9 | 72.45
6.5 | 0.9 | 4.75 | 63.55 | 64.6 | 1.2006 | 6.1 | 64.02 | 65.1 | 1.6 | 9.2 | 64.975 | 66.15 | 3.001 | 17.45 | 67.775 | 68.977 | 4.4 | 43.8 | 70.625 | 72
6.6 | 0.9 | 4.6667 | 62.875 | 64.0333 | 1.2006 | 5.9 | 63.4333 | 64.65 | 1.6 | 8.7333 | 64.625 | 65.725 | 3.001 | 16.1 | 67.3 | 68.5771 | 4.4 | 37.75 | 70.15 | 71.8
6.7 | 0.9 | 4.6 | 62.24 | 63.44 | 1.1005 | 5.6333 | 63 | 64.05 | 1.5 | 8.4 | 64.05 | 65.2333 | 3.001 | 14.9 | 66.9 | 68.0773 | 4.2 | 32.1 | 69.75 | 71.3
6.8 | 0.9 | 4.5 | 61.7 | 62.7667 | 1.1005 | 5.5 | 62.2 | 63.55 | 1.5 | 8.15 | 63.45 | 64.55 | 2.901 | 13.975 | 66.35 | 67.5275 | 4.1 | 27.275 | 69.26 | 70.8
6.9 | 0.9 | 4.4 | 60.9 | 62.1 | 1.1005 | 5.375 | 61.55 | 62.95 | 1.5 | 7.9 | 62.9 | 64 | 2.8009 | 13.3333 | 65.9 | 66.9777 | 4 | 23.3667 | 68.85 | 70.225
7 | 0.9 | 4.3 | 59.75 | 61 | 1.1005 | 5.2 | 60.9 | 62.25 | 1.5 | 7.7 | 62.3 | 63.55 | 2.7009 | 12.85 | 65.36 | 66.6278 | 4 | 20.8 | 68.4 | 69.8
7.1 |  | 4.2 | 58.4 | 59.7 | 1.1005 | 5.1 | 60.1 | 61.475 | 1.5 | 7.5 | 61.7 | 62.98 | 2.5008 | 12.3667 | 64.92 | 65.978 | 4 | 19.3 | 67.9 | 69.52
7.2 |  | 4.1 | 56.22 | 57.6 |  | 5 | 59.2 | 60.475 | 1.4 | 7.25 | 61 | 62.4 | 2.5008 | 11.95 | 64.375 | 65.4532 | 4 | 18.0333 | 67.4 | 69.05
7.3 |  | 4.025 | 53 | 54.2 |  | 4.8 | 57.95 | 59.4 |  | 7.1 | 60.3 | 61.7 | 2.5008 | 11.7 | 63.85 | 65.0283 | 3.9 | 17.2 | 66.98 | 68.525
7.4 |  | 4 | 47.875 | 50 |  | 4.7 | 56.3 | 57.6667 |  | 6.9 | 59.7 | 60.85 |  | 11.3 | 63.35 | 64.5285 | 3.9 | 16.3333 | 66.45 | 67.9
7.5 |  | 3.9 | 41.875 | 43.9 |  | 4.6 | 53.775 | 55.2667 |  | 6.675 | 58.9667 | 60.275 |  | 11.1 | 62.76 | 64.0453 | 3.8 | 15.9 | 66.05 | 67.6667
7.6 |  | 3.9 | 35.325 | 38.025 |  | 4.5 | 49.825 | 52 |  | 6.525 | 57.8 | 59.3 |  | 10.8333 | 62.175 | 63.5455 | 3.8 | 15.3 | 65.65 | 67.1
7.7 |  | 3.8 | 28.325 | 32.025 |  | 4.4 | 44.825 | 47.0667 |  | 6.35 | 56.55 | 58.0667 |  | 10.6 | 61.7 | 62.9457 |  | 14.9667 | 65.15 | 66.7333
7.8 |  | 3.7 | 22.3 | 25.8667 |  | 4.3667 | 39.3 | 41.6 |  | 6.2 | 54.8333 | 56.4 |  | 10.3 | 61.1 | 62.6458 |  | 14.66 | 64.625 | 66.1667
7.9 |  | 3.6333 | 17.2 | 20.925 |  | 4.3 | 33.3 | 35.8 |  | 6.025 | 52.3667 | 53.9 |  | 10.16 | 60.5 | 62.0793 |  | 14.42 | 64.2 | 65.6
8 |  | 3.6 | 13.2 | 17.32 |  | 4.2 | 27.05 | 29.75 |  | 5.9 | 48.3 | 50.5667 |  | 9.9 | 59.8 | 61.1796 |  | 14 | 63.7 | 65.25
8.1 |  | 3.6 | 10.95 | 15.1 |  | 4.1 | 21.8 | 24.9 |  | 5.75 | 43.4 | 45.8667 |  | 9.6333 | 59.2 | 60.6798 |  | 13.65 | 63.2333 | 64.7667
8.2 |  | 3.6 | 9.7 | 13.6 |  | 4.025 | 16.85 | 20.65 |  | 5.7 | 38.1 | 40.5 |  | 9.45 | 58.5 | 59.93 |  | 13.4 | 62.5667 | 64.3333
8.3 |  | 3.6 | 8.75 | 12.9 |  | 4 | 13.6 | 17.475 |  | 5.6 | 32.7333 | 35.4 |  | 9.3 | 57.5 | 59.1803 |  | 13.275 | 62.1 | 63.9
8.4 |  | 3.6 | 8.3 | 12.4 |  | 3.9 | 11.7 | 15.75 |  | 5.475 | 27.1 | 30.3333 |  | 9 | 56.55 | 58.2139 |  | 12.9333 | 61.65 | 63.3
8.5 |  | 3.55 | 8 | 12.0333 |  | 3.9 | 10.725 | 14.6 |  | 5.3 | 22.0667 | 25.75 |  | 8.9 | 55.075 | 56.7144 |  | 12.6 | 61.15 | 62.85
8.6 |  | 3.5 | 7.7 | 11.8 |  | 3.9 | 9.9 | 13.88 |  | 5.225 | 18.1 | 22.0667 |  | 8.675 | 53.125 | 54.9067 |  | 12.28 | 60.6333 | 62.25
8.7 |  | 3.5 | 7.56 | 11.575 |  | 3.8 | 9.35 | 13.3 |  | 5.1333 | 15.6 | 19.5667 |  | 8.5 | 50.25 | 52.3825 |  | 12.15 | 59.9667 | 61.75
8.8 |  | 3.5 | 7.425 | 11.35 |  | 3.8 | 9.025 | 12.96 |  | 5.1 | 13.8 | 17.9333 |  | 8.3333 | 46.525 | 48.7504 |  | 11.9 | 59.3 | 61.15
8.9 |  | 3.425 | 7.3 | 11.225 |  | 3.8 | 8.675 | 12.7 |  | 5 | 12.7 | 17 |  | 8.225 | 42.14 | 44.3519 |  | 11.7333 | 58.5333 | 60.5
9 |  | 3.4 | 7.2 | 11.15 |  | 3.8 | 8.425 | 12.45 |  | 4.975 | 12.0333 | 16.2333 |  | 8.1 | 36.98 | 40.1866 |  | 11.55 | 57.9 | 59.7
9.1 |  | 3.4 | 7.175 | 11.1 |  | 3.7 | 8.3 | 12.2 |  | 4.9 | 11.5 | 15.6667 |  | 8 | 32.4 | 35.7214 |  | 11.425 | 56.92 | 59.0333
9.2 |  | 3.4 | 7.1 | 11 |  | 3.7 | 8.15 | 12.05 |  | 4.9 | 10.9667 | 15.3 |  | 7.8 | 27.6 | 31.2896 |  | 11.2333 | 55.525 | 58
9.3 |  | 3.3 | 7.1 | 11 |  | 3.6 | 8 | 11.92 |  | 4.8 | 10.7 | 15 |  | 7.65 | 23.6 | 27.6241 |  | 11 | 54 | 56.3
9.4 |  | 3.3 | 7.1 | 10.9 |  | 3.6 | 7.9 | 11.8 |  | 4.8 | 10.5333 | 14.7333 |  | 7.52 | 20.5 | 24.7918 |  | 10.8333 | 51.85 | 54.5
9.5 |  | 3.3 | 7.1 | 10.9 |  | 3.6 | 7.8 | 11.7 |  | 4.7 | 10.3 | 14.425 |  | 7.4 | 18.4 | 23.1923 |  | 10.6667 | 48.8333 | 51.7
9.6 |  | 3.2667 | 7 | 10.9 |  | 3.6 | 7.7 | 11.575 |  | 4.6 | 10 | 14.25 |  | 7.35 | 17.2667 | 21.8927 |  | 10.5 | 45.3 | 48.3
9.7 |  | 3.2 | 7 | 10.9 |  | 3.5 | 7.6 | 11.5 |  | 4.5 | 9.8 | 14.025 |  | 7.2 | 16.275 | 21.068 |  | 10.4 | 40.7 | 44.55
9.8 |  | 3.2 | 7 | 10.85 |  | 3.4 | 7.55 | 11.4 |  | 4.5 | 9.6 | 13.8 |  | 7.2 | 15.55 | 20.3932 |  | 10.2667 | 36.8 | 40.625
9.9 |  | 3.1 | 7 | 10.8 |  | 3.4 | 7.5 | 11.3 |  | 4.5 | 9.475 | 13.625 |  | 7.1 | 15.125 | 19.8434 |  | 9.975 | 32.1333 | 36.5667
10 |  | 3.1 | 7 | 10.8 |  | 3.325 | 7.4333 | 11.3 |  | 4.4 | 9.3 | 13.5 |  | 7 | 14.68 | 19.4602 |  | 9.74 | 28.3 | 33.025
10.1 |  |  | 6.975 | 10.8 |  | 3.3 | 7.4 | 11.2 |  | 4.375 | 9.2 | 13.4 |  | 6.9 | 14.4 | 19.1436 |  | 9.5 | 24.8667 | 30.05
10.2 |  |  | 6.9 | 10.7 |  | 3.3 | 7.38 | 11.2 |  | 4.3 | 9.1 | 13.3 |  | 6.8 | 13.98 | 18.7937 |  | 9.475 | 22.7 | 27.825
10.3 |  |  | 6.9 | 10.7 |  |  | 7.3 | 11.2 |  | 4.2667 | 9 | 13.175 |  | 6.7 | 13.6 | 18.5271 |  | 9.4 | 20.8 | 26.3
10.4 |  |  | 6.8 | 10.6 |  |  | 7.2 | 11.1 |  | 4.2 | 8.825 | 13 |  | 6.6667 | 13.45 | 18.194 |  | 9.3 | 19.66 | 25.3
10.5 |  |  | 6.8 | 10.6 |  |  | 7.2 | 10.9 |  |  | 8.8 | 12.9 |  | 6.6 | 13.2 | 17.994 |  | 9.2 | 18.95 | 24.4
10.6 |  |  | 6.7 | 10.5 |  |  | 7.2 | 10.9 |  |  | 8.7 | 12.8 |  | 6.5 | 13.05 | 17.6691 |  | 9.2 | 18.35 | 23.9333
10.7 |  |  | 6.7 | 10.475 |  |  | 7.1 | 10.8 |  |  | 8.6 | 12.7 |  | 6.4 | 12.9 | 17.4942 |  | 9 | 17.85 | 23.4333
10.8 |  |  | 6.6 | 10.4 |  |  | 7 | 10.8 |  |  | 8.6 | 12.6 |  | 6.3 | 12.65 | 17.3192 |  | 9 | 17.4 | 23.05
10.9 |  |  | 6.5 | 10.3 |  |  | 6.92 | 10.7 |  |  | 8.4 | 12.525 |  | 6.2 | 12.4 | 17.0943 |  | 8.95 | 17.175 | 22.75
11 |  |  | 6.45 | 10.2 |  |  | 6.9 | 10.6 |  |  | 8.3 | 12.4 |  |  | 12.3 | 16.9943 |  | 8.775 | 16.825 | 22.55
11.1 |  |  | 6.4 | 10.2 |  |  | 6.8 | 10.5 |  |  | 8.2 | 12.35 |  |  | 12.1 | 16.8444 |  | 8.725 | 16.58 | 22.2
11.2 |  |  | 6.3 | 10.1 |  |  | 6.8 | 10.5 |  |  | 8.1 | 12.2 |  |  | 11.9 | 16.6444 |  | 8.625 | 16.25 | 21.8333
11.3 |  |  | 6.2 | 10 |  |  | 6.7 | 10.4 |  |  | 8 | 12.1 |  |  | 11.7 | 16.3945 |  | 8.3333 | 16.0333 | 21.5
11.4 |  |  | 6.2 | 9.9 |  |  | 6.6 | 10.3 |  |  | 7.9667 | 11.9 |  |  | 11.6 | 16.2613 |  | 8.3 | 15.85 | 21.2
11.5 |  |  | 6.1 | 9.825 |  |  | 6.5333 | 10.2 |  |  | 7.8 | 11.8 |  |  | 11.4 | 16.0946 |  |  | 15.6 | 21
11.6 |  |  | 6 | 9.8 |  |  | 6.5 | 10.1 |  |  | 7.775 | 11.7 |  |  | 11.3333 | 15.8947 |  |  | 15.3 | 20.8
11.7 |  |  | 5.9667 | 9.6333 |  |  | 6.4 | 10.05 |  |  | 7.7 | 11.6 |  |  | 11.2 | 15.7947 |  |  | 15.0667 | 20.625
11.8 |  |  | 5.9 | 9.575 |  |  | 6.275 | 9.9 |  |  | 7.6 | 11.52 |  |  | 11 | 15.6198 |  |  | 14.9 | 20.375
11.9 |  |  | 5.825 | 9.5 |  |  | 6.2 | 9.875 |  |  | 7.475 | 11.425 |  |  | 10.8 | 15.3449 |  |  | 14.7 | 20.125
12 |  |  | 5.775 | 9.375 |  |  | 6.1 | 9.8 |  |  | 7.4 | 11.35 |  |  | 10.7667 | 15.2282 |  |  | 14.55 | 19.925
12.1 |  |  | 5.7 | 9.3 |  |  | 6.1 | 9.65 |  |  | 7.3 | 11.25 |  |  | 10.6 | 15.095 |  |  | 14.4 | 19.7
12.2 |  |  | 5.6 | 9.2 |  |  | 6 | 9.6 |  |  | 7.2 | 11.1 |  |  | 10.55 | 14.995 |  |  | 14.1667 | 19.5
12.3 |  |  | 5.6 | 9.1 |  |  | 5.9 | 9.5 |  |  | 7.1 | 11 |  |  | 10.4 | 14.895 |  |  | 13.9 | 19.3
12.4 |  |  | 5.5 | 9 |  |  | 5.8 | 9.4 |  |  | 7.1 | 10.9 |  |  | 10.25 | 14.7951 |  |  | 13.775 | 19.2
12.5 |  |  | 5.4 | 8.9 |  |  | 5.8 | 9.3 |  |  | 6.95 | 10.8 |  |  | 10.1 | 14.5951 |  |  | 13.68 | 19
12.6 | [6 empty] | 5.66 | 9.2 |  |  | 6.9 | 10.7 |  |  | 10 | 14.4952 |  |  | 13.525 | 18.9
12.7 | [6 empty] | 5.6 | 9.1 |  |  | 6.8 | 10.6 |  |  | 9.9 | 14.1953 |  |  | 13.3 | 18.625
12.8 | [6 empty] | 5.5 | 9.025 |  |  | 6.7 | 10.5 |  |  | 9.8 | 14.1153 |  |  | 13.14 | 18.6
12.9 | [10 empty] | 6.6 | 10.4 |  |  | 9.7 | 13.9953 |  |  | 13 | 18.35
13 | [10 empty] | 6.5 | 10.3 |  |  | 9.6 | 13.8954 |  |  | 12.75 | 18.2
13.1 | [10 empty] | 6.4 | 10.2 |  |  | 9.4 | 13.7621 |  |  | 12.62 | 18.0333
13.2 | [10 empty] | 6.3 | 10.1 |  |  | 9.3 | 13.6621 |  |  | 12.56 | 17.9
13.3 | [14 empty] | 9.1 | 13.4705 |  |  | 12.4 | 17.65
13.4 | [14 empty] | 9.0667 | 13.2956 |  |  | 12.3 | 17.55
13.5 | [14 empty] | 8.9 | 13.1756 |  |  | 12.2 | 17.35
13.6 | [14 empty] | 8.8 | 13.0956 |  |  | 12.1 | 17.2
13.7 | [14 empty] | 8.7 | 12.9957 |  |  | 11.9 | 17.0667
13.8 | [14 empty] | 8.6 | 12.8357 |  |  | 11.8 | 16.925
13.9 | [14 empty] | 8.575 | 12.7458 |  |  | 11.7 | 16.6667
14 | [18 empty] | 11.6 | 16.5
14.1 | [18 empty] | 11.45 | 16.325
14.2 | [18 empty] | 11.3 | 16.2
14.3 | [18 empty] | 11.2 | 16.1
14.4 | [18 empty] | 11.1 | 16
14.5 | [18 empty] | 10.95 | 15.8
14.6 | [18 empty] | 10.8 | 15.7
14.7 | [18 empty] | 10.8 | 15.475

## Wedge 60 Longitudinal Profiles
Longitudinal Profiles
Wedge 60
SSD=100cm
Field Size=40x15cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 12.6
-30.8 |  |  |  |  | 12.6667
-30.7 |  |  |  |  | 12.725
-30.6 |  |  |  |  | 12.9375
-30.5 |  |  |  |  | 13.0625
-30.4 |  |  |  |  | 13.22
-30.3 |  |  |  |  | 13.4
-30.2 |  |  |  |  | 13.4571
-30.1 |  |  |  |  | 13.7
-30 |  |  |  |  | 13.825
-29.9 |  |  |  |  | 13.95
-29.8 |  |  |  |  | 14.15
-29.7 |  |  |  |  | 14.3
-29.6 |  |  |  |  | 14.5375
-29.5 |  |  |  |  | 14.6714
-29.4 |  |  |  |  | 14.95
-29.3 |  |  |  |  | 15.1571
-29.2 |  |  |  |  | 15.3625
-29.1 |  |  |  |  | 15.4875
-29 |  |  |  |  | 15.5
-28.9 |  |  |  | 10.6 | 15.85
-28.8 |  |  |  | 10.7 | 16.0571
-28.7 |  |  |  | 10.825 | 16.2286
-28.6 |  |  |  | 10.95 | 16.4
-28.5 |  |  |  | 11.0857 | 16.6889
-28.4 |  |  |  | 11.26 | 16.8833
-28.3 |  |  |  | 11.425 | 17.1286
-28.2 |  |  |  | 11.5 | 17.4333
-28.1 |  |  |  | 11.5667 | 17.64
-28 |  |  |  | 11.8 | 18
-27.9 |  |  |  | 11.9125 | 18.18
-27.8 |  |  |  | 12.05 | 18.4429
-27.7 |  |  |  | 12.2143 | 18.65
-27.6 |  |  |  | 12.3667 | 18.9286
-27.5 |  |  |  | 12.5 | 19.2
-27.4 |  |  |  | 12.65 | 19.5
-27.3 |  |  |  | 12.775 | 19.9778
-27.2 |  |  |  | 13.0143 | 20.3667
-27.1 |  |  |  | 13.2333 | 20.78
-27 |  |  |  | 13.42 | 21.2
-26.9 |  |  | 8.5 | 13.6 | 21.6667
-26.8 |  |  | 8.6 | 13.825 | 22.4167
-26.7 |  |  | 8.6857 | 13.9889 | 23.48
-26.6 |  |  | 8.8 | 14.225 | 24.9
-26.5 |  |  | 8.9 | 14.425 | 26.5
-26.4 |  |  | 8.9 | 14.6857 | 28.65
-26.3 |  |  | 9 | 14.875 | 31.6286
-26.2 |  |  | 9.1 | 15.125 | 35.2833
-26.1 |  |  | 9.2167 | 15.375 | 39.1
-26 |  |  | 9.3 | 15.625 | 43.6
-25.9 |  | 7.5 | 9.4429 | 15.9 | 48.5
-25.8 |  | 7.6 | 9.5 | 16.1 | 53.325
-25.7 |  | 7.6286 | 9.625 | 16.48 | 57.8857
-25.6 |  | 7.7 | 9.7 | 16.7667 | 62
-25.5 |  | 7.8 | 9.86 | 17.1 | 65.1625
-25.4 |  | 7.8429 | 9.94 | 17.4 | 67.6
-25.3 |  | 7.9 | 10.1 | 17.6667 | 69.5125
-25.2 | 7.2 | 7.9 | 10.25 | 17.9857 | 70.9167
-25.1 | 7.2571 | 8 | 10.4667 | 18.4857 | 72.0833
-25 | 7.3167 | 8.0833 | 10.525 | 18.95 | 73.0714
-24.9 | 7.4 | 8.1 | 10.75 | 19.45 | 73.6
-24.8 | 7.4 | 8.1667 | 10.8 | 20.2 | 74.15
-24.7 | 7.5 | 8.25 | 10.9571 | 21.2 | 74.7
-24.6 | 7.5 | 8.3 | 11.1 | 22.3 | 75.15
-24.5 | 7.5 | 8.4 | 11.35 | 24.5857 | 75.5333
-24.4 | 7.5125 | 8.4625 | 11.6143 | 27.075 | 76
-24.3 | 7.6 | 8.5 | 11.78 | 30.5 | 76.3667
-24.2 | 7.6333 | 8.6 | 12.0429 | 34.675 | 76.6429
-24.1 | 7.7 | 8.7 | 12.1857 | 39.55 | 77
-24 | 7.7 | 8.7667 | 12.3333 | 45.3143 | 77.2857
-23.9 | 7.75 | 8.9167 | 12.625 | 50.7333 | 77.6
-23.8 | 7.8 | 9.1 | 12.875 | 56.0286 | 77.78
-23.7 | 7.8857 | 9.16 | 13.1167 | 61.6857 | 78.1714
-23.6 | 7.9 | 9.3286 | 13.3 | 65.725 | 78.5143
-23.5 | 7.9 | 9.4 | 13.6143 | 69.1333 | 78.7333
-23.4 | 8 | 9.5333 | 13.8667 | 72.1143 | 79
-23.3 | 8 | 9.7833 | 14.1714 | 74.0667 | 79.1857
-23.2 | 8.1 | 9.95 | 14.55 | 75.4333 | 79.45
-23.1 | 8.1833 | 10.1 | 14.925 | 76.6 | 79.75
-23 | 8.26 | 10.2 | 15.52 | 77.2857 | 80
-22.9 | 8.3 | 10.4333 | 16 | 77.9 | 80.325
-22.8 | 8.45 | 10.55 | 16.8 | 78.6286 | 80.625
-22.7 | 8.6 | 10.7333 | 17.84 | 79.1333 | 80.875
-22.6 | 8.7 | 10.9714 | 19.2571 | 79.4875 | 81.0833
-22.5 | 8.8 | 11.225 | 21.4875 | 79.8625 | 81.3375
-22.4 | 8.9 | 11.4 | 24.6333 | 80.1375 | 81.65
-22.3 | 9 | 11.65 | 28.9 | 80.5125 | 81.74
-22.2 | 9.2 | 11.9667 | 33.9857 | 80.8667 | 81.9
-22.1 | 9.3333 | 12.3 | 39.775 | 81.1857 | 82.1286
-22 | 9.5 | 12.7143 | 46.4429 | 81.5333 | 82.3667
-21.9 | 9.6333 | 13.1833 | 53.75 | 81.8857 | 82.6333
-21.8 | 9.7333 | 13.975 | 60.6286 | 82.1625 | 82.8667
-21.7 | 9.95 | 14.8 | 67.175 | 82.4 | 83.15
-21.6 | 10.125 | 16.5429 | 72.15 | 82.6571 | 83.34
-21.5 | 10.36 | 18.675 | 76.2 | 82.9429 | 83.6
-21.4 | 10.6571 | 22.25 | 78.95 | 83.2857 | 83.8333
-21.3 | 10.9667 | 27.0667 | 80.75 | 83.5 | 84.0167
-21.2 | 11.42 | 33.1429 | 82.1 | 83.7667 | 84.1833
-21.1 | 12.0429 | 40.0286 | 83.0833 | 84.0429 | 84.3286
-21 | 12.9833 | 47.2 | 83.8125 | 84.2714 | 84.5667
-20.9 | 14.5857 | 55.45 | 84.24 | 84.4857 | 84.7286
-20.8 | 17.1143 | 63.2714 | 84.7 | 84.7333 | 84.9667
-20.7 | 20.9714 | 70.5571 | 85.2333 | 85.025 | 85.25
-20.6 | 26.3 | 76.22 | 85.65 | 85.2 | 85.56
-20.5 | 33.3429 | 79.9333 | 85.8857 | 85.4143 | 85.6429
-20.4 | 40.9333 | 82.8857 | 86.2571 | 85.5778 | 85.7857
-20.3 | 49.2 | 84.7 | 86.6 | 85.8 | 85.95
-20.2 | 57.95 | 86 | 86.9 | 86.05 | 86.0875
-20.1 | 66.45 | 86.8125 | 87.0667 | 86.3333 | 86.2571
-20 | 74.0143 | 87.45 | 87.4857 | 86.5778 | 86.4833
-19.9 | 79.7333 | 87.9 | 87.8 | 86.7286 | 86.725
-19.8 | 83.9 | 88.3667 | 88.0667 | 86.9571 | 86.9375
-19.7 | 86.25 | 88.7667 | 88.25 | 87.0875 | 87.0625
-19.6 | 87.9 | 89.05 | 88.4571 | 87.3286 | 87.1875
-19.5 | 88.9 | 89.3167 | 88.6143 | 87.5 | 87.3
-19.4 | 89.7125 | 89.4714 | 88.8333 | 87.675 | 87.55
-19.3 | 90.22 | 89.75 | 89.0333 | 87.7286 | 87.75
-19.2 | 90.5333 | 90 | 89.3 | 87.8625 | 87.9
-19.1 | 90.825 | 90.2 | 89.4714 | 88.1333 | 88.125
-19 | 91.0667 | 90.45 | 89.6286 | 88.4667 | 88.25
-18.9 | 91.3 | 90.6 | 89.8429 | 88.6167 | 88.375
-18.8 | 91.4667 | 90.75 | 90.025 | 88.7 | 88.4
-18.7 | 91.64 | 90.8875 | 90.1571 | 88.9429 | 88.5667
-18.6 | 91.8 | 91.16 | 90.3167 | 89.1 | 88.7
-18.5 | 91.95 | 91.25 | 90.525 | 89.3125 | 88.7625
-18.4 | 92.1143 | 91.3875 | 90.6875 | 89.4375 | 89
-18.3 | 92.3143 | 91.6 | 90.8333 | 89.5625 | 89.25
-18.2 | 92.5 | 91.68 | 91 | 89.8 | 89.3
-18.1 | 92.5833 | 91.8 | 91.1429 | 89.9 | 89.45
-18 | 92.64 | 91.9 | 91.3714 | 89.9857 | 89.7333
-17.9 | 92.74 | 92.15 | 91.5429 | 90.1 | 89.8
-17.8 | 93 | 92.3143 | 91.675 | 90.2286 | 89.9
-17.7 | 93.1 | 92.4 | 91.8333 | 90.3833 | 90
-17.6 | 93.2 | 92.54 | 92 | 90.525 | 90.0667
-17.5 | 93.3 | 92.7286 | 92.1833 | 90.68 | 90.1
-17.4 | 93.3833 | 92.9 | 92.2333 | 90.85 | 90.24
-17.3 | 93.55 | 93 | 92.325 | 91.0167 | 90.4333
-17.2 | 93.6 | 93.0833 | 92.5 | 91.1833 | 90.6714
-17.1 | 93.74 | 93.1 | 92.6125 | 91.2429 | 90.8286
-17 | 93.9 | 93.2333 | 92.75 | 91.3857 | 90.9
-16.9 | 93.9 | 93.3143 | 92.9 | 91.425 | 91.05
-16.8 | 94.06 | 93.4571 | 93.0375 | 91.5571 | 91.2
-16.7 | 94.2 | 93.54 | 93.1 | 91.6875 | 91.3
-16.6 | 94.2833 | 93.62 | 93.2 | 91.9125 | 91.3857
-16.5 | 94.3 | 93.8 | 93.35 | 92.0429 | 91.6
-16.4 | 94.4 | 93.9 | 93.4 | 92.1857 | 91.7
-16.3 | 94.4571 | 94.0143 | 93.5 | 92.3125 | 91.8
-16.2 | 94.5143 | 94.1667 | 93.5571 | 92.4 | 91.8889
-16.1 | 94.6667 | 94.2429 | 93.7 | 92.4857 | 92.0625
-16 | 94.7 | 94.3167 | 93.9 | 92.6 | 92.1
-15.9 | 94.75 | 94.4625 | 93.9429 | 92.75 | 92.2
-15.8 | 94.8 | 94.6 | 94 | 92.84 | 92.275
-15.7 | 95 | 94.6833 | 94.1 | 92.9143 | 92.4
-15.6 | 95.15 | 94.8 | 94.1714 | 93 | 92.6
-15.5 | 95.275 | 94.9 | 94.3167 | 93.1333 | 92.75
-15.4 | 95.3167 | 94.9857 | 94.4 | 93.3 | 92.875
-15.3 | 95.4 | 95.06 | 94.5 | 93.4125 | 93
-15.2 | 95.45 | 95.1143 | 94.5 | 93.5429 | 93.1
-15.1 | 95.5 | 95.2 | 94.6143 | 93.6167 | 93.1857
-15 | 95.5857 | 95.3 | 94.7 | 93.7833 | 93.3333
-14.9 | 95.7 | 95.3 | 94.8429 | 93.8286 | 93.4667
-14.8 | 95.7 | 95.4 | 94.9857 | 93.9 | 93.5
-14.7 | 95.8 | 95.4667 | 95.0429 | 94 | 93.5125
-14.6 | 95.8833 | 95.5 | 95.2 | 94.1167 | 93.675
-14.5 | 95.9333 | 95.5625 | 95.3 | 94.2 | 93.9111
-14.4 | 96 | 95.72 | 95.3714 | 94.3667 | 94.075
-14.3 | 96.1333 | 95.8 | 95.4167 | 94.44 | 94.2
-14.2 | 96.2 | 95.8286 | 95.5714 | 94.6 | 94.2429
-14.1 | 96.3 | 95.9 | 95.6286 | 94.6333 | 94.3
-14 | 96.3 | 95.925 | 95.7 | 94.7857 | 94.3143
-13.9 | 96.4 | 96.0667 | 95.8 | 94.9 | 94.4571
-13.8 | 96.42 | 96.14 | 95.85 | 95.0143 | 94.5143
-13.7 | 96.5 | 96.3 | 95.9 | 95.1444 | 94.6667
-13.6 | 96.6333 | 96.4 | 96.0286 | 95.2 | 94.8429
-13.5 | 96.7 | 96.5 | 96.1 | 95.3 | 94.9
-13.4 | 96.7625 | 96.6 | 96.2 | 95.38 | 95
-13.3 | 96.8 | 96.6 | 96.2667 | 95.46 | 95
-13.2 | 96.9 | 96.6 | 96.3 | 95.5 | 95
-13.1 | 96.9875 | 96.7 | 96.3667 | 95.55 | 95.1
-13 | 97.0667 | 96.7714 | 96.44 | 95.6 | 95.1333
-12.9 | 97.1 | 96.8 | 96.5143 | 95.6714 | 95.2
-12.8 | 97.2 | 96.9 | 96.6667 | 95.7143 | 95.3429
-12.7 | 97.25 | 97 | 96.7333 | 95.9 | 95.475
-12.6 | 97.3 | 97.1 | 96.8 | 96.0143 | 95.62
-12.5 | 97.4 | 97.2 | 96.8857 | 96.1 | 95.8
-12.4 | 97.4667 | 97.2 | 96.92 | 96.1125 | 95.8
-12.3 | 97.55 | 97.275 | 97.075 | 96.2 | 95.9286
-12.2 | 97.6 | 97.4 | 97.1 | 96.2667 | 96
-12.1 | 97.6 | 97.4714 | 97.2 | 96.3 | 96.1
-12 | 97.7 | 97.6 | 97.25 | 96.3667 | 96.1571
-11.9 | 97.7 | 97.6571 | 97.3857 | 96.5167 | 96.2
-11.8 | 97.8 | 97.7 | 97.4333 | 96.6 | 96.3429
-11.7 | 97.8 | 97.7 | 97.5 | 96.7 | 96.4
-11.6 | 97.95 | 97.7857 | 97.5714 | 96.72 | 96.45
-11.5 | 98.1 | 97.8375 | 97.62 | 96.8 | 96.54
-11.4 | 98.1 | 97.9 | 97.7 | 96.9 | 96.7
-11.3 | 98.26 | 98 | 97.8 | 96.9857 | 96.7
-11.2 | 98.3 | 98 | 97.9 | 97.0429 | 96.775
-11.1 | 98.4 | 98 | 97.9 | 97.1 | 96.8167
-11 | 98.4 | 98.1 | 98 | 97.1143 | 96.9
-10.9 | 98.4571 | 98.18 | 98 | 97.25 | 96.9667
-10.8 | 98.5 | 98.2429 | 98.0714 | 97.3 | 97.1
-10.7 | 98.5 | 98.3 | 98.1 | 97.4 | 97.1875
-10.6 | 98.5 | 98.4 | 98.18 | 97.4875 | 97.2167
-10.5 | 98.6 | 98.4 | 98.25 | 97.5 | 97.3
-10.4 | 98.6333 | 98.4167 | 98.3 | 97.575 | 97.3875
-10.3 | 98.7125 | 98.5 | 98.3833 | 97.7 | 97.45
-10.2 | 98.8 | 98.5 | 98.425 | 97.7375 | 97.5
-10.1 | 98.9 | 98.5333 | 98.5 | 97.8556 | 97.6
-10 | 98.9 | 98.6 | 98.5 | 98 | 97.6
-9.9 | 98.9 | 98.6571 | 98.5714 | 98 | 97.7
-9.8 | 99 | 98.8 | 98.6333 | 98.0571 | 97.8
-9.7 | 99 | 98.8571 | 98.7 | 98.1 | 97.84
-9.6 | 99.1 | 98.9 | 98.7 | 98.1857 | 98
-9.5 | 99.2 | 99 | 98.8 | 98.24 | 98
-9.4 | 99.2 | 99 | 98.8 | 98.3 | 98.1
-9.3 | 99.2857 | 99 | 98.85 | 98.3 | 98.1714
-9.2 | 99.35 | 99.0875 | 98.9 | 98.3167 | 98.2125
-9.1 | 99.4 | 99.2 | 99 | 98.3286 | 98.3
-9 | 99.4 | 99.2 | 99.1 | 98.4 | 98.3
-8.9 | 99.5 | 99.3 | 99.1167 | 98.5143 | 98.4
-8.8 | 99.5333 | 99.3143 | 99.2 | 98.6 | 98.4667
-8.7 | 99.6 | 99.4 | 99.2167 | 98.7 | 98.5
-8.6 | 99.7 | 99.4 | 99.3 | 98.7 | 98.55
-8.5 | 99.7 | 99.5 | 99.3 | 98.85 | 98.675
-8.4 | 99.8 | 99.58 | 99.4 | 98.9 | 98.8
-8.3 | 99.8 | 99.6 | 99.4 | 99 | 98.8
-8.2 | 99.8333 | 99.6 | 99.4 | 99.08 | 98.9167
-8.1 | 99.9 | 99.6 | 99.42 | 99.0571 | 99
-8 | 99.9 | 99.7 | 99.5 | 99 | 99.1
-7.9 | 99.9 | 99.8 | 99.6 | 99 | 99.1
-7.8 | 100 | 99.9 | 99.7 | 99 | 99.2
-7.7 | 100 | 99.9 | 99.7 | 99 | 99.2
-7.6 | 100.1 | 100 | 99.8 | 99.1 | 99.3
-7.5 | 100.1 | 100 | 99.8 | 99.2 | 99.3
-7.4 | 100.2 | 100 | 99.9 | 99.2 | 99.3
-7.3 | 100.2 | 100 | 99.9 | 99.2 | 99.3143
-7.2 | 100.2 | 100.1 | 99.95 | 99.2875 | 99.4
-7.1 | 100.2857 | 100.1 | 100 | 99.3 | 99.4
-7 | 100.3 | 100.1 | 100 | 99.4 | 99.4286
-6.9 | 100.3 | 100.2 | 100 | 99.4 | 99.5
-6.8 | 100.3 | 100.2 | 100 | 99.5 | 99.5
-6.7 | 100.4 | 100.2 | 100.0857 | 99.5 | 99.5
-6.6 | 100.4667 | 100.3 | 100.1429 | 99.6 | 99.5857
-6.5 | 100.5 | 100.3 | 100.2 | 99.6875 | 99.6
-6.4 | 100.5 | 100.3 | 100.2 | 99.7 | 99.7
-6.3 | 100.5857 | 100.3 | 100.2333 | 99.7625 | 99.7
-6.2 | 100.5667 | 100.4 | 100.3 | 99.8 | 99.7
-6.1 | 100.6 | 100.5 | 100.3 | 99.8375 | 99.7
-6 | 100.54 | 100.5 | 100.3 | 99.9 | 99.6
-5.9 | 100.5 | 100.6 | 100.3 | 99.9 | 99.6
-5.8 | 100.5 | 100.6 | 100.3333 | 99.9 | 99.7
-5.7 | 100.5 | 100.5167 | 100.4 | 99.9 | 99.7
-5.6 | 100.5 | 100.5 | 100.4167 | 99.9 | 99.7
-5.5 | 100.6 | 100.5 | 100.5 | 99.9 | 99.8
-5.4 | 100.6 | 100.45 | 100.5 | 99.9 | 99.8
-5.3 | 100.7 | 100.4 | 100.5 | 99.9 | 99.8
-5.2 | 100.7 | 100.4 | 100.5 | 99.8667 | 99.8286
-5.1 | 100.7 | 100.5 | 100.5 | 99.8 | 99.9625
-5 | 100.7 | 100.5 | 100.5 | 99.9 | 100
-4.9 | 100.7 | 100.5714 | 100.5 | 99.9 | 100.1
-4.8 | 100.6714 | 100.6 | 100.5 | 99.9125 | 100.1
-4.7 | 100.6 | 100.5 | 100.5 | 100 | 100
-4.6 | 100.7 | 100.4 | 100.4 | 100 | 100
-4.5 | 100.7 | 100.5 | 100.5 | 100 | 100
-4.4 | 100.64 | 100.5 | 100.4 | 100 | 100
-4.3 | 100.6 | 100.475 | 100.4 | 100 | 100
-4.2 | 100.6 | 100.4 | 100.4 | 100.1 | 100
-4.1 | 100.5 | 100.4 | 100.3 | 100.1 | 100.0571
-4 | 100.5 | 100.35 | 100.2 | 100.1 | 100
-3.9 | 100.45 | 100.3 | 100.2 | 100.1 | 100.0667
-3.8 | 100.4 | 100.3 | 100.2 | 100.1 | 100.15
-3.7 | 100.4 | 100.3 | 100.2 | 100.2 | 100.2
-3.6 | 100.4 | 100.3 | 100.3 | 100.2 | 100.2
-3.5 | 100.4 | 100.3 | 100.3286 | 100.2 | 100.2571
-3.4 | 100.4 | 100.3 | 100.4 | 100.2 | 100.2
-3.3 | 100.4 | 100.3 | 100.4 | 100.2 | 100.26
-3.2 | 100.4 | 100.3 | 100.4 | 100.2333 | 100.3
-3.1 | 100.44 | 100.3 | 100.4 | 100.3 | 100.3
-3 | 100.4333 | 100.3 | 100.4 | 100.3 | 100.2429
-2.9 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
-2.8 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2875
-2.7 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
-2.6 | 100.5 | 100.4 | 100.4 | 100.4 | 100.2
-2.5 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
-2.4 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
-2.3 | 100.5 | 100.32 | 100.4 | 100.2875 | 100.2571
-2.2 | 100.5 | 100.3 | 100.4 | 100.2 | 100.3
-2.1 | 100.5 | 100.3 | 100.4 | 100.2 | 100.36
-2 | 100.5 | 100.3 | 100.3333 | 100.2 | 100.4
-1.9 | 100.5 | 100.2833 | 100.25 | 100.125 | 100.4
-1.8 | 100.46 | 100.2 | 100.2 | 100.2 | 100.3
-1.7 | 100.4 | 100.2 | 100.2 | 100.2 | 100.3
-1.6 | 100.3167 | 100.2 | 100.3 | 100.2 | 100.3
-1.5 | 100.3 | 100.2 | 100.3 | 100.2 | 100.3
-1.4 | 100.3 | 100.2 | 100.3 | 100.2 | 100.3
-1.3 | 100.3 | 100.2 | 100.2857 | 100.2 | 100.3
-1.2 | 100.2 | 100.2 | 100.2 | 100.1571 | 100.3
-1.1 | 100.2 | 100.1333 | 100.1 | 100.1 | 100.2125
-1 | 100.1429 | 100.1 | 100.1 | 100.1 | 100.1667
-0.9 | 100.1 | 100.1 | 100 | 100.1 | 100
-0.8 | 100.1 | 100.0833 | 100 | 100.1 | 100
-0.7 | 100.1 | 100 | 100 | 100.1 | 100
-0.6 | 100.0333 | 100 | 100 | 100.1 | 99.92
-0.5 | 100 | 100 | 100 | 100.04 | 99.9
-0.4 | 100 | 100 | 100 | 99.96 | 99.9
-0.3 | 100 | 100 | 100.06 | 99.9 | 99.9
-0.2 | 99.9 | 100 | 100 | 99.9667 | 99.9
-0.1 | 99.95 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 99.95 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 99.9429 | 99.9
0.3 | 100 | 100 | 100 | 99.9 | 99.9
0.4 | 100 | 100 | 100 | 99.9833 | 99.9
0.5 | 100 | 100 | 100 | 100.0667 | 99.9
0.6 | 100.0333 | 100 | 100 | 100.1 | 99.9
0.7 | 100.1 | 100 | 100 | 100.1 | 100
0.8 | 100.1 | 100.0833 | 100 | 100.1 | 100
0.9 | 100.1 | 100.1 | 100 | 100.1 | 100
1 | 100.1429 | 100.1 | 100.1 | 100.1 | 100.1333
1.1 | 100.2 | 100.1333 | 100.1 | 100.1 | 100.2
1.2 | 100.2 | 100.2 | 100.2 | 100.1 | 100.3
1.3 | 100.3 | 100.2 | 100.2143 | 100.2 | 100.3
1.4 | 100.3 | 100.2 | 100.3 | 100.2 | 100.3
1.5 | 100.3 | 100.2 | 100.3 | 100.2 | 100.3
1.6 | 100.3167 | 100.2 | 100.3 | 100.2 | 100.3
1.7 | 100.4 | 100.2 | 100.25 | 100.2 | 100.3
1.8 | 100.46 | 100.2 | 100.2 | 100.2 | 100.3
1.9 | 100.5 | 100.2833 | 100.2714 | 100.1 | 100.3714
2 | 100.5 | 100.3 | 100.36 | 100.2 | 100.4
2.1 | 100.5 | 100.3 | 100.4 | 100.2 | 100.3833
2.2 | 100.5 | 100.3 | 100.4 | 100.2 | 100.3
2.3 | 100.5 | 100.32 | 100.4 | 100.3 | 100.2
2.4 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
2.5 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2
2.6 | 100.5 | 100.4 | 100.4 | 100.4 | 100.2
2.7 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2286
2.8 | 100.5 | 100.4 | 100.4 | 100.3 | 100.3
2.9 | 100.5 | 100.4 | 100.4 | 100.3 | 100.2125
3 | 100.4333 | 100.3 | 100.4 | 100.3 | 100.2143
3.1 | 100.44 | 100.3 | 100.4 | 100.3 | 100.3
3.2 | 100.4 | 100.3 | 100.4 | 100.2 | 100.225
3.3 | 100.4 | 100.3 | 100.4 | 100.2 | 100.2333
3.4 | 100.4 | 100.3 | 100.4 | 100.2 | 100.225
3.5 | 100.4 | 100.3 | 100.4 | 100.2 | 100.2625
3.6 | 100.4 | 100.3 | 100.3 | 100.2 | 100.2
3.7 | 100.4 | 100.3 | 100.2 | 100.2 | 100.2
3.8 | 100.4 | 100.3 | 100.2 | 100.1 | 100.18
3.9 | 100.45 | 100.3 | 100.2 | 100.1 | 100.1
4 | 100.5 | 100.35 | 100.2 | 100.1 | 100
4.1 | 100.5 | 100.4 | 100.3167 | 100.1 | 100.0375
4.2 | 100.6 | 100.4 | 100.4 | 100.0714 | 100.0167
4.3 | 100.6 | 100.475 | 100.4 | 100 | 100
4.4 | 100.64 | 100.5 | 100.42 | 100 | 100
4.5 | 100.7 | 100.5 | 100.5 | 100 | 100
4.6 | 100.7 | 100.4 | 100.4 | 100 | 100
4.7 | 100.6 | 100.5 | 100.5 | 100 | 100.0714
4.8 | 100.6714 | 100.6 | 100.5 | 99.9 | 100.1
4.9 | 100.7 | 100.5714 | 100.5 | 99.9 | 100.1
5 | 100.7 | 100.5 | 100.5 | 99.9 | 100
5.1 | 100.7 | 100.5 | 100.5 | 99.8 | 99.8875
5.2 | 100.7 | 100.4 | 100.5 | 99.8833 | 99.8
5.3 | 100.7 | 100.4 | 100.5 | 99.9 | 99.8
5.4 | 100.6 | 100.45 | 100.5 | 99.9 | 99.8
5.5 | 100.6 | 100.5 | 100.5 | 99.9 | 99.8
5.6 | 100.5 | 100.5 | 100.4833 | 99.9 | 99.7125
5.7 | 100.5 | 100.5167 | 100.4 | 99.9 | 99.7
5.8 | 100.5 | 100.6 | 100.3167 | 99.9 | 99.7
5.9 | 100.5 | 100.6 | 100.3 | 99.9 | 99.6
6 | 100.54 | 100.5 | 100.3 | 99.9 | 99.6
6.1 | 100.6 | 100.5 | 100.3 | 99.8125 | 99.7
6.2 | 100.5667 | 100.4 | 100.3 | 99.8 | 99.7
6.3 | 100.5857 | 100.3 | 100.3 | 99.8 | 99.7
6.4 | 100.5 | 100.3 | 100.2 | 99.7 | 99.7
6.5 | 100.5 | 100.3 | 100.2 | 99.6625 | 99.6
6.6 | 100.4667 | 100.3 | 100.2 | 99.6 | 99.6
6.7 | 100.4 | 100.2 | 100.0571 | 99.5 | 99.5
6.8 | 100.3 | 100.2 | 100 | 99.4714 | 99.5
6.9 | 100.3 | 100.2 | 100 | 99.4 | 99.5
7 | 100.3 | 100.1 | 100 | 99.4 | 99.4571
7.1 | 100.2857 | 100.1 | 100 | 99.3333 | 99.4
7.2 | 100.2 | 100.1 | 99.9286 | 99.2625 | 99.4
7.3 | 100.2 | 100 | 99.9 | 99.2 | 99.3333
7.4 | 100.2 | 100 | 99.9 | 99.2 | 99.3
7.5 | 100.1 | 100 | 99.8 | 99.2 | 99.3
7.6 | 100.1 | 100 | 99.8 | 99.1 | 99.2833
7.7 | 100 | 99.9 | 99.7 | 99.075 | 99.2
7.8 | 100 | 99.9 | 99.6857 | 99 | 99.16
7.9 | 99.9 | 99.8 | 99.6 | 99 | 99.1
8 | 99.9 | 99.7 | 99.5 | 99 | 99.1
8.1 | 99.9 | 99.6 | 99.4 | 99.075 | 99
8.2 | 99.8333 | 99.6 | 99.4 | 99.05 | 98.9
8.3 | 99.8 | 99.6 | 99.4 | 98.9833 | 98.8
8.4 | 99.8 | 99.58 | 99.3714 | 98.9 | 98.8
8.5 | 99.7 | 99.5 | 99.3 | 98.82 | 98.6889
8.6 | 99.7 | 99.4 | 99.3 | 98.7 | 98.575
8.7 | 99.6 | 99.4 | 99.3 | 98.7 | 98.5
8.8 | 99.5333 | 99.3143 | 99.2 | 98.6 | 98.4857
8.9 | 99.5 | 99.3 | 99.1 | 98.4857 | 98.4
9 | 99.4 | 99.2 | 99.1 | 98.4 | 98.325
9.1 | 99.4 | 99.2 | 99 | 98.3 | 98.3
9.2 | 99.35 | 99.0875 | 98.9 | 98.3 | 98.2375
9.3 | 99.2857 | 99 | 98.8286 | 98.3 | 98.2
9.4 | 99.2 | 99 | 98.8 | 98.3 | 98.1
9.5 | 99.2 | 99 | 98.7833 | 98.22 | 98
9.6 | 99.1 | 98.9 | 98.7 | 98.1571 | 97.9333
9.7 | 99 | 98.8571 | 98.7 | 98.1 | 97.86
9.8 | 99 | 98.8 | 98.7 | 98.1 | 97.8
9.9 | 98.9 | 98.6571 | 98.6 | 98 | 97.7
10 | 98.9 | 98.6 | 98.5125 | 97.9667 | 97.6
10.1 | 98.9 | 98.5333 | 98.5 | 97.825 | 97.6
10.2 | 98.8 | 98.5 | 98.4 | 97.7111 | 97.52
10.3 | 98.7125 | 98.5 | 98.4 | 97.675 | 97.4
10.4 | 98.6333 | 98.4167 | 98.3 | 97.55 | 97.4
10.5 | 98.6 | 98.4 | 98.2167 | 97.5 | 97.3
10.6 | 98.5 | 98.4 | 98.15 | 97.4625 | 97.25
10.7 | 98.5 | 98.3 | 98.1 | 97.4 | 97.2
10.8 | 98.5 | 98.2429 | 98.05 | 97.38 | 97.1
10.9 | 98.4571 | 98.18 | 98 | 97.225 | 96.9857
11 | 98.4 | 98.1 | 98 | 97.1 | 96.9
11.1 | 98.4 | 98 | 97.9 | 97.1 | 96.8429
11.2 | 98.3 | 98 | 97.9 | 97.025 | 96.8
11.3 | 98.26 | 98 | 97.8 | 96.9667 | 96.7
11.4 | 98.1 | 97.9 | 97.7714 | 96.9 | 96.625
11.5 | 98.1 | 97.8375 | 97.6 | 96.8 | 96.56
11.6 | 97.95 | 97.7857 | 97.6 | 96.8 | 96.4833
11.7 | 97.8 | 97.7 | 97.5 | 96.7 | 96.4
11.8 | 97.8 | 97.7 | 97.4167 | 96.6 | 96.3625
11.9 | 97.7 | 97.6571 | 97.3571 | 96.5 | 96.2286
12 | 97.7 | 97.6 | 97.225 | 96.4333 | 96.175
12.1 | 97.6 | 97.4714 | 97.2 | 96.3 | 96.1
12.2 | 97.6 | 97.4 | 97.1286 | 96.2375 | 95.9857
12.3 | 97.55 | 97.275 | 97.0571 | 96.2 | 95.9
12.4 | 97.4667 | 97.2 | 96.9 | 96.1875 | 95.8143
12.5 | 97.4 | 97.2 | 96.8571 | 96.1 | 95.7333
12.6 | 97.3 | 97.1 | 96.8 | 96 | 95.65
12.7 | 97.25 | 97 | 96.7 | 95.8571 | 95.5
12.8 | 97.2 | 96.9 | 96.65 | 95.7 | 95.3667
12.9 | 97.1 | 96.8 | 96.5833 | 95.7 | 95.2
13 | 97.0667 | 96.7714 | 96.4167 | 95.6 | 95.1556
13.1 | 96.9875 | 96.7 | 96.34 | 95.5286 | 95.0375
13.2 | 96.9 | 96.6 | 96.3 | 95.5 | 95
13.3 | 96.8 | 96.6 | 96.2429 | 95.4333 | 95
13.4 | 96.7625 | 96.6 | 96.2 | 95.4 | 95
13.5 | 96.7 | 96.5 | 96.1 | 95.3 | 94.9
13.6 | 96.6333 | 96.4 | 96 | 95.25 | 94.8571
13.7 | 96.5 | 96.3 | 95.9 | 95.1222 | 94.6
13.8 | 96.42 | 96.14 | 95.8286 | 95 | 94.5429
13.9 | 96.4 | 96.0667 | 95.7875 | 94.9 | 94.4
14 | 96.3 | 95.925 | 95.7 | 94.8 | 94.35
14.1 | 96.3 | 95.9 | 95.6 | 94.7167 | 94.3
14.2 | 96.2 | 95.8286 | 95.55 | 94.5714 | 94.2714
14.3 | 96.1333 | 95.8 | 95.5 | 94.4167 | 94.1286
14.4 | 96 | 95.72 | 95.4 | 94.3429 | 94.1
14.5 | 95.9333 | 95.5625 | 95.3 | 94.2667 | 93.8556
14.6 | 95.8833 | 95.5 | 95.1667 | 94.1 | 93.6625
14.7 | 95.8 | 95.4667 | 95.025 | 94 | 93.5375
14.8 | 95.7 | 95.4 | 94.9667 | 93.96 | 93.5
14.9 | 95.7 | 95.3 | 94.825 | 93.9 | 93.5
15 | 95.5857 | 95.3 | 94.74 | 93.7571 | 93.35
15.1 | 95.5 | 95.2 | 94.5857 | 93.6833 | 93.2125
15.2 | 95.45 | 95.1143 | 94.5 | 93.5167 | 93.06
15.3 | 95.4 | 95.06 | 94.5 | 93.3889 | 92.9222
15.4 | 95.3167 | 94.9857 | 94.4667 | 93.2714 | 92.9
15.5 | 95.275 | 94.9 | 94.4 | 93.1 | 92.6714
15.6 | 95.15 | 94.8 | 94.2429 | 93.0333 | 92.5286
15.7 | 95 | 94.6833 | 94.1 | 92.9 | 92.4
15.8 | 94.8 | 94.6 | 94 | 92.8167 | 92.2556
15.9 | 94.75 | 94.4625 | 93.9143 | 92.7333 | 92.2
16 | 94.7 | 94.3167 | 93.9 | 92.6 | 92.1125
16.1 | 94.6667 | 94.2429 | 93.7429 | 92.4667 | 91.9875
16.2 | 94.5143 | 94.1667 | 93.6 | 92.4 | 91.8667
16.3 | 94.4571 | 94.0143 | 93.5 | 92.2875 | 91.8
16.4 | 94.4 | 93.9 | 93.4 | 92.1571 | 91.7
16.5 | 94.3 | 93.8 | 93.3286 | 92.0143 | 91.6429
16.6 | 94.2833 | 93.62 | 93.275 | 91.8875 | 91.4167
16.7 | 94.2 | 93.54 | 93.1429 | 91.7625 | 91.2714
16.8 | 94.06 | 93.4571 | 93 | 91.6286 | 91.1286
16.9 | 93.9 | 93.3143 | 92.8714 | 91.5 | 91.0714
17 | 93.9 | 93.2333 | 92.7286 | 91.3571 | 90.9
17.1 | 93.74 | 93.1 | 92.5857 | 91.225 | 90.85
17.2 | 93.6 | 93.0833 | 92.5 | 91.1667 | 90.7
17.3 | 93.55 | 93 | 92.3833 | 91 | 90.45
17.4 | 93.3833 | 92.9 | 92.2167 | 90.82 | 90.2667
17.5 | 93.3 | 92.7286 | 92.1571 | 90.65 | 90.125
17.6 | 93.2 | 92.54 | 91.98 | 90.5 | 90
17.7 | 93.1 | 92.4 | 91.8 | 90.4 | 89.92
17.8 | 93 | 92.3143 | 91.6571 | 90.225 | 89.9
17.9 | 92.74 | 92.15 | 91.525 | 90.0778 | 89.8143
18 | 92.64 | 91.9 | 91.3143 | 89.9625 | 89.66
18.1 | 92.5833 | 91.8 | 91.1143 | 89.9 | 89.48
18.2 | 92.5 | 91.68 | 91 | 89.7714 | 89.3
18.3 | 92.3143 | 91.6 | 90.9167 | 89.575 | 89.25
18.4 | 92.1143 | 91.3875 | 90.6625 | 89.4125 | 88.9667
18.5 | 91.95 | 91.25 | 90.475 | 89.2889 | 88.7875
18.6 | 91.8 | 91.16 | 90.3 | 89.15 | 88.7
18.7 | 91.64 | 90.8875 | 90.2286 | 88.9429 | 88.5857
18.8 | 91.4667 | 90.75 | 90 | 88.78 | 88.4333
18.9 | 91.3 | 90.6 | 89.85 | 88.5875 | 88.3
19 | 91.0667 | 90.45 | 89.6 | 88.4571 | 88.275
19.1 | 90.825 | 90.2 | 89.4429 | 88.1667 | 88.15
19.2 | 90.5333 | 90 | 89.2833 | 87.9375 | 87.9286
19.3 | 90.22 | 89.75 | 89.0333 | 87.8125 | 87.75
19.4 | 89.7125 | 89.4714 | 88.7857 | 87.6625 | 87.5
19.5 | 88.9 | 89.3167 | 88.6 | 87.5286 | 87.25
19.6 | 87.9 | 89.05 | 88.4571 | 87.2875 | 87.1111
19.7 | 86.25 | 88.7667 | 88.3 | 87.125 | 87
19.8 | 83.9 | 88.3667 | 88.0333 | 86.9286 | 86.925
19.9 | 79.7333 | 87.9 | 87.8429 | 86.6875 | 86.675
20 | 74.0143 | 87.45 | 87.45 | 86.5333 | 86.5
20.1 | 66.45 | 86.8125 | 87.08 | 86.2667 | 86.2571
20.2 | 57.95 | 86 | 86.9429 | 86 | 86.1143
20.3 | 49.2 | 84.7 | 86.6571 | 85.7556 | 85.98
20.4 | 40.9333 | 82.8857 | 86.2143 | 85.5625 | 85.8
20.5 | 33.3429 | 79.9333 | 85.9286 | 85.375 | 85.6714
20.6 | 26.3 | 76.22 | 85.6 | 85.1833 | 85.5
20.7 | 20.9714 | 70.5571 | 85.3 | 84.975 | 85.2
20.8 | 17.1143 | 63.2714 | 84.7 | 84.7 | 84.9286
20.9 | 14.5857 | 55.45 | 84.2 | 84.5167 | 84.65
21 | 12.9833 | 47.2 | 83.75 | 84.2667 | 84.4833
21.1 | 12.0429 | 40.0286 | 83.1 | 84.0286 | 84.3375
21.2 | 11.42 | 33.1429 | 82.2 | 83.8 | 84.2
21.3 | 10.9667 | 27.0667 | 80.8571 | 83.5667 | 84.05
21.4 | 10.6571 | 22.25 | 78.875 | 83.2333 | 83.7833
21.5 | 10.36 | 18.675 | 76.0833 | 83 | 83.6167
21.6 | 10.125 | 16.5429 | 72.2143 | 82.725 | 83.4
21.7 | 9.95 | 14.8 | 67.3 | 82.4429 | 83.1
21.8 | 9.7333 | 13.975 | 60.9333 | 82.175 | 82.8571
21.9 | 9.6333 | 13.1833 | 53.725 | 81.8 | 82.6
22 | 9.5 | 12.7143 | 46.8 | 81.4667 | 82.3333
22.1 | 9.3333 | 12.3 | 39.75 | 81.1333 | 82.1143
22.2 | 9.2 | 11.9667 | 33.8571 | 80.9 | 81.925
22.3 | 9 | 11.65 | 28.9375 | 80.5375 | 81.76
22.4 | 8.9 | 11.4 | 24.6286 | 80.1625 | 81.6333
22.5 | 8.8 | 11.225 | 21.4857 | 79.825 | 81.3
22.6 | 8.7 | 10.9714 | 19.3333 | 79.4778 | 81
22.7 | 8.6 | 10.7333 | 17.8 | 79.1 | 80.825
22.8 | 8.45 | 10.55 | 16.7 | 78.65 | 80.575
22.9 | 8.3 | 10.4333 | 15.9857 | 77.8667 | 80.3556
23 | 8.26 | 10.2 | 15.5333 | 77.3 | 79.95
23.1 | 8.1833 | 10.1 | 14.9714 | 76.6 | 79.7429
23.2 | 8.1 | 9.95 | 14.575 | 75.4286 | 79.5
23.3 | 8 | 9.7833 | 14.2333 | 73.9333 | 79.2429
23.4 | 8 | 9.5333 | 13.9 | 72.05 | 78.9571
23.5 | 7.9 | 9.4 | 13.5667 | 69.16 | 78.7
23.6 | 7.9 | 9.3286 | 13.3333 | 65.9 | 78.4571
23.7 | 7.8857 | 9.16 | 13.1 | 61.6143 | 78.1667
23.8 | 7.8 | 9.1 | 12.825 | 55.9571 | 77.8
23.9 | 7.75 | 8.9167 | 12.575 | 50.5667 | 77.62
24 | 7.7 | 8.7667 | 12.4 | 45.0333 | 77.3143
24.1 | 7.7 | 8.7 | 12.1571 | 39.6286 | 76.9571
24.2 | 7.6333 | 8.6 | 12.0143 | 34.7 | 76.6429
24.3 | 7.6 | 8.5 | 11.8 | 30.475 | 76.4
24.4 | 7.5125 | 8.4625 | 11.5714 | 27.1 | 76.04
24.5 | 7.5 | 8.4 | 11.3333 | 24.6143 | 75.5
24.6 | 7.5 | 8.3 | 11.15 | 22.45 | 75.1667
24.7 | 7.5 | 8.25 | 10.9333 | 21.2 | 74.7
24.8 | 7.4 | 8.1667 | 10.8 | 20.225 | 74.2
24.9 | 7.4 | 8.1 | 10.7167 | 19.5125 | 73.56
25 | 7.3167 | 8.0833 | 10.5857 | 18.9857 | 73.0333
25.1 | 7.2571 | 8 | 10.4429 | 18.5 | 72.1
25.2 | 7.2 | 7.9 | 10.2 | 18.0375 | 70.95
25.3 |  | 7.9 | 10.0714 | 17.68 | 69.4625
25.4 |  | 7.8429 | 9.9167 | 17.3429 | 67.5571
25.5 |  | 7.8 | 9.8333 | 17.0333 | 65.0333
25.6 |  | 7.7 | 9.75 | 16.7286 | 62.0571
25.7 |  | 7.6286 | 9.5875 | 16.4333 | 58.0429
25.8 |  | 7.6 | 9.5 | 16.1333 | 53.5
25.9 |  | 7.5 | 9.425 | 15.8833 | 48.55
26 |  |  | 9.3667 | 15.675 | 43.6
26.1 |  |  | 9.2 | 15.3375 | 39.1
26.2 |  |  | 9.1333 | 15.0857 | 35
26.3 |  |  | 9.05 | 14.9 | 31.7714
26.4 |  |  | 8.9 | 14.7143 | 28.6
26.5 |  |  | 8.8833 | 14.4667 | 26.55
26.6 |  |  | 8.8 | 14.2375 | 24.85
26.7 |  |  | 8.6714 | 14.025 | 23.5
26.8 |  |  | 8.6 | 13.8 | 22.42
26.9 |  |  | 8.5 | 13.6333 | 21.7
27 |  |  |  | 13.4 | 21.18
27.1 |  |  |  | 13.2429 | 20.75
27.2 |  |  |  | 12.9857 | 20.4
27.3 |  |  |  | 12.85 | 19.95
27.4 |  |  |  | 12.6667 | 19.5571
27.5 |  |  |  | 12.4857 | 19.25
27.6 |  |  |  | 12.3429 | 18.9857
27.7 |  |  |  | 12.1833 | 18.7
27.8 |  |  |  | 12.0167 | 18.425
27.9 |  |  |  | 11.8875 | 18.1333
28 |  |  |  | 11.7625 | 17.94
28.1 |  |  |  | 11.6444 | 17.7
28.2 |  |  |  | 11.525 | 17.4714
28.3 |  |  |  | 11.4 | 17.1857
28.4 |  |  |  | 11.24 | 16.925
28.5 |  |  |  | 11.0667 | 16.6429
28.6 |  |  |  | 10.925 | 16.4167
28.7 |  |  |  | 10.9 | 16.25
28.8 |  |  |  | 10.775 | 16.0714
28.9 |  |  |  | 10.65 | 15.8
29 |  |  |  |  | 15.55
29.1 |  |  |  |  | 15.4111
29.2 |  |  |  |  | 15.375
29.3 |  |  |  |  | 15.15
29.4 |  |  |  |  | 14.9667
29.5 |  |  |  |  | 14.7
29.6 |  |  |  |  | 14.5625
29.7 |  |  |  |  | 14.3333
29.8 |  |  |  |  | 14.2
29.9 |  |  |  |  | 13.975
30 |  |  |  |  | 13.8
30.1 |  |  |  |  | 13.625
30.2 |  |  |  |  | 13.475
30.3 |  |  |  |  | 13.3333
30.4 |  |  |  |  | 13.24
30.5 |  |  |  |  | 13.075
30.6 |  |  |  |  | 12.8625
30.7 |  |  |  |  | 12.7375
30.8 |  |  |  |  | 12.7
30.9 |  |  |  |  | 12.6

## Wedge 60 Output Factors
Output factors
Wedge 60 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15
| 3 | 0.878 | 0.888 | 0.896 | 0.906 | 0.913 | 0.921
| 4 | 0.89 | 0.903 | 0.912 | 0.924 | 0.934 | 0.943
| 5 | 0.897 | 0.913 | 0.924 | 0.939 | 0.95 | 0.962
| 7 | 0.91 | 0.929 | 0.942 | 0.96 | 0.976 | 0.992
| 10 | 0.921 | 0.942 | 0.958 | 0.981 | 1 | 1.022
| 15 | 0.931 | 0.955 | 0.972 | 0.999 | 1.027 | 1.055
| 20 | 0.936 | 0.962 | 0.983 | 1.012 | 1.043 | 1.077
| 30 | 0.944 | 0.972 | 0.995 | 1.028 | 1.064 | 1.107
| 40 | 0.949 | 0.978 | 1.001 | 1.037 | 1.077 | 1.124